## Index
Thesis Title | Development of a novel bioreactor and systems for suspension cell culture in biopharmaceutical production |  |  | Worksheet details
|  |  |  | 1 | Design sheet and Users' requirement specifications (URS)
Chapter Title | 4. Design, Fabrication and abiotic characterisation of HTB-(Results and Discussion - Part I) |  |  | 2 | Impeller immersion ratios
|  |  |  | 3 | Data for calculating mixing time through conductivity and F (t) curve
Authors | Rajesh Sharma |  |  | 4 | Mixing time through conductivity method and modelling
|  |  |  | 5 | Determination of minimum agitation speed at different impeller immersion ratios from STR model
Institution | Centre for Bioprocess Engineering Research (CeBER), Department of Chemical Engineering, University of Cape Town, Private Bag X3, Rondebosch 7700, South Africa |  |  | 6 | Determination of minimum agitation speed at different impeller immersion ratios from HTB model
|  |  |  | 7 | Modelling of oxygen mass transfer (Buckingham Pi method)
|  |  |  | 8 | Sample example for the determination of mass transfer coefficient (Kla) for 2 L (42 % impeller imersion).
Index | Reference to article | Sheet no. | Title of items | 9 | Sample calculations for overall heat transfer coefficient (UA) for HTB at 2 L and 3 L volume
| Figure 4-13 | 3 | The F (t) curves representing the conductivity profile of different fill volumes at a different impeller speed | 10 | Sample calculation for power utilisation in HTB at 2 L and 3 L with 200 rpm
| Figure 4-14 | 4 | Relationship between mixing time and impeller speed at different fill volumes.
| Figure 4-15 | 4 | Best-fit trendline elucidating the relationships of mixing time with (a) impeller immersion and (b) with impeller speed
| Figure 4-16 | 4 | Means of observation of the two factors depicting the interaction
| Figure 4-17 | 4 | Residual plots for all the mixing models with respect to impeller immersion and impeller speed
| Figure 4-18 | 4 | Relationship of mixing time with (a) impeller immersion and (b) impeller speed in comparison with model 3
| Figure 4-19 | 7 | model comparison to experimental values for 2 L (left), and 3 L (right) volumes
| Figure 4-20 | 7 | model comparison to experimental values for 2 L (left), and 3 L (right) volumes
| Figure 4-21 | 6 | Relationship between Zwietering's S coefficient and impeller immersion.
Figure 4-22
| Figure 4-23 | 6 | Residual plot of modified Zwietering equation with respect to impeller immersion
| Figure 4-24 | 6 | Experimental and predicted values of minimum suspension speed with respect to impeller immersion.
| Figure 4-25 | 10 | Power consumption profile of HTB at different volumes at various agitation speeds
| Figure 4-26 | 10 | Mixing time Vs power consumption with different impeller immersion ratios at different impeller speeds
| Figure 4-27 | 10 | Mass transfer efficiency Vs power input at (a) 42 % and (b) 79 % impeller immersion
| Table 4-2 | 1 | URS
| Table 4-5 | 4 | Analysis of mixing time data with Two-way ANOVA with replication
Table 4-6
| Table 4-10 | 8 | Experimentally calculated values of (h-1) for 2 L volume (42 % impeller immersion)
| Table 4-11 | 8 | Experimentally calculated values of (h-1) for 3 L volume (79 % impeller immersion)
| Table 4-12 | 4 | Predicted values of power numbers and circulation flow numbers

## 1
| Length of the vessel (m) | [12 empty] | Length of the vessel (m)
L/D | 0.35 |  | 0.4 |  | 0.45 |  | 0.5 |  | 0.55 |  | 0.6 |  | L/D | 0.35 |  | 0.4 |  | 0.45 |  | 0.5 |  | 0.55 |  | 0.6 |  | Tan θ | 15 | 30 | 45 | 60 | π
1.0 | D=H | 0.35 | D=H | 0.4 | D=H | 0.45 | D=H | 0.5 | D=H | 0.55 | D=H | 0.6 | 1.8 | D=H | 0.19444 | D=H | 0.22222 | D=H | 0.25 | D=H | 0.27778 | D=H | 0.30556 | D=H | 0.33333 |  | 0.308 | 0.577 | 1 | 1.732 | 3.14
| d=D/2.2 | 0.15909 | d=D/2.2 | 0.18182 | d=D/2.2 | 0.20455 | d=D/2.2 | 0.22727 | d=D/2.2 | 0.25 | d=D/2.2 | 0.27273 |  | d=D/2.2 | 0.088384 | d=D/2.2 | 0.10101 | d=D/2.2 | 0.11364 | d=D/2.2 | 0.12626 | d=D/2.2 | 0.13889 | d=D/2.2 | 0.15152
| h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 |  | h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 | L/D | dm (Mean Diameter)
| D=H=2.2d | 0.35 | D=H=2.2d | 0.4 | D=H=2.2d | 0.45 | D=H=2.2d | 0.5 | D=H=2.2d | 0.55 | D=H=2.2d | 0.6 |  | D=H=2.2d | 0.19444 | D=H=2.2d | 0.22222 | D=H=2.2d | 0.25 | D=H=2.2d | 0.27778 | D=H=2.2d | 0.30556 | D=H=2.2d | 0.33333 |  | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
| ds=0.2d | 0.031818 | ds=0.2d | 0.036364 | ds=0.2d | 0.040909 | ds=0.2d | 0.045455 | ds=0.2d | 0.05 | ds=0.2d | 0.054545 |  | ds=0.2d | 0.017677 | ds=0.2d | 0.020202 | ds=0.2d | 0.022727 | ds=0.2d | 0.025253 | ds=0.2d | 0.027778 | ds=0.2d | 0.030303 | 1 | 0.095455 | 0.10909 | 0.12273 | 0.13636 | 0.15 | 0.16364
| C=0.05d | 0.0079545 | C=0.05d | 0.0090909 | C=0.05d | 0.010227 | C=0.05d | 0.011364 | C=0.05d | 0.0125 | C=0.05d | 0.013636 |  | C=0.05d | 0.0044192 | C=0.05d | 0.0050505 | C=0.05d | 0.0056818 | C=0.05d | 0.0063131 | C=0.05d | 0.0069444 | C=0.05d | 0.0075758 | 1.2 | 0.079545 | 0.090909 | 0.10227 | 0.11364 | 0.125 | 0.13636
| S=0.6d | 0.095455 | S=0.6d | 0.10909 | S=0.6d | 0.12273 | S=0.6d | 0.13636 | S=0.6d | 0.15 | S=0.6d | 0.16364 |  | S=0.6d | 0.05303 | S=0.6d | 0.060606 | S=0.6d | 0.068182 | S=0.6d | 0.075758 | S=0.6d | 0.083333 | S=0.6d | 0.090909 | 1.5 | 0.063636 | 0.072727 | 0.081818 | 0.090909 | 0.1 | 0.10909
| N (Turns) | 2.969 | N (Turns) | 2.969 | N (Turns) | 2.969 | N (Turns) | 2.969 | N (Turns) | 2.969 | N (Turns) | 2.969 |  | N (Turns) | 5.3441 | N (Turns) | 5.3441 | N (Turns) | 5.3441 | N (Turns) | 5.3441 | N (Turns) | 5.3441 | N (Turns) | 5.3441 | 1.8 | 0.05303 | 0.060606 | 0.068182 | 0.075758 | 0.083333 | 0.090909
| W=(d-ds)/2 | 0.063636 | W=(d-ds)/2 | 0.072727 | W=(d-ds)/2 | 0.081818 | W=(d-ds)/2 | 0.090909 | W=(d-ds)/2 | 0.1 | W=(d-ds)/2 | 0.10909 |  | W=(d-ds)/2 | 0.035354 | W=(d-ds)/2 | 0.040404 | W=(d-ds)/2 | 0.045455 | W=(d-ds)/2 | 0.050505 | W=(d-ds)/2 | 0.055556 | W=(d-ds)/2 | 0.060606 | 2 | 0.047727 | 0.054545 | 0.061364 | 0.068182 | 0.075 | 0.081818
| r | 0.175 | r | 0.2 | r | 0.225 | r | 0.25 | r | 0.275 | r | 0.3 |  | r | 0.097222 | r | 0.11111 | r | 0.125 | r | 0.13889 | r | 0.15278 | r | 0.16667 | 2.5 | 0.038182 | 0.043636 | 0.049091 | 0.054545 | 0.06 | 0.065455
| V (m^3) | 0.033657 | V (m^3) | 0.05024 | V (m^3) | 0.071533 | V (m^3) | 0.098125 | V (m^3) | 0.1306 | V (m^3) | 0.16956 |  | V (m^3) | 0.010388 | V (m^3) | 0.015506 | V (m^3) | 0.022078 | V (m^3) | 0.030285 | V (m^3) | 0.04031 | V (m^3) | 0.052333 | 3 | 0.031818 | 0.036364 | 0.040909 | 0.045455 | 0.05 | 0.054545
| V (m^3)70% | 0.02356 | V (m^3)70% | 0.035168 | V (m^3)70% | 0.050073 | V (m^3)70% | 0.068687 | V (m^3)70% | 0.091423 | V (m^3)70% | 0.11869 |  | V (m^3)70% | 0.0072715 | V (m^3)70% | 0.010854 | V (m^3)70% | 0.015455 | V (m^3)70% | 0.0212 | V (m^3)70% | 0.028217 | V (m^3)70% | 0.036633 | 3.5 | 0.027273 | 0.031169 | 0.035065 | 0.038961 | 0.042857 | 0.046753
| HL (mm) | 234 | HL (mm) | 262 | HL (mm) | 296 | HL (mm) | 331 | HL (mm) | 361 | HL (mm) | 396 |  | HL (mm) | 124 | HL (mm) | 148 | HL (mm) | 160 | HL (mm) | 181 | HL (mm) | 200 | HL (mm) | 221 | 4 | 0.023864 | 0.027273 | 0.030682 | 0.034091 | 0.0375 | 0.040909
[26 empty] | 4.3 | 0.022199 | 0.02537 | 0.028541 | 0.031712 | 0.034884 | 0.038055
1.2 | D=H | 0.29167 | D=H | 0.33333 | D=H | 0.375 | D=H | 0.41667 | D=H | 0.45833 | D=H | 0.5 | 2.0 | D=H | 0.175 | D=H | 0.2 | D=H | 0.225 | D=H | 0.25 | D=H | 0.275 | D=H | 0.3 | 4.6 | 0.020751 | 0.023715 | 0.02668 | 0.029644 | 0.032609 | 0.035573
| d=D/2.2 | 0.13258 | d=D/2.2 | 0.15152 | d=D/2.2 | 0.17045 | d=D/2.2 | 0.18939 | d=D/2.2 | 0.20833 | d=D/2.2 | 0.22727 |  | d=D/2.2 | 0.079545 | d=D/2.2 | 0.090909 | d=D/2.2 | 0.10227 | d=D/2.2 | 0.11364 | d=D/2.2 | 0.125 | d=D/2.2 | 0.13636 | 5 | 0.019091 | 0.021818 | 0.024545 | 0.027273 | 0.03 | 0.032727
| h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 |  | h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583
| D=H=2.2d | 0.29167 | D=H=2.2d | 0.33333 | D=H=2.2d | 0.375 | D=H=2.2d | 0.41667 | D=H=2.2d | 0.45833 | D=H=2.2d | 0.5 |  | D=H=2.2d | 0.175 | D=H=2.2d | 0.2 | D=H=2.2d | 0.225 | D=H=2.2d | 0.25 | D=H=2.2d | 0.275 | D=H=2.2d | 0.3 | L/D | Pitch (S) at 15 deg | [5 empty] | L/D | No. of spiral turns on the shaft (n) at 15 deg
| ds=0.2d | 0.026515 | ds=0.2d | 0.030303 | ds=0.2d | 0.034091 | ds=0.2d | 0.037879 | ds=0.2d | 0.041667 | ds=0.2d | 0.045455 |  | ds=0.2d | 0.015909 | ds=0.2d | 0.018182 | ds=0.2d | 0.020455 | ds=0.2d | 0.022727 | ds=0.2d | 0.025 | ds=0.2d | 0.027273 |  | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 |  | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
| C=0.05d | 0.0066288 | C=0.05d | 0.0075758 | C=0.05d | 0.0085227 | C=0.05d | 0.0094697 | C=0.05d | 0.010417 | C=0.05d | 0.011364 |  | C=0.05d | 0.0039773 | C=0.05d | 0.0045455 | C=0.05d | 0.0051136 | C=0.05d | 0.0056818 | C=0.05d | 0.00625 | C=0.05d | 0.0068182 | 1 | 0.092316 | 0.1055 | 0.11869 | 0.13188 | 0.14507 | 0.15826 | 1 | 3.0699 | 3.0699 | 3.0699 | 3.0699 | 3.0699 | 3.0699
| S=0.6d | 0.079545 | S=0.6d | 0.090909 | S=0.6d | 0.10227 | S=0.6d | 0.11364 | S=0.6d | 0.125 | S=0.6d | 0.13636 |  | S=0.6d | 0.047727 | S=0.6d | 0.054545 | S=0.6d | 0.061364 | S=0.6d | 0.068182 | S=0.6d | 0.075 | S=0.6d | 0.081818 | 1.2 | 0.07693 | 0.08792 | 0.09891 | 0.1099 | 0.12089 | 0.13188 | 1.2 | 3.6839 | 3.6839 | 3.6839 | 3.6839 | 3.6839 | 3.6839
| N (Turns) | 3.5628 | N (Turns) | 3.5628 | N (Turns) | 3.5628 | N (Turns) | 3.5628 | N (Turns) | 3.5628 | N (Turns) | 3.5628 |  | N (Turns) | 5.9379 | N (Turns) | 5.9379 | N (Turns) | 5.9379 | N (Turns) | 5.9379 | N (Turns) | 5.9379 | N (Turns) | 5.9379 | 1.5 | 0.061544 | 0.070336 | 0.079128 | 0.08792 | 0.096712 | 0.1055 | 1.5 | 4.6048 | 4.6048 | 4.6048 | 4.6048 | 4.6048 | 4.6048
| W=(d-ds)/2 | 0.05303 | W=(d-ds)/2 | 0.060606 | W=(d-ds)/2 | 0.068182 | W=(d-ds)/2 | 0.075758 | W=(d-ds)/2 | 0.083333 | W=(d-ds)/2 | 0.090909 |  | W=(d-ds)/2 | 0.031818 | W=(d-ds)/2 | 0.036364 | W=(d-ds)/2 | 0.040909 | W=(d-ds)/2 | 0.045455 | W=(d-ds)/2 | 0.05 | W=(d-ds)/2 | 0.054545 | 1.8 | 0.051287 | 0.058613 | 0.06594 | 0.073267 | 0.080593 | 0.08792 | 1.8 | 5.5258 | 5.5258 | 5.5258 | 5.5258 | 5.5258 | 5.5258
| r | 0.14583 | r | 0.16667 | r | 0.1875 | r | 0.20833 | r | 0.22917 | r | 0.25 |  | r | 0.0875 | r | 0.1 | r | 0.1125 | r | 0.125 | r | 0.1375 | r | 0.15 | 2 | 0.046158 | 0.052752 | 0.059346 | 0.06594 | 0.072534 | 0.079128 | 2 | 6.1398 | 6.1398 | 6.1398 | 6.1398 | 6.1398 | 6.1398
| V (m^3) | 0.023373 | V (m^3) | 0.034889 | V (m^3) | 0.049676 | V (m^3) | 0.068142 | V (m^3) | 0.090697 | V (m^3) | 0.11775 |  | V (m^3) | 0.0084142 | V (m^3) | 0.01256 | V (m^3) | 0.017883 | V (m^3) | 0.024531 | V (m^3) | 0.032651 | V (m^3) | 0.04239 | 2.5 | 0.036926 | 0.042202 | 0.047477 | 0.052752 | 0.058027 | 0.063302 | 2.5 | 7.6747 | 7.6747 | 7.6747 | 7.6747 | 7.6747 | 7.6747
| V (m^3)70% | 0.016361 | V (m^3)70% | 0.024422 | V (m^3)70% | 0.034773 | V (m^3)70% | 0.0477 | V (m^3)70% | 0.063488 | V (m^3)70% | 0.082425 |  | V (m^3)70% | 0.00589 | V (m^3)70% | 0.008792 | V (m^3)70% | 0.012518 | V (m^3)70% | 0.017172 | V (m^3)70% | 0.022856 | V (m^3)70% | 0.029673 | 3 | 0.030772 | 0.035168 | 0.039564 | 0.04396 | 0.048356 | 0.052752 | 3 | 9.2097 | 9.2097 | 9.2097 | 9.2097 | 9.2097 | 9.2097
| HL (mm) | 188 | HL (mm) | 216 | HL (mm) | 248 | HL (mm) | 276 | HL (mm) | 300 | HL (mm) | 328 |  | HL (mm) | 117 | HL (mm) | 134 | HL (mm) | 153 | HL (mm) | 163 | HL (mm) | 182 | HL (mm) | 199 | 3.5 | 0.026376 | 0.030144 | 0.033912 | 0.03768 | 0.041448 | 0.045216 | 3.5 | 10.7446 | 10.7446 | 10.7446 | 10.7446 | 10.7446 | 10.7446
[26 empty] | 4 | 0.023079 | 0.026376 | 0.029673 | 0.03297 | 0.036267 | 0.039564 | 4 | 12.2796 | 12.2796 | 12.2796 | 12.2796 | 12.2796 | 12.2796
1.5 | D=H | 0.23333 | D=H | 0.26667 | D=H | 0.3 | D=H | 0.33333 | D=H | 0.36667 | D=H | 0.4 | 2.5 | D=H | 0.14 | D=H | 0.16 | D=H | 0.18 | D=H | 0.2 | D=H | 0.22 | D=H | 0.24 | 4.3 | 0.021469 | 0.024536 | 0.027603 | 0.03067 | 0.033737 | 0.036804 | 4.3 | 13.2006 | 13.2006 | 13.2006 | 13.2006 | 13.2006 | 13.2006
| d=D/2.2 | 0.10606 | d=D/2.2 | 0.12121 | d=D/2.2 | 0.13636 | d=D/2.2 | 0.15152 | d=D/2.2 | 0.16667 | d=D/2.2 | 0.18182 |  | d=D/2.2 | 0.063636 | d=D/2.2 | 0.072727 | d=D/2.2 | 0.081818 | d=D/2.2 | 0.090909 | d=D/2.2 | 0.1 | d=D/2.2 | 0.10909 | 4.6 | 0.020069 | 0.022936 | 0.025803 | 0.02867 | 0.031537 | 0.034403 | 4.6 | 14.1215 | 14.1215 | 14.1215 | 14.1215 | 14.1215 | 14.1215
| h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 |  | h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 | 5 | 0.018463 | 0.021101 | 0.023738 | 0.026376 | 0.029014 | 0.031651 | 5 | 15.3495 | 15.3495 | 15.3495 | 15.3495 | 15.3495 | 15.3495
| D=H=2.2d | 0.23333 | D=H=2.2d | 0.26667 | D=H=2.2d | 0.3 | D=H=2.2d | 0.33333 | D=H=2.2d | 0.36667 | D=H=2.2d | 0.4 |  | D=H=2.2d | 0.14 | D=H=2.2d | 0.16 | D=H=2.2d | 0.18 | D=H=2.2d | 0.2 | D=H=2.2d | 0.22 | D=H=2.2d | 0.24
| ds=0.2d | 0.021212 | ds=0.2d | 0.024242 | ds=0.2d | 0.027273 | ds=0.2d | 0.030303 | ds=0.2d | 0.033333 | ds=0.2d | 0.036364 |  | ds=0.2d | 0.012727 | ds=0.2d | 0.014545 | ds=0.2d | 0.016364 | ds=0.2d | 0.018182 | ds=0.2d | 0.02 | ds=0.2d | 0.021818 | L/D | Pitch (S) at 30 deg | [5 empty] | L/D | No. of spiral turns on the shaft (n) at 30 deg
| C=0.05d | 0.005303 | C=0.05d | 0.0060606 | C=0.05d | 0.0068182 | C=0.05d | 0.0075758 | C=0.05d | 0.0083333 | C=0.05d | 0.0090909 |  | C=0.05d | 0.0031818 | C=0.05d | 0.0036364 | C=0.05d | 0.0040909 | C=0.05d | 0.0045455 | C=0.05d | 0.005 | C=0.05d | 0.0054545 |  | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 |  | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
| S=0.6d | 0.063636 | S=0.6d | 0.072727 | S=0.6d | 0.081818 | S=0.6d | 0.090909 | S=0.6d | 0.1 | S=0.6d | 0.10909 |  | S=0.6d | 0.038182 | S=0.6d | 0.043636 | S=0.6d | 0.049091 | S=0.6d | 0.054545 | S=0.6d | 0.06 | S=0.6d | 0.065455 | 1 | 0.17294 | 0.19765 | 0.22235 | 0.24706 | 0.27177 | 0.29647 | 1 | 1.6387 | 1.6387 | 1.6387 | 1.6387 | 1.6387 | 1.6387
| N (Turns) | 4.4534 | N (Turns) | 4.4534 | N (Turns) | 4.4534 | N (Turns) | 4.4534 | N (Turns) | 4.4534 | N (Turns) | 4.4534 |  | N (Turns) | 7.4224 | N (Turns) | 7.4224 | N (Turns) | 7.4224 | N (Turns) | 7.4224 | N (Turns) | 7.4224 | N (Turns) | 7.4224 | 1.2 | 0.14412 | 0.16471 | 0.1853 | 0.20588 | 0.22647 | 0.24706 | 1.2 | 1.9664 | 1.9664 | 1.9664 | 1.9664 | 1.9664 | 1.9664
| W=(d-ds)/2 | 0.042424 | W=(d-ds)/2 | 0.048485 | W=(d-ds)/2 | 0.054545 | W=(d-ds)/2 | 0.060606 | W=(d-ds)/2 | 0.066667 | W=(d-ds)/2 | 0.072727 |  | W=(d-ds)/2 | 0.025455 | W=(d-ds)/2 | 0.029091 | W=(d-ds)/2 | 0.032727 | W=(d-ds)/2 | 0.036364 | W=(d-ds)/2 | 0.04 | W=(d-ds)/2 | 0.043636 | 1.5 | 0.1153 | 0.13177 | 0.14824 | 0.16471 | 0.18118 | 0.19765 | 1.5 | 2.458 | 2.458 | 2.458 | 2.458 | 2.458 | 2.458
| r | 0.11667 | r | 0.13333 | r | 0.15 | r | 0.16667 | r | 0.18333 | r | 0.2 |  | r | 0.07 | r | 0.08 | r | 0.09 | r | 0.1 | r | 0.11 | r | 0.12 | 1.8 | 0.096079 | 0.1098 | 0.12353 | 0.13726 | 0.15098 | 0.16471 | 1.8 | 2.9497 | 2.9497 | 2.9497 | 2.9497 | 2.9497 | 2.9497
| V (m^3) | 0.014959 | V (m^3) | 0.022329 | V (m^3) | 0.031793 | V (m^3) | 0.043611 | V (m^3) | 0.058046 | V (m^3) | 0.07536 |  | V (m^3) | 0.0053851 | V (m^3) | 0.0080384 | V (m^3) | 0.011445 | V (m^3) | 0.0157 | V (m^3) | 0.020897 | V (m^3) | 0.02713 | 2 | 0.086471 | 0.098824 | 0.11118 | 0.12353 | 0.13588 | 0.14824 | 2 | 3.2774 | 3.2774 | 3.2774 | 3.2774 | 3.2774 | 3.2774
| V (m^3)70% | 0.010471 | V (m^3)70% | 0.01563 | V (m^3)70% | 0.022255 | V (m^3)70% | 0.030528 | V (m^3)70% | 0.040632 | V (m^3)70% | 0.052752 |  | V (m^3)70% | 0.0037696 | V (m^3)70% | 0.0056269 | V (m^3)70% | 0.0080117 | V (m^3)70% | 0.01099 | V (m^3)70% | 0.014628 | V (m^3)70% | 0.018991 | 2.5 | 0.069177 | 0.079059 | 0.088942 | 0.098824 | 0.10871 | 0.11859 | 2.5 | 4.0967 | 4.0967 | 4.0967 | 4.0967 | 4.0967 | 4.0967
| HL (mm) | 148 | HL (mm) | 179 | HL (mm) | 195 | HL (mm) | 223 | HL (mm) | 243 | HL (mm) | 264 |  | HL (mm) | 97.3 | HL (mm) | 111.7 | HL (mm) | 118.5 | HL (mm) | 132 | HL (mm) | 148.3 | HL (mm) | 158.3 | 3 | 0.057648 | 0.065883 | 0.074118 | 0.082354 | 0.090589 | 0.098824 | 3 | 4.9161 | 4.9161 | 4.9161 | 4.9161 | 4.9161 | 4.9161
[26 empty] | 3.5 | 0.049412 | 0.056471 | 0.06353 | 0.070589 | 0.077648 | 0.084707 | 3.5 | 5.7354 | 5.7354 | 5.7354 | 5.7354 | 5.7354 | 5.7354
| Length of the vessel (m) | [12 empty] | Length of the vessel (m) | [11 empty] | 4 | 0.043236 | 0.049412 | 0.055589 | 0.061765 | 0.067942 | 0.074118 | 4 | 6.5548 | 6.5548 | 6.5548 | 6.5548 | 6.5548 | 6.5548
L/D | 0.35 |  | 0.4 |  | 0.45 |  | 0.5 |  | 0.55 |  | 0.6 |  | L/D | 0.35 |  | 0.4 |  | 0.45 |  | 0.5 |  | 0.55 |  | 0.6 |  | 4.3 | 0.040219 | 0.045965 | 0.05171 | 0.057456 | 0.063202 | 0.068947 | 4.3 | 7.0464 | 7.0464 | 7.0464 | 7.0464 | 7.0464 | 7.0464
3 | D=H | 0.11667 | D=H | 0.13333 | D=H | 0.15 | D=H | 0.16667 | D=H | 0.18333 | D=H | 0.2 | 4.3 | D=H | 0.081395 | D=H | 0.093023 | D=H | 0.10465 | D=H | 0.11628 | D=H | 0.12791 | D=H | 0.13953 | 4.6 | 0.037596 | 0.042967 | 0.048338 | 0.053709 | 0.05908 | 0.064451 | 4.6 | 7.538 | 7.538 | 7.538 | 7.538 | 7.538 | 7.538
| d=D/2.2 | 0.05303 | d=D/2.2 | 0.060606 | d=D/2.2 | 0.068182 | d=D/2.2 | 0.075758 | d=D/2.2 | 0.083333 | d=D/2.2 | 0.090909 |  | d=D/2.2 | 0.036998 | d=D/2.2 | 0.042283 | d=D/2.2 | 0.047569 | d=D/2.2 | 0.052854 | d=D/2.2 | 0.05814 | d=D/2.2 | 0.063425 | 5 | 0.034589 | 0.03953 | 0.044471 | 0.049412 | 0.054353 | 0.059295 | 5 | 8.1935 | 8.1935 | 8.1935 | 8.1935 | 8.1935 | 8.1935
| h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 |  | h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583
| D=H=2.2d | 0.11667 | D=H=2.2d | 0.13333 | D=H=2.2d | 0.15 | D=H=2.2d | 0.16667 | D=H=2.2d | 0.18333 | D=H=2.2d | 0.2 |  | D=H=2.2d | 0.081395 | D=H=2.2d | 0.093023 | D=H=2.2d | 0.10465 | D=H=2.2d | 0.11628 | D=H=2.2d | 0.12791 | D=H=2.2d | 0.13953 | L/D | Pitch (S) at 45 deg | [5 empty] | L/D | No. of spiral turns on the shaft (n) at 45 deg
| ds=0.2d | 0.010606 | ds=0.2d | 0.012121 | ds=0.2d | 0.013636 | ds=0.2d | 0.015152 | ds=0.2d | 0.016667 | ds=0.2d | 0.018182 |  | ds=0.2d | 0.0073996 | ds=0.2d | 0.0084567 | ds=0.2d | 0.0095137 | ds=0.2d | 0.010571 | ds=0.2d | 0.011628 | ds=0.2d | 0.012685 |  | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 |  | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
| C=0.05d | 0.0026515 | C=0.05d | 0.0030303 | C=0.05d | 0.0034091 | C=0.05d | 0.0037879 | C=0.05d | 0.0041667 | C=0.05d | 0.0045455 |  | C=0.05d | 0.0018499 | C=0.05d | 0.0021142 | C=0.05d | 0.0023784 | C=0.05d | 0.0026427 | C=0.05d | 0.002907 | C=0.05d | 0.0031712 | 1 | 0.29973 | 0.34255 | 0.38536 | 0.42818 | 0.471 | 0.51382 | 1 | 0.94553 | 0.94553 | 0.94553 | 0.94553 | 0.94553 | 0.94553
| S=0.6d | 0.031818 | S=0.6d | 0.036364 | S=0.6d | 0.040909 | S=0.6d | 0.045455 | S=0.6d | 0.05 | S=0.6d | 0.054545 |  | S=0.6d | 0.022199 | S=0.6d | 0.02537 | S=0.6d | 0.028541 | S=0.6d | 0.031712 | S=0.6d | 0.034884 | S=0.6d | 0.038055 | 1.2 | 0.24977 | 0.28545 | 0.32114 | 0.35682 | 0.3925 | 0.42818 | 1.2 | 1.1346 | 1.1346 | 1.1346 | 1.1346 | 1.1346 | 1.1346
| N (Turns) | 8.9069 | N (Turns) | 8.9069 | N (Turns) | 8.9069 | N (Turns) | 8.9069 | N (Turns) | 8.9069 | N (Turns) | 8.9069 |  | N (Turns) | 12.7665 | N (Turns) | 12.7665 | N (Turns) | 12.7665 | N (Turns) | 12.7665 | N (Turns) | 12.7665 | N (Turns) | 12.7665 | 1.5 | 0.19982 | 0.22836 | 0.25691 | 0.28545 | 0.314 | 0.34255 | 1.5 | 1.4183 | 1.4183 | 1.4183 | 1.4183 | 1.4183 | 1.4183
| W=(d-ds)/2 | 0.021212 | W=(d-ds)/2 | 0.024242 | W=(d-ds)/2 | 0.027273 | W=(d-ds)/2 | 0.030303 | W=(d-ds)/2 | 0.033333 | W=(d-ds)/2 | 0.036364 |  | W=(d-ds)/2 | 0.014799 | W=(d-ds)/2 | 0.016913 | W=(d-ds)/2 | 0.019027 | W=(d-ds)/2 | 0.021142 | W=(d-ds)/2 | 0.023256 | W=(d-ds)/2 | 0.02537 | 1.8 | 0.16652 | 0.1903 | 0.21409 | 0.23788 | 0.26167 | 0.28545 | 1.8 | 1.702 | 1.702 | 1.702 | 1.702 | 1.702 | 1.702
| r | 0.058333 | r | 0.066667 | r | 0.075 | r | 0.083333 | r | 0.091667 | r | 0.1 |  | r | 0.040698 | r | 0.046512 | r | 0.052326 | r | 0.05814 | r | 0.063953 | r | 0.069767 | 2 | 0.14986 | 0.17127 | 0.19268 | 0.21409 | 0.2355 | 0.25691 | 2 | 1.8911 | 1.8911 | 1.8911 | 1.8911 | 1.8911 | 1.8911
| V (m^3) | 0.0037397 | V (m^3) | 0.0055822 | V (m^3) | 0.0079481 | V (m^3) | 0.010903 | V (m^3) | 0.014512 | V (m^3) | 0.01884 |  | V (m^3) | 0.0018203 | V (m^3) | 0.0027171 | V (m^3) | 0.0038687 | V (m^3) | 0.0053069 | V (m^3) | 0.0070635 | V (m^3) | 0.0091704 | 2.5 | 0.11989 | 0.13702 | 0.15415 | 0.17127 | 0.1884 | 0.20553 | 2.5 | 2.3638 | 2.3638 | 2.3638 | 2.3638 | 2.3638 | 2.3638
| V (m^3)70% | 0.0026178 | V (m^3)70% | 0.0039076 | V (m^3)70% | 0.0055637 | V (m^3)70% | 0.0076319 | V (m^3)70% | 0.010158 | V (m^3)70% | 0.013188 |  | V (m^3)70% | 0.0012742 | V (m^3)70% | 0.001902 | V (m^3)70% | 0.0027081 | V (m^3)70% | 0.0037148 | V (m^3)70% | 0.0049445 | V (m^3)70% | 0.0064193 | 3 | 0.099909 | 0.11418 | 0.12845 | 0.14273 | 0.157 | 0.17127 | 3 | 2.8366 | 2.8366 | 2.8366 | 2.8366 | 2.8366 | 2.8366
| HL (mm) | 234 | HL (mm) | 262 | HL (mm) | 296 | HL (mm) | 331 | HL (mm) | 361 | HL (mm) | 396 |  | HL (mm) | 124 | HL (mm) | 148 | HL (mm) | 160 | HL (mm) | 181 | HL (mm) | 200 | HL (mm) | 221 | 3.5 | 0.085636 | 0.09787 | 0.1101 | 0.12234 | 0.13457 | 0.14681 | 3.5 | 3.3094 | 3.3094 | 3.3094 | 3.3094 | 3.3094 | 3.3094
[26 empty] | 4 | 0.074932 | 0.085636 | 0.096341 | 0.10705 | 0.11775 | 0.12845 | 4 | 3.7821 | 3.7821 | 3.7821 | 3.7821 | 3.7821 | 3.7821
3.5 | D=H | 0.1 | D=H | 0.11429 | D=H | 0.12857 | D=H | 0.14286 | D=H | 0.15714 | D=H | 0.17143 | 4.6 | D=H | 0.076087 | D=H | 0.086957 | D=H | 0.097826 | D=H | 0.1087 | D=H | 0.11957 | D=H | 0.13043 | 4.3 | 0.069704 | 0.079662 | 0.089619 | 0.099577 | 0.10953 | 0.11949 | 4.3 | 4.0658 | 4.0658 | 4.0658 | 4.0658 | 4.0658 | 4.0658
| d=D/2.2 | 0.045455 | d=D/2.2 | 0.051948 | d=D/2.2 | 0.058442 | d=D/2.2 | 0.064935 | d=D/2.2 | 0.071429 | d=D/2.2 | 0.077922 |  | d=D/2.2 | 0.034585 | d=D/2.2 | 0.039526 | d=D/2.2 | 0.044466 | d=D/2.2 | 0.049407 | d=D/2.2 | 0.054348 | d=D/2.2 | 0.059289 | 4.6 | 0.065158 | 0.074466 | 0.083775 | 0.093083 | 0.10239 | 0.1117 | 4.6 | 4.3494 | 4.3494 | 4.3494 | 4.3494 | 4.3494 | 4.3494
| h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 |  | h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 | 5 | 0.059945 | 0.068509 | 0.077073 | 0.085636 | 0.0942 | 0.10276 | 5 | 4.7276 | 4.7276 | 4.7276 | 4.7276 | 4.7276 | 4.7276
| D=H=2.2d | 0.1 | D=H=2.2d | 0.11429 | D=H=2.2d | 0.12857 | D=H=2.2d | 0.14286 | D=H=2.2d | 0.15714 | D=H=2.2d | 0.17143 |  | D=H=2.2d | 0.076087 | D=H=2.2d | 0.086957 | D=H=2.2d | 0.097826 | D=H=2.2d | 0.1087 | D=H=2.2d | 0.11957 | D=H=2.2d | 0.13043
| ds=0.2d | 0.0090909 | ds=0.2d | 0.01039 | ds=0.2d | 0.011688 | ds=0.2d | 0.012987 | ds=0.2d | 0.014286 | ds=0.2d | 0.015584 |  | ds=0.2d | 0.006917 | ds=0.2d | 0.0079051 | ds=0.2d | 0.0088933 | ds=0.2d | 0.0098814 | ds=0.2d | 0.01087 | ds=0.2d | 0.011858 | L/D | Pitch (S) at 60 deg | [5 empty] | L/D | No. of spiral turns on the shaft (n) at 60 deg
| C=0.05d | 0.0022727 | C=0.05d | 0.0025974 | C=0.05d | 0.0029221 | C=0.05d | 0.0032468 | C=0.05d | 0.0035714 | C=0.05d | 0.0038961 |  | C=0.05d | 0.0017292 | C=0.05d | 0.0019763 | C=0.05d | 0.0022233 | C=0.05d | 0.0024704 | C=0.05d | 0.0027174 | C=0.05d | 0.0029644 |  | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6 |  | 0.35 | 0.4 | 0.45 | 0.5 | 0.55 | 0.6
| S=0.6d | 0.027273 | S=0.6d | 0.031169 | S=0.6d | 0.035065 | S=0.6d | 0.038961 | S=0.6d | 0.042857 | S=0.6d | 0.046753 |  | S=0.6d | 0.020751 | S=0.6d | 0.023715 | S=0.6d | 0.02668 | S=0.6d | 0.029644 | S=0.6d | 0.032609 | S=0.6d | 0.035573 | 1 | 0.51913 | 0.59329 | 0.66745 | 0.74161 | 0.81577 | 0.88993 | 1 | 0.54592 | 0.54592 | 0.54592 | 0.54592 | 0.54592 | 0.54592
| N (Turns) | 10.3914 | N (Turns) | 10.3914 | N (Turns) | 10.3914 | N (Turns) | 10.3914 | N (Turns) | 10.3914 | N (Turns) | 10.3914 |  | N (Turns) | 13.6572 | N (Turns) | 13.6572 | N (Turns) | 13.6572 | N (Turns) | 13.6572 | N (Turns) | 13.6572 | N (Turns) | 13.6572 | 1.2 | 0.43261 | 0.49441 | 0.55621 | 0.61801 | 0.67981 | 0.74161 | 1.2 | 0.6551 | 0.6551 | 0.6551 | 0.6551 | 0.6551 | 0.6551
| W=(d-ds)/2 | 0.018182 | W=(d-ds)/2 | 0.020779 | W=(d-ds)/2 | 0.023377 | W=(d-ds)/2 | 0.025974 | W=(d-ds)/2 | 0.028571 | W=(d-ds)/2 | 0.031169 |  | W=(d-ds)/2 | 0.013834 | W=(d-ds)/2 | 0.01581 | W=(d-ds)/2 | 0.017787 | W=(d-ds)/2 | 0.019763 | W=(d-ds)/2 | 0.021739 | W=(d-ds)/2 | 0.023715 | 1.5 | 0.34609 | 0.39553 | 0.44497 | 0.49441 | 0.54385 | 0.59329 | 1.5 | 0.81888 | 0.81888 | 0.81888 | 0.81888 | 0.81888 | 0.81888
| r | 0.05 | r | 0.057143 | r | 0.064286 | r | 0.071429 | r | 0.078571 | r | 0.085714 |  | r | 0.038043 | r | 0.043478 | r | 0.048913 | r | 0.054348 | r | 0.059783 | r | 0.065217 | 1.8 | 0.2884 | 0.3296 | 0.37081 | 0.41201 | 0.45321 | 0.49441 | 1.8 | 0.98265 | 0.98265 | 0.98265 | 0.98265 | 0.98265 | 0.98265
| V (m^3) | 0.0027475 | V (m^3) | 0.0041012 | V (m^3) | 0.0058394 | V (m^3) | 0.0080102 | V (m^3) | 0.010662 | V (m^3) | 0.013842 |  | V (m^3) | 0.0015906 | V (m^3) | 0.0023743 | V (m^3) | 0.0033806 | V (m^3) | 0.0046373 | V (m^3) | 0.0061722 | V (m^3) | 0.0080132 | 2 | 0.25956 | 0.29664 | 0.33372 | 0.37081 | 0.40789 | 0.44497 | 2 | 1.0918 | 1.0918 | 1.0918 | 1.0918 | 1.0918 | 1.0918
| V (m^3)70% | 0.0019232 | V (m^3)70% | 0.0028709 | V (m^3)70% | 0.0040876 | V (m^3)70% | 0.0056071 | V (m^3)70% | 0.0074631 | V (m^3)70% | 0.0096891 |  | V (m^3)70% | 0.0011134 | V (m^3)70% | 0.001662 | V (m^3)70% | 0.0023664 | V (m^3)70% | 0.0032461 | V (m^3)70% | 0.0043206 | V (m^3)70% | 0.0056093 | 2.5 | 0.20765 | 0.23732 | 0.26698 | 0.29664 | 0.32631 | 0.35597 | 2.5 | 1.3648 | 1.3648 | 1.3648 | 1.3648 | 1.3648 | 1.3648
| HL (mm) | 188 | HL (mm) | 216 | HL (mm) | 248 | HL (mm) | 276 | HL (mm) | 300 | HL (mm) | 328 |  | HL (mm) | 50 | HL (mm) | 58.4 | HL (mm) | 65.2 | HL (mm) | 71.6 | HL (mm) | 73.5 | HL (mm) | 91.6 | 3 | 0.17304 | 0.19776 | 0.22248 | 0.2472 | 0.27192 | 0.29664 | 3 | 1.6378 | 1.6378 | 1.6378 | 1.6378 | 1.6378 | 1.6378
[26 empty] | 3.5 | 0.14832 | 0.16951 | 0.1907 | 0.21189 | 0.23308 | 0.25427 | 3.5 | 1.9107 | 1.9107 | 1.9107 | 1.9107 | 1.9107 | 1.9107
4 | D=H | 0.0875 | D=H | 0.1 | D=H | 0.1125 | D=H | 0.125 | D=H | 0.1375 | D=H | 0.15 | 5 | D=H | 0.07 | D=H | 0.08 | D=H | 0.09 | D=H | 0.1 | D=H | 0.11 | D=H | 0.12 | 4 | 0.12978 | 0.14832 | 0.16686 | 0.1854 | 0.20394 | 0.22248 | 4 | 2.1837 | 2.1837 | 2.1837 | 2.1837 | 2.1837 | 2.1837
| d=D/2.2 | 0.039773 | d=D/2.2 | 0.045455 | d=D/2.2 | 0.051136 | d=D/2.2 | 0.056818 | d=D/2.2 | 0.0625 | d=D/2.2 | 0.068182 |  | d=D/2.2 | 0.031818 | d=D/2.2 | 0.036364 | d=D/2.2 | 0.040909 | d=D/2.2 | 0.045455 | d=D/2.2 | 0.05 | d=D/2.2 | 0.054545 | 4.3 | 0.12073 | 0.13797 | 0.15522 | 0.17247 | 0.18971 | 0.20696 | 4.3 | 2.3474 | 2.3474 | 2.3474 | 2.3474 | 2.3474 | 2.3474
| h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 |  | h=L/1.235 | 0.2834 | h=L/1.235 | 0.32389 | h=L/1.235 | 0.36437 | h=L/1.235 | 0.40486 | h=L/1.235 | 0.44534 | h=L/1.235 | 0.48583 | 4.6 | 0.11285 | 0.12898 | 0.1451 | 0.16122 | 0.17734 | 0.19346 | 4.6 | 2.5112 | 2.5112 | 2.5112 | 2.5112 | 2.5112 | 2.5112
| D=H=2.2d | 0.0875 | D=H=2.2d | 0.1 | D=H=2.2d | 0.1125 | D=H=2.2d | 0.125 | D=H=2.2d | 0.1375 | D=H=2.2d | 0.15 |  | D=H=2.2d | 0.07 | D=H=2.2d | 0.08 | D=H=2.2d | 0.09 | D=H=2.2d | 0.1 | D=H=2.2d | 0.11 | D=H=2.2d | 0.12 | 5 | 0.10383 | 0.11866 | 0.13349 | 0.14832 | 0.16315 | 0.17799 | 5 | 2.7296 | 2.7296 | 2.7296 | 2.7296 | 2.7296 | 2.7296
| ds=0.2d | 0.0079545 | ds=0.2d | 0.0090909 | ds=0.2d | 0.010227 | ds=0.2d | 0.011364 | ds=0.2d | 0.0125 | ds=0.2d | 0.013636 |  | ds=0.2d | 0.0063636 | ds=0.2d | 0.0072727 | ds=0.2d | 0.0081818 | ds=0.2d | 0.0090909 | ds=0.2d | 0.01 | ds=0.2d | 0.010909
| C=0.05d | 0.0019886 | C=0.05d | 0.0022727 | C=0.05d | 0.0025568 | C=0.05d | 0.0028409 | C=0.05d | 0.003125 | C=0.05d | 0.0034091 |  | C=0.05d | 0.0015909 | C=0.05d | 0.0018182 | C=0.05d | 0.0020455 | C=0.05d | 0.0022727 | C=0.05d | 0.0025 | C=0.05d | 0.0027273
| S=0.6d | 0.023864 | S=0.6d | 0.027273 | S=0.6d | 0.030682 | S=0.6d | 0.034091 | S=0.6d | 0.0375 | S=0.6d | 0.040909 |  | S=0.6d | 0.019091 | S=0.6d | 0.021818 | S=0.6d | 0.024545 | S=0.6d | 0.027273 | S=0.6d | 0.03 | S=0.6d | 0.032727
| N (Turns) | 11.8758 | N (Turns) | 11.8758 | N (Turns) | 11.8758 | N (Turns) | 11.8758 | N (Turns) | 11.8758 | N (Turns) | 11.8758 |  | N (Turns) | 14.8448 | N (Turns) | 14.8448 | N (Turns) | 14.8448 | N (Turns) | 14.8448 | N (Turns) | 14.8448 | N (Turns) | 14.8448
| W=(d-ds)/2 | 0.015909 | W=(d-ds)/2 | 0.018182 | W=(d-ds)/2 | 0.020455 | W=(d-ds)/2 | 0.022727 | W=(d-ds)/2 | 0.025 | W=(d-ds)/2 | 0.027273 |  | W=(d-ds)/2 | 0.012727 | W=(d-ds)/2 | 0.014545 | W=(d-ds)/2 | 0.016364 | W=(d-ds)/2 | 0.018182 | W=(d-ds)/2 | 0.02 | W=(d-ds)/2 | 0.021818
| r | 0.04375 | r | 0.05 | r | 0.05625 | r | 0.0625 | r | 0.06875 | r | 0.075 |  | r | 0.035 | r | 0.04 | r | 0.045 | r | 0.05 | r | 0.055 | r | 0.06
| V (m^3) | 0.0021036 | V (m^3) | 0.00314 | V (m^3) | 0.0044708 | V (m^3) | 0.0061328 | V (m^3) | 0.0081628 | V (m^3) | 0.010598 |  | V (m^3) | 0.0013463 | V (m^3) | 0.0020096 | V (m^3) | 0.0028613 | V (m^3) | 0.003925 | V (m^3) | 0.0052242 | V (m^3) | 0.0067824
| V (m^3)70% | 0.0014725 | V (m^3)70% | 0.002198 | V (m^3)70% | 0.0031296 | V (m^3)70% | 0.004293 | V (m^3)70% | 0.0057139 | V (m^3)70% | 0.0074183 |  | V (m^3)70% | 0.00094239 | V (m^3)70% | 0.0014067 | V (m^3)70% | 0.0020029 | V (m^3)70% | 0.0027475 | V (m^3)70% | 0.0036569 | V (m^3)70% | 0.0047477
| HL (mm) | 58.4 | HL (mm) | 66 | HL (mm) | 74 | HL (mm) | 82.5 | HL (mm) | 90 | HL (mm) | 98.7 |  | HL (mm) | 97.3 | HL (mm) | 111.7 | HL (mm) | 118.5 | HL (mm) | 132 | HL (mm) | 148.3 | HL (mm) | 158.3
Users' Requirement Specifications
Design Configuration
| Configuration | Aspect ratios | Measurements | Units
| Length of the vessel |  | 500 | mm
| Internal diameter of the vessel |  | 108.6957 | mm
| Impeller diameter |  | 50 | mm
| Impeller axial length |  | 404.8583 | mm
| The diameter of the shaft |  | 10 | mm
| Pitch of the impeller |  | 28 | mm
| Number of spiral turns |  | 14.4592 | mm
| Clearance |  | 2.5 | mm
| Thickness of the blade |  | 2 | mm
| Height of the Vane |  | 20 | mm
| Total volume / 70 % working volume |  | 4.6 L/3.2 L | L
| Blade angle | θ | 17 | [o]
(𝐿)
(𝐷)
(
(𝑑)
(𝑆)
〖(𝐿〗_𝑖)
(𝑁)
(𝐶)
)
(𝑊)
(𝑉)
〖(𝑑〗_𝑠)
〖(𝑇〗_𝑏)
(𝐿/4.6)
(𝐷/2.2)
(𝐿/1.235)
(0.2∗𝑑)
(𝐿_𝑖/𝑆)
(0.05∗𝑑)
(4 % 𝑜𝑓 𝑑)
(𝑑−𝑑_𝑠)/2
(𝐿)

## 2
Impeller immersion calculations
|  |  |  | Liquid volume (L) | Actual Volume (L) | Height of liquid (hL) (m) | Height of impeller from bottomh,i (m) | Height of impeller/impeller diameter (i/d) m
Parameters | Measurements | Units
Vessl internal diameter (ID) | 0.1087 | m |  | 1 | 1.111 | 0.03259 | 0.0032422 | 0.064843
Radius of the vessel - (R) | 0.054348 | m |  | 2 | 2.111 | 0.05295 | 0.023602 | 0.47204
Length of the vessel (L) | 0.5 | m |  | 3 | 3.111 | 0.07389 | 0.044542 | 0.89084
Total volume (V) | 4.6396 | L |  | 4 | 4.111 | 0.08977 | 0.060422 | 1.2084
Impeller diameter (d) | 0.05 | m
Impeller diameter (r) | 0.025 | m
R-r | 0.029348 | m
Volume of heating finger (HF) | [8 empty] | r |  |  | Heating finger
Diameter of the HF (d) | 0.022 | m | [6 empty] | R
Length of the HF (L) | 0.292 | m | [8 empty] | ID
Volume(V) | 0.000111 | m3 | [8 empty] | d
| 0.111 | L | [7 empty] | R-r
Volume (L) | m3
1 | 0.001 |  |  | a b
2 | 0.002 |  |  | Figure depicting the fluid covering in the HTB (a) volume filled less than radus of the reactor (b) more than the radious of the reactor |  |  | Reference:
3 | 0.003 | [5 empty] | http://www.ajdesigner.com/phphydraulicradius/hydraulic_radius_equation_pipe.php
4 | 0.004
From equation (a), when the working VL is less than 1/2 of the total volume
Let θ is the central angle, then
Let the area of the circular segmented area (Blue) |  | K
Working volume
Substituting
putting up the values
Solution for the above equation for h at each volume | VL | r | L | RHS | LHS | Variable (h)
| 0.001 | 0.054348 | 0.5 | 0.00099833 | 0.001 | 0.029091
| 0.002 | 0.054348 | 0.5 | 0.0019961 | 0.002 | 0.04838
| 0.003 | 0.054348 | 0.5 | 0.0029931 | 0.003 | 0.066848
| 0.004 | 0.05 | 0.5 | #NUM! | 0.004 | 0.1
This indicated that the height of liquid (h) is less than the clearance (C)
Let the impeller clearance is "C" |  | 0.029348 | m | h<C (liquid eght does not reach to the impeller)
Volume displaced by heating element
Volume displaced
0.000111
For 1 L
1 L | 0.001111 | m3
2 L | 0.002111 | m3
3 L | 0.003111 | m3
4 L | 0.004111 | m3
Liquid height for modified volume m3
| VL | r | L | RHS | LHS | Variable (h)
| 0.001111 | 0.054348 | 0.5 | 0.0011091 | 0.001111 | 0.031366
| 0.002111 | 0.054348 | 0.5 | 0.0021069 | 0.002111 | 0.050426
| 0.003111 | 0.054348 | 0.5 | 0.0031038 | 0.003111 | 0.068951
| 0.004111 | 0.054348 | 0.5 | 0.0040984 | 0.004111 | 0.089774
for 1L,h>C | 0.0020185 | m | 2.0185 | mm
| 0.021078 |  | 21.0782
| 0.039604 |  | 39.6036
| 0.060426 |  | 60.4261
in %
1 L | 0.04037 | 4.037%
2 L | 0.42156 | 42.16%
3 L | 0.79207 | 79.21%
4 L | 1.2085 | 120.9%
4 % (1 L)
42 % (2 L)
79 % (3 L)
Fully immersed (4 L)
< V
𝑽_(𝑳 )< 𝟏/𝟐 V
θ = 2arccos(r - h /r)
𝑽_(𝑳 )=𝐊∗𝐋
𝐾=𝑟^2/2(θ−𝑠𝑖𝑛θ)
𝑉_𝐿=𝑟^2/2 {2 cos^(−1)⁡((𝑟−ℎ)/𝑟)−sin⁡[2 cos^(−1)⁡((𝑟−ℎ)/𝑟) ] }𝐿
0.001=〖(0.0545)〗^2/2 {2 cos^(−1)⁡((0.0545−ℎ)/0.0545)−sin⁡[2 cos^(−1)⁡((0.0545−ℎ)/0.0545) ] }0.5
𝐶= (𝐷−𝑑)/2
𝑉_𝐻𝐸
𝑉_𝐻𝐸=𝜋 (〖0.022〗^2×0.292)/4=0.000111 𝑚^3
= Volume in the bioreactor (0.001) +
𝑉_𝐿 = Volume in the bioreactor (0.001) + 𝑉_𝐻𝐸
0.001111 =〖(0.0545)〗^2/2 {2 cos^(−1)⁡((0.0545−ℎ)/0.0545)−sin⁡[2 cos^(−1)⁡((0.0545−ℎ)/0.0545) ] }0.5
𝐼𝑚𝑝𝑒𝑙𝑙𝑒𝑟 𝑖𝑚𝑚𝑒𝑟𝑠𝑖𝑜𝑛 𝑟𝑎𝑡𝑖𝑜=ℎ_(𝐿,𝑑)/𝑑

## 3
Data for calculating mixing time through conductivity and F (t) curve
Sample calculations for determining mixing time through conductivity method and F(t) graphs | [14 empty] | Raw Conductivity |  |  | Co
Raw data for 1 L at 100 rpm (run 1) | [14 empty] | t (s) | Run 1 | Run 2 | Run 3 | average conductivity | conc | Int | E(t) | F(t) |  | 4 % (1 L) |  |  |  |  | 42 % (2 L) | [5 empty] | 79 % (3 L) | [5 empty] | Fully immersed (4 L)
--time-- | Cond | unit | TDS | unit | Tcon | unit | [8 empty] | 0 | 0 | 0 | 0 | 0 | 1.5 | 0 | 0.025623 | 0 | RPM | 100 | 200 | 300 | 400 | 500 | RPM | 100 | 200 | 300 | 400 | 500 | RPM | 100 | 200 | 300 | 400 | 500 | RPM | 100 | 200 | 300 | 400 | 500
16:15 | 20.3 | uS | 10.2 | ppm | 22.6 | C | [8 empty] | 1 | 0.1 | 2 | 0 | 0.7 | 1.5 | 1.5 | 0.025623 | 0.0015453 | Mixing time | 659 | 162 | 137 | 85 | 82 | Mixing time | 72 | 41 | 28 | 21 | 20 | Mixing time | 71 | 25 | 20 | 18 | 16 | Mixing time | 135 | 69 | 46 | 35 | 33
16:15 | 20.3 | uS | 10.2 | ppm | 22.6 | C | [8 empty] | 2 | 0 | -1 | 1 | 0 | 1.5 | 1.5 | 0.025623 | 0 | time (s) | F(t) | F(t) | F(t) | F(t) | F(t) | time (s) | F(t) | F(t) | F(t) | F(t) | F(t) | time (s) | F(t) | F(t) | F(t) | F(t) | F(t) | time (s) | F(t) | F(t) | F(t) | F(t) | F(t)
16:15 | 20.3 | uS | 10.2 | ppm | 22.6 | C | [8 empty] | 3 | 0 | 2 | 0 | 0.66667 | 1.5 | 1.5 | 0.025623 | 0.0014717 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024691 | 0.04065 | 0.033023 | 0.015873 | 0 | 0.015504 | 0.0081301 | 0.0083333 | 0.02439 | 0.019658 | 0 | 0 | 0 | 0 | 0.0083333 | 0.008547
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 4 | 0 | -1 | -1 | -0.66667 | 1.5 | 1.5 | 0.025623 | -0.0014717 | 1 | 0.0015453 | 0 | 0 | 0 | 0 | 1 | -0.0081301 | 0.024691 | 0.04065 | 0.033023 | 0.015873 | 1 | 0.015504 | 0.0081301 | 0.0083333 | 0.02439 | 0.019658 | 1 | 0 | 0 | 0 | 0.0083333 | 0.008547
16:15 | 20.3 | uS | 10.2 | ppm | 22.6 | C | [8 empty] | 5 | 0.1 | 2 | 0 | 0.7 | 1.5 | 1.5 | 0.025623 | 0.0015453 | 2 | 0 | 0 | 0 | 0 | -0.016878 | 2 | -0.0081301 | 0.024691 | 0.04065 | 0.033023 | 0.015873 | 2 | 0.015504 | 0.0081301 | 0.0083333 | 0.02439 | 0.019658 | 2 | 0 | 0 | 0 | 0.0083333 | 0.008547
16:15 | 20.3 | uS | 10.2 | ppm | 22.6 | C | [8 empty] | 6 | 0.1 | -1 | 0 | -0.3 | 1.5 | 1.5 | 0.025623 | -0.00066225 | 3 | 0.0014717 | 0 | 0 | 0 | -0.016878 | 3 | 0 | 0.024691 | 0.04065 | 0.033023 | 0.015873 | 3 | 0.015504 | 0.0081301 | 0.0083333 | 0.02439 | 0.019658 | 3 | 0 | 0 | 0.0083333 | 0.0083333 | 0
16:15 | 20.3 | uS | 10.2 | ppm | 22.6 | C | [8 empty] | 7 | 0 | 2 | 2 | 1.3333 | 1.5 | 1.5 | 0.025623 | 0.0029433 | 4 | -0.0014717 | 0 | 0 | 0 | -0.025316 | 4 | -0.0081301 | 0.024691 | 0.04065 | 0.033023 | 0.015873 | 4 | 0.015504 | 0.0081301 | 0.0083333 | 0.02439 | 0.019658 | 4 | 0 | 0 | 0.0083333 | 0.0083333 | 0.008547
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 8 | 0 | 0 | 0 | 0 | 1.5 | 1.5 | 0.025623 | 0 | 5 | 0.0015453 | 0 | 0 | 0 | 0 | 5 | -0.0081301 | 0.024691 | 0.04065 | 0.033023 | 0.015873 | 5 | 0.015504 | 0.0081301 | 0.0083333 | 0.02439 | 0.019658 | 5 | 0 | 0 | 0.0083333 | 0.016667 | 0.008547
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 9 | 0.1 | 2 | 0 | 0.7 | 1.5 | 1.5 | 0.025623 | 0.0015453 | 6 | -0.00066225 | 0.00086806 | 0 | 0 | 0 | 6 | -0.01626 | 0.024691 | 0.04065 | 0.033023 | 0.079365 | 6 | 0.015504 | 0.0081301 | 0.0083333 | 0.02439 | 0.019658 | 6 | 0 | 0 | 0.0083333 | 0.016667 | 0
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 10 | 0 | 0 | 0 | 0 | 1.5 | 1.5 | 0.025623 | 0 | 7 | 0.0029433 | 0.0017361 | 0 | 0 | -0.0084388 | 7 | -0.0081301 | 0.024691 | 0.04065 | 0.033023 | 0.19841 | 7 | 0.015504 | 0.0081301 | 0.0083333 | 0.02439 | 0.019658 | 7 | 0 | 0 | 0.016667 | 0.033333 | 0.008547
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 11 | 0 | 2 | 0 | 0.66667 | 1.5 | 1.5 | 0.025623 | 0.0014717 | 8 | 0 | 0.011285 | 0 | 0 | -0.0084388 | 8 | -0.0081301 | 0.024691 | 0.04065 | 0.20155 | 0.3254 | 8 | 0.015504 | 0.0081301 | 0.0083333 | 0.13008 | 0.30769 | 8 | 0 | 0 | 0.016667 | 0.05 | 0.034188
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 12 | 0.1 | 0 | 0 | 0.033333 | 1.5 | 1.5 | 0.025623 | 7.3584e-05 | 9 | 0.0015453 | 0.012153 | 0 | 0 | -0.0084388 | 9 | 0 | 0.024691 | 0.10569 | 0.31783 | 0.47619 | 9 | 0.015504 | 0.0081301 | 0.041667 | 0.30894 | 0.49573 | 9 | 0.01626 | 0 | 0.0083333 | 0.066667 | 0.059829
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 13 | 0 | 2 | 0 | 0.66667 | 1.5 | 1.5 | 0.025623 | 0.0014717 | 10 | 0 | 0.013889 | 0 | 0 | 0 | 10 | 0 | 0.041152 | 0.20325 | 0.42636 | 0.5873 | 10 | 0.015504 | 0.0081301 | 0.15833 | 0.50407 | 0.64103 | 10 | 0.0081301 | 0 | 0.016667 | 0.10833 | 0.18803
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 14 | 0 | 1 | 0 | 0.33333 | 1.5 | 1.5 | 0.025623 | 0.00073584 | 11 | 0.0014717 | 0.015625 | -0.0075758 | 0 | 0 | 11 | 0 | 0.049383 | 0.28455 | 0.53488 | 0.6746 | 11 | 0.015504 | 0.02439 | 0.325 | 0.63415 | 0.7265 | 11 | 0 | 0 | 0.041667 | 0.15 | 0.23932
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 15 | 0 | 2 | 1 | 1 | 1.5 | 1.5 | 0.025623 | 0.0022075 | 12 | 7.3584e-05 | 0.025174 | 0 | 0 | -0.0084388 | 12 | -0.0081301 | 0.074074 | 0.37398 | 0.6124 | 0.7381 | 12 | 0.015504 | 0.056911 | 0.49167 | 0.73171 | 0.81197 | 12 | 0.0081301 | 0 | 0.066667 | 0.20833 | 0.2735
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 16 | 0 | 1 | 1 | 0.66667 | 1.5 | 1.5 | 0.025623 | 0.0014717 | 13 | 0.0014717 | 0.017361 | 0 | 0 | -0.016878 | 13 | -0.0081301 | 0.11523 | 0.47154 | 0.68217 | 0.80952 | 13 | 0.015504 | 0.12195 | 0.64167 | 0.78862 | 0.8547 | 13 | 0.0081301 | 0 | 0.1 | 0.25833 | 0.41026
16:15 | 20.3 | uS | 10.2 | ppm | 22.7 | C |  | Mixing time (s) | (x - xmean)^2 |  |  | Legends |  |  | 17 | 0 | 2 | 1 | 1 | 1.5 | 1.5 | 0.025623 | 0.0022075 | 14 | 0.00073584 | 0.019965 | 0 | 0 | -0.0084388 | 14 | -0.02439 | 0.17284 | 0.55285 | 0.73643 | 0.84921 | 14 | 0.015504 | 0.21951 | 0.75 | 0.8374 | 0.90598 | 14 | 0 | 0.0081301 | 0.13333 | 0.325 | 0.46154
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | Run |  |  |  |  | 100 rpm |  |  | 18 | 0 | 2 | 1 | 1 | 1.5 | 1.5 | 0.025623 | 0.0022075 | 15 | 0.0022075 | 0.03125 | -0.0075758 | 0 | 0 | 15 | 0.0081301 | 0.23045 | 0.61789 | 0.7907 | 0.89683 | 15 | 0.015504 | 0.3252 | 0.83333 | 0.87805 | 0.92308 | 15 | 0 | 0.02439 | 0.175 | 0.38333 | 0.49573
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 1 | 575.5 | 6972.25 |  |  | 200 rpm |  |  | 19 | 0 | 2 | 1 | 1 | 1.5 | 1.5 | 0.025623 | 0.0022075 | 16 | 0.0014717 | 0.034722 | -0.0075758 | 0 | -0.0084388 | 16 | 0.01626 | 0.27984 | 0.69106 | 0.83721 | 0.92857 | 16 | 0.015504 | 0.44715 | 0.875 | 0.89431 | 0.94872 | 16 | 0.01626 | 0.04065 | 0.23333 | 0.43333 | 0.62393
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 2 | 561 | 9604 |  |  | 300 rpm |  |  | 20 | 0 | 2 | 1 | 1 | 1.5 | 1.5 | 0.025623 | 0.0022075 | 17 | 0.0022075 | 0.044271 | 0 | 0 | 0 | 17 | 0.02439 | 0.32099 | 0.74797 | 0.86047 | 0.95238 | 17 | 0.015504 | 0.55285 | 0.90833 | 0.92683 | 0.97436 | 17 | 0 | 0.04878 | 0.29167 | 0.48333 | 0.66667
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 3 | 840.5 | 32942.25 |  |  | 400 rpm |  |  | 21 | 0 | 2 | -2 | 0 | 1.5 | 1.5 | 0.025623 | 0 | 18 | 0.0022075 | 0.037326 | 0 | 0.0079365 | 0.0084388 | 18 | 0.03252 | 0.37037 | 0.78862 | 0.88372 | 0.96032 | 18 | 0.015504 | 0.65854 | 0.93333 | 0.93496 | 0.97436 | 18 | 0 | 0.065041 | 0.33333 | 0.54167 | 0.70085
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | Mean | 659 | 49518.5 |  |  | 500 rpm |  |  | 22 | 0 | 2 | -2 | 0 | 1.5 | 1.5 | 0.025623 | 0 | 19 | 0.0022075 | 0.052083 | -0.0075758 | 0.0079365 | 0 | 19 | 0.04878 | 0.42798 | 0.82927 | 0.91473 | 0.98413 | 19 | 0.015504 | 0.73171 | 0.95833 | 0.94309 | 0.99145 | 19 | 0 | 0.097561 | 0.39167 | 0.575 | 0.7265
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | N | 3 | [6 empty] | 23 | 0 | 3 | 0 | 1 | 1.5 | 1.5 | 0.025623 | 0.0022075 | 20 | 0.0022075 | 0.065104 | -0.0075758 | 0.015873 | 0.0084388 | 20 | 0.073171 | 0.4856 | 0.85366 | 0.92248 | 0.99206 | 20 | 0.015504 | 0.79675 | 0.96667 | 0.93496 | 0.99145 | 20 | 0.0081301 | 0.12195 | 0.43333 | 0.61667 | 0.83761
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | Standard deviation | 157.3507 | [6 empty] | 24 | 0 | 2 | 0 | 0.66667 | 1.5 | 1.5 | 0.025623 | 0.0014717 | 21 | 0 | 0.071181 | 0.0075758 | 0.02381 | 0.025316 | 21 | 0.089431 | 0.51029 | 0.87805 | 0.93798 | 1.0079 | 21 | 0.023256 | 0.85366 | 0.975 | 0.95122 | 1 | 21 | 0 | 0.15447 | 0.475 | 0.65833 | 0.86325
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | Standard error | 90.8465 | [6 empty] | 25 | 0 | 3 | 0 | 1 | 1 | 1.25 | 0.021352 | 0.0022075 | 22 | 0 | 0.082465 | 0.0075758 | 0.039683 | 0.025316 | 22 | 0.13008 | 0.55144 | 0.91057 | 0.95349 | 1.0079 | 22 | 0.03876 | 0.90244 | 0.98333 | 0.95935 | 1 | 22 | 0.0081301 | 0.19512 | 0.53333 | 0.68333 | 0.88034
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | U(95%) | 181.693 | [6 empty] | 26 | 0 | 2 | 0 | 0.66667 | 1.5 | 1.25 | 0.021352 | 0.0014717 | 23 | 0.0022075 | 0.095486 | 0.0075758 | 0.055556 | 0.042194 | 23 | 0.14634 | 0.59259 | 0.93496 | 0.95349 | 1.0159 | 23 | 0.046512 | 0.92683 | 0.98333 | 0.95935 | 1 | 23 | 0.0081301 | 0.21951 | 0.575 | 0.71667 | 0.89744
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | Mixing time | 659 ± 182 | [6 empty] | 27 | 0 | 3 | 0 | 1 | 1 | 1.25 | 0.021352 | 0.0022075 | 24 | 0.0014717 | 0.1059 | 0.015152 | 0.071429 | 0.067511 | 24 | 0.18699 | 0.65021 | 0.93496 | 0.97674 | 1.0079 | 24 | 0.069767 | 0.94309 | 0.99167 | 0.95935 | 1 | 24 | 0.0081301 | 0.25203 | 0.61667 | 0.74167 | 0.92308
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 28 | 0 | 3 | 0 | 1 | 1 | 1 | 0.017082 | 0.0022075 | 25 | 0.0022075 | 0.11198 | 0.015152 | 0.087302 | 0.084388 | 25 | 0.19512 | 0.68313 | 0.93496 | 0.96899 | 1.0238 | 25 | 0.077519 | 0.95935 | 1 | 0.96748 | 1 | 25 | 0.02439 | 0.30081 | 0.65833 | 0.775 | 0.92308
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 29 | 0.1 | 3 | 1 | 1.3667 | 0.73171 | 0.86585 | 0.01479 | 0.0030169 | 26 | 0.0014717 | 0.11892 | 0.022727 | 0.10317 | 0.1097 | 26 | 0.23577 | 0.70782 | 0.95935 | 0.97674 | 1.0238 | 26 | 0.10853 | 0.96748 | 1 | 0.95935 | 1 | 26 | 0.02439 | 0.3252 | 0.68333 | 0.79167 | 0.93162
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 30 | 0.1 | 3 | 1 | 1.3667 | 0.73171 | 0.73171 | 0.012499 | 0.0030169 | 27 | 0.0022075 | 0.13455 | 0.030303 | 0.12698 | 0.11814 | 27 | 0.25203 | 0.72428 | 0.96748 | 0.9845 | 1.0238 | 27 | 0.11628 | 0.98374 | 0.99167 | 0.97561 | 1 | 27 | 0.02439 | 0.35772 | 0.73333 | 0.825 | 0.94017
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 31 | 0.1 | 3 | 0 | 1.0333 | 0.96774 | 0.84972 | 0.014515 | 0.0022811 | 28 | 0.0022075 | 0.14931 | 0.030303 | 0.15873 | 0.14346 | 28 | 0.28455 | 0.7572 | 0.99187 | 0.9845 | 1.0317 | 28 | 0.14729 | 0.98374 | 1 | 0.95935 | 0.99145 | 28 | 0.03252 | 0.39024 | 0.75 | 0.825 | 0.94872
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 32 | 0.1 | 3 | 0 | 1.0333 | 0.96774 | 0.96774 | 0.016531 | 0.0022811 | 29 | 0.0030169 | 0.15625 | 0.037879 | 0.18254 | 0.16878 | 29 | 0.3252 | 0.79012 | 0.99187 | 0.9845 | 1.0397 | 29 | 0.18605 | 0.98374 | 1.0083 | 0.96748 | 1 | 29 | 0.03252 | 0.41463 | 0.76667 | 0.84167 | 0.95726
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 33 | 0.1 | 3 | 1 | 1.3667 | 0.73171 | 0.84972 | 0.014515 | 0.0030169 | 30 | 0.0030169 | 0.1684 | 0.060606 | 0.21429 | 0.20253 | 30 | 0.35772 | 0.80658 | 1 | 0.96899 | 1.0397 | 30 | 0.20155 | 0.98374 | 1.0083 | 0.96748 | 0.98291 | 30 | 0.03252 | 0.45528 | 0.80833 | 0.85833 | 0.95726
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 34 | 0 | 4 | 1 | 1.6667 | 0.6 | 0.66585 | 0.011374 | 0.0036792 | 31 | 0.0022811 | 0.18663 | 0.060606 | 0.24603 | 0.21941 | 31 | 0.38211 | 0.82305 | 1 | 0.97674 | 1.0397 | 31 | 0.23256 | 0.99187 | 1.0083 | 0.96748 | 0.98291 | 31 | 0.04878 | 0.47967 | 0.80833 | 0.875 | 0.96581
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 35 | 0.1 | 3 | -1 | 0.7 | 1.4286 | 1.0143 | 0.017326 | 0.0015453 | 32 | 0.0022811 | 0.19184 | 0.068182 | 0.26984 | 0.24473 | 32 | 0.39837 | 0.83951 | 1 | 0.97674 | 1.0476 | 32 | 0.26357 | 0.98374 | 1.0083 | 0.97561 | 0.98291 | 32 | 0.056911 | 0.5122 | 0.83333 | 0.89167 | 0.96581
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 36 | 0.1 | 4 | -1 | 1.0333 | 0.96774 | 1.1982 | 0.020467 | 0.0022811 | 33 | 0.0030169 | 0.2066 | 0.090909 | 0.29365 | 0.27004 | 33 | 0.41463 | 0.85597 | 1 | 0.97674 | 1.0397 | 33 | 0.29457 | 0.98374 | 1.0083 | 0.96748 | 0.99145 | 33 | 0.065041 | 0.54472 | 0.84167 | 0.9 | 0.96581
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 37 | 0.1 | 4 | 1 | 1.7 | 0.58824 | 0.77799 | 0.01329 | 0.0037528 | 34 | 0.0036792 | 0.22483 | 0.10606 | 0.31746 | 0.29536 | 34 | 0.45528 | 0.87243 | 1 | 0.96899 | 1.0476 | 34 | 0.31783 | 0.99187 | 1 | 0.97561 | 0.98291 | 34 | 0.081301 | 0.56911 | 0.86667 | 0.90833 | 0.98291
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 38 | 0.1 | 5 | 1 | 2.0333 | 0.4918 | 0.54002 | 0.0092246 | 0.0044886 | 35 | 0.0015453 | 0.2283 | 0.11364 | 0.35714 | 0.32068 | 35 | 0.47967 | 0.88066 | 1.0081 | 0.97674 | 1.0397 | 35 | 0.34884 | 0.99187 | 1 | 0.95935 | 0.98291 | 35 | 0.089431 | 0.5935 | 0.88333 | 0.90833 | 0.98291
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 39 | 0.1 | 5 | 0 | 1.7 | 0.58824 | 0.54002 | 0.0092246 | 0.0037528 | 36 | 0.0022811 | 0.23872 | 0.12879 | 0.37302 | 0.34599 | 36 | 0.52033 | 0.88889 | 1.0081 | 0.97674 | 1.0397 | 36 | 0.37984 | 0.99187 | 1.0167 | 0.95935 | 0.98291 | 36 | 0.10569 | 0.61789 | 0.88333 | 0.93333 | 0.99145
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 40 | 0.2 | 6 | 0 | 2.0667 | 0.48387 | 0.53605 | 0.0091568 | 0.0045622 | 37 | 0.0037528 | 0.25521 | 0.13636 | 0.39683 | 0.37131 | 37 | 0.54472 | 0.89712 | 1.0244 | 0.9845 | 1.0397 | 37 | 0.42636 | 0.98374 | 1.0083 | 0.95935 | 0.98291 | 37 | 0.12195 | 0.62602 | 0.90833 | 0.93333 | 0.99145
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 41 | 0.2 | 6 | 1 | 2.4 | 0.41667 | 0.45027 | 0.0076914 | 0.005298 | 38 | 0.0044886 | 0.2717 | 0.15152 | 0.42857 | 0.38819 | 38 | 0.56911 | 0.91358 | 1.0244 | 0.99225 | 1.0397 | 38 | 0.45736 | 1 | 1 | 0.95935 | 0.97436 | 38 | 0.13821 | 0.65041 | 0.90833 | 0.94167 | 0.98291
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 42 | 0.2 | 8 | 1 | 3.0667 | 0.32609 | 0.37138 | 0.0063438 | 0.0067697 | 39 | 0.0037528 | 0.2908 | 0.16667 | 0.44444 | 0.42194 | 39 | 0.5935 | 0.93827 | 1.0163 | 1 | 1.0317 | 39 | 0.49612 | 0.99187 | 1.0083 | 0.95935 | 0.99145 | 39 | 0.1626 | 0.66667 | 0.93333 | 0.93333 | 0.99145
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 43 | 0.3 | 8 | 1 | 3.1 | 0.32258 | 0.32433 | 0.0055402 | 0.0068433 | 40 | 0.0045622 | 0.29601 | 0.18182 | 0.47619 | 0.44726 | 40 | 0.60976 | 0.93827 | 1.0081 | 0.97674 | 1.0317 | 40 | 0.52713 | 0.99187 | 1.0083 | 0.96748 | 0.97436 | 40 | 0.1626 | 0.69106 | 0.94167 | 0.94167 | 0.99145
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 44 | 0.4 | 10 | 1 | 3.8 | 0.26316 | 0.29287 | 0.0050028 | 0.0083885 | 41 | 0.005298 | 0.32378 | 0.18939 | 0.50794 | 0.47257 | 41 | 0.61789 | 0.93827 | 1.0081 | 0.97674 | 1.0317 | 41 | 0.57364 | 1 | 1.0083 | 0.95935 | 0.99145 | 41 | 0.17886 | 0.70732 | 0.95833 | 0.93333 | 0.99145
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 45 | 0.5 | 11 | 0 | 3.8333 | 0.26087 | 0.26201 | 0.0044757 | 0.0084621 | 42 | 0.0067697 | 0.33247 | 0.21212 | 0.53175 | 0.49789 | 42 | 0.65854 | 0.93827 | 1.0163 | 0.97674 | 1.0317 | 42 | 0.5969 | 1 | 1.0083 | 0.98374 | 0.99145 | 42 | 0.19512 | 0.73171 | 0.95833 | 0.94167 | 0.99145
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 46 | 0.5 | 12 | 0 | 4.1667 | 0.24 | 0.25043 | 0.0042779 | 0.0091979 | 43 | 0.0068433 | 0.33941 | 0.24242 | 0.54762 | 0.51477 | 43 | 0.6748 | 0.9465 | 1.0081 | 0.97674 | 1.0317 | 43 | 0.63566 | 0.99187 | 0.99167 | 0.97561 | 0.98291 | 43 | 0.21138 | 0.73984 | 0.96667 | 0.93333 | 0.99145
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 47 | 0.6 | 15 | 1 | 5.5333 | 0.18072 | 0.21036 | 0.0035934 | 0.012215 | 44 | 0.0083885 | 0.36024 | 0.25 | 0.56349 | 0.54008 | 44 | 0.69919 | 0.96296 | 1.0244 | 0.97674 | 1.0317 | 44 | 0.65891 | 0.98374 | 0.99167 | 0.97561 | 0.99145 | 44 | 0.21951 | 0.7561 | 0.96667 | 0.94167 | 0.99145
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 48 | 0.9 | 18 | 1 | 6.6333 | 0.15075 | 0.16574 | 0.0028311 | 0.014643 | 45 | 0.0084621 | 0.37153 | 0.25758 | 0.57937 | 0.54852 | 45 | 0.69106 | 0.97942 | 1.0244 | 0.97674 | 1.0317 | 45 | 0.66667 | 1 | 1 | 0.98374 | 0.98291 | 45 | 0.2439 | 0.76423 | 0.975 | 0.95 | 0.99145
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | [8 empty] | 49 | 1 | 19 | 1 | 7 | 0.14286 | 0.14681 | 0.0025077 | 0.015453 | 46 | 0.0091979 | 0.3776 | 0.2803 | 0.61111 | 0.57384 | 46 | 0.71545 | 0.97119 | 1.0244 | 0.97674 | 1.0317 | 46 | 0.68992 | 0.98374 | 1.0083 | 0.97561 | 0.98291 | 46 | 0.25203 | 0.78049 | 0.98333 | 0.95 | 0.99145
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | Injection | init cond | 20.3 | uS |  |  |  |  | 50 | 1.1 | 21 | 1 | 7.7 | 0.12987 | 0.13636 | 0.0023294 | 0.016998 | 47 | 0.012215 | 0.38976 | 0.30303 | 0.61905 | 0.59072 | 47 | 0.73984 | 0.97119 | 1.0244 | 0.97674 | 1.0397 | 47 | 0.71318 | 0.98374 | 1.0083 | 0.96748 | 0.98291 | 47 | 0.27642 | 0.79675 | 0.99167 | 0.95 | 0.99145
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | t (s) | Conductivity | Conc | Conc normalized | Conductivity diff | Unit | Homogeneity |  | 51 | 1.3 | 22 | 3 | 8.7667 | 0.11407 | 0.12197 | 0.0020835 | 0.019352 | 48 | 0.014643 | 0.39583 | 0.31818 | 0.64286 | 0.61603 | 48 | 0.7561 | 0.96296 | 1.0163 | 0.96899 | 1.0397 | 48 | 0.72868 | 0.98374 | 1 | 0.96748 | 0.98291 | 48 | 0.29268 | 0.80488 | 0.99167 | 0.95833 | 0.99145
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 0 | 20.3 | 0.049261 | 0 | 0 | µS | ct(0) | 20.3 | 52 | 1.5 | 21 | 3 | 8.5 | 0.11765 | 0.11586 | 0.0019791 | 0.018764 | 49 | 0.015453 | 0.40885 | 0.34091 | 0.65873 | 0.62447 | 49 | 0.76423 | 0.97942 | 1.0244 | 0.9845 | 1.0317 | 49 | 0.75194 | 0.98374 | 1.0083 | 0.96341 | 0.98291 | 49 | 0.29268 | 0.81301 | 1 | 0.95833 | 0.99145
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 1 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS | ct(inf) | 458 | 53 | 1.6 | 22 | 0 | 7.8667 | 0.12712 | 0.12238 | 0.0020905 | 0.017366 | 50 | 0.016998 | 0.42361 | 0.35606 | 0.66667 | 0.64979 | 50 | 0.76423 | 0.97119 | 1.0163 | 0.9845 | 1.0317 | 50 | 0.77519 | 0.99187 | 1 | 0.96341 | 0.98291 | 50 | 0.31707 | 0.82927 | 1 | 0.95 | 0.99145
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 2 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 54 | 1.9 | 23 | 0 | 8.3 | 0.12048 | 0.1238 | 0.0021147 | 0.018322 | 51 | 0.019352 | 0.44184 | 0.37121 | 0.68254 | 0.66667 | 51 | 0.78862 | 0.97942 | 1.0163 | 0.99225 | 1.0397 | 51 | 0.78295 | 1 | 1 | 0.96341 | 0.99145 | 51 | 0.33333 | 0.8374 | 1.0083 | 0.95833 | 0.99145
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 3 | 20.3 | 0.049261 | 0 | 0 | µS | c(t) | 436.115 | 55 | 2.1 | 25 | 2 | 9.7 | 0.10309 | 0.11179 | 0.0019095 | 0.021413 | 52 | 0.018764 | 0.46007 | 0.38636 | 0.69841 | 0.67511 | 52 | 0.78862 | 0.97942 | 1.0163 | 0.9845 | 1.0397 | 52 | 0.79845 | 0.99187 | 1.0083 | 0.96341 | 0.99145 | 52 | 0.34146 | 0.85366 | 1 | 0.95833 | 0.99145
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 4 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 56 | 2.4 | 28 | 2 | 10.8 | 0.092593 | 0.097843 | 0.0016713 | 0.023841 | 53 | 0.017366 | 0.46354 | 0.40152 | 0.72222 | 0.69198 | 53 | 0.81301 | 0.98765 | 1.0244 | 0.99225 | 1.0317 | 53 | 0.81395 | 0.98374 | 1.0083 | 0.97561 | 0.99145 | 53 | 0.35772 | 0.85366 | 1.0083 | 0.95833 | 0.99145
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 5 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 57 | 2.8 | 29 | 2 | 11.2667 | 0.088757 | 0.090675 | 0.0015489 | 0.024871 | 54 | 0.018322 | 0.47483 | 0.42424 | 0.73016 | 0.69198 | 54 | 0.81301 | 1.0041 | 1.0244 | 0.99225 | 1.0238 | 54 | 0.82171 | 0.98374 | 1 | 0.97561 | 0.99145 | 54 | 0.37398 | 0.85366 | 1.0083 | 0.96667 | 1.0269
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 6 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 58 | 3.3 | 32 | 2 | 12.4333 | 0.080429 | 0.084593 | 0.001445 | 0.027447 | 55 | 0.021413 | 0.49045 | 0.43182 | 0.74603 | 0.70886 | 55 | 0.82927 | 0.99588 | 1.0163 | 0.9845 | 1.0397 | 55 | 0.84496 | 0.99187 | 1.0083 | 0.96341 | 1 | 55 | 0.38211 | 0.86179 | 1.0083 | 0.96667 | 1.0355
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 7 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 59 | 3.5 | 33 | 2 | 12.8333 | 0.077922 | 0.079176 | 0.0013525 | 0.02833 | 56 | 0.023841 | 0.4974 | 0.44697 | 0.7619 | 0.72574 | 56 | 0.8374 | 1.0123 | 1.0244 | 0.9845 | 1.0397 | 56 | 0.84496 | 0.99187 | 1.0083 | 0.97561 | 0.99145 | 56 | 0.39837 | 0.86992 | 1.0167 | 0.95833 | 1.0355
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 8 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 60 | 3.7 | 37 | 2 | 14.2333 | 0.070258 | 0.07409 | 0.0012656 | 0.03142 | 57 | 0.024871 | 0.51476 | 0.4697 | 0.76984 | 0.74262 | 57 | 0.85366 | 1.0123 | 1.0163 | 0.97674 | 1.0476 | 57 | 0.86047 | 0.99187 | 1.0083 | 0.96341 | 1 | 57 | 0.42276 | 0.87805 | 1.0167 | 0.95833 | 1.0355
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 9 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 61 | 4.2 | 38 | 3 | 15.0667 | 0.066372 | 0.068315 | 0.0011669 | 0.03326 | 58 | 0.027447 | 0.52778 | 0.47727 | 0.77778 | 0.75105 | 58 | 0.85366 | 1.0123 | 1.0163 | 0.97674 | 1.0397 | 58 | 0.86822 | 0.99187 | 1.0083 | 0.96341 | 0.99145 | 58 | 0.42276 | 0.88618 | 1.0167 | 0.95833 | 1
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 10 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 62 | 4.4 | 41 | 3 | 16.1333 | 0.061983 | 0.064178 | 0.0010963 | 0.035614 | 59 | 0.02833 | 0.53733 | 0.49242 | 0.77778 | 0.75949 | 59 | 0.86179 | 1.0041 | 1.0244 | 0.97674 | 1.0476 | 59 | 0.87597 | 0.99187 | 1 | 0.96341 | 1 | 59 | 0.44715 | 0.89431 | 1.025 | 0.95 | 1.0269
16:16 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 11 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 63 | 4.9 | 40 | 3 | 15.9667 | 0.06263 | 0.062307 | 0.0010643 | 0.035247 | 60 | 0.03142 | 0.55122 | 0.5 | 0.77778 | 0.78481 | 60 | 0.86179 | 1.0041 | 1.0244 | 0.96899 | 1.0397 | 60 | 0.89147 | 0.98374 | 0.99167 | 0.97561 | 0.99145 | 60 | 0.46341 | 0.90244 | 1.025 | 0.96667 | 1.0355
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 12 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 64 | 5.7 | 45 | 3 | 17.9 | 0.055866 | 0.059248 | 0.0010121 | 0.039514 | 61 | 0.03326 | 0.55729 | 0.50758 | 0.79365 | 0.78481 | 61 | 0.87805 | 1.0041 | 1.0163 | 0.97674 | 1.0556 | 61 | 0.89147 | 0.99187 | 0.99167 | 0.97561 | 0.99145 | 61 | 0.47154 | 0.90244 | 1.0333 | 0.95833 | 1
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 13 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 65 | 6.6 | 46 | 4 | 18.8667 | 0.053004 | 0.054435 | 0.00092985 | 0.041648 | 62 | 0.035614 | 0.56771 | 0.5303 | 0.80159 | 0.81013 | 62 | 0.88618 | 1.0206 | 1.0244 | 0.97674 | 1.0397 | 62 | 0.89147 | 0.98374 | 1.0083 | 0.9878 | 1.0128 | 62 | 0.47967 | 0.91057 | 1.025 | 0.95 | 1
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 14 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 66 | 7 | 51 | 4 | 20.6667 | 0.048387 | 0.050695 | 0.00086597 | 0.045622 | 63 | 0.035247 | 0.5816 | 0.53788 | 0.80952 | 0.81013 | 63 | 0.88618 | 1.0041 | 1.0163 | 0.96899 | 1.0556 | 63 | 0.89922 | 0.98374 | 1.0083 | 0.97561 | 1 | 63 | 0.49593 | 0.91057 | 1.0333 | 0.96667 | 1.044
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 15 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 67 | 7.2 | 50 | 1 | 19.4 | 0.051546 | 0.049967 | 0.00085353 | 0.042826 | 64 | 0.039514 | 0.58941 | 0.53788 | 0.8254 | 0.81013 | 64 | 0.89431 | 0.99588 | 1.0163 | 0.97674 | 1.0397 | 64 | 0.90698 | 0.98374 | 1.0083 | 0.97561 | 1.0128 | 64 | 0.49593 | 0.9187 | 1.0333 | 0.96667 | 1.0355
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 16 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 68 | 8 | 52 | 1 | 20.3333 | 0.04918 | 0.050363 | 0.0008603 | 0.044886 | 65 | 0.041648 | 0.59375 | 0.55303 | 0.84127 | 0.81857 | 65 | 0.89431 | 0.99588 | 1.0244 | 0.97674 | 1.0556 | 65 | 0.90698 | 0.98374 | 1.0083 | 0.9878 | 1 | 65 | 0.52033 | 0.9187 | 1.0167 | 0.95833 | 1.044
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 17 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 69 | 8.5 | 51 | 3 | 20.8333 | 0.048 | 0.04859 | 0.00083001 | 0.04599 | 66 | 0.045622 | 0.60851 | 0.56818 | 0.84127 | 0.83544 | 66 | 0.91057 | 1.0041 | 1.0081 | 0.97674 | 1.0476 | 66 | 0.92248 | 1.0081 | 1 | 0.9878 | 1 | 66 | 0.52846 | 0.9187 | 1.025 | 0.975 | 1.044
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 18 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 70 | 8.9 | 53 | 3 | 21.6333 | 0.046225 | 0.047112 | 0.00080477 | 0.047756 | 67 | 0.042826 | 0.58681 | 0.58333 | 0.84921 | 0.84388 | 67 | 0.9187 | 1.0123 | 1.0244 | 0.97674 | 1.0476 | 67 | 0.92248 | 1 | 1.0083 | 0.9878 | 1 | 67 | 0.53659 | 0.9187 | 1.0167 | 0.975 | 1.0355
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 19 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 71 | 13.2 | 53 | 3 | 23.0667 | 0.043353 | 0.044789 | 0.00076508 | 0.05092 | 68 | 0.044886 | 0.62413 | 0.59091 | 0.84921 | 0.86076 | 68 | 0.92683 | 1.0041 | 1.0325 | 0.9845 | 1.0397 | 68 | 0.92248 | 0.98374 | 1.0083 | 0.96341 | 1 | 68 | 0.55285 | 0.92683 | 1.025 | 0.95 | 1
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 20 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 72 | 13.9 | 56 | 3 | 24.3 | 0.041152 | 0.042252 | 0.00072175 | 0.053642 | 69 | 0.04599 | 0.61198 | 0.58333 | 0.85714 | 0.86076 | 69 | 0.92683 | 0.99588 | 1.0244 | 0.97674 | 1.0595 | 69 | 0.93023 | 0.99187 | 1.0083 | 0.96341 | 1 | 69 | 0.56098 | 0.93496 | 1.025 | 0.975 | 1.044
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 21 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 73 | 14.2 | 57 | 4 | 25.0667 | 0.039894 | 0.040523 | 0.00069221 | 0.055335 | 70 | 0.047756 | 0.63021 | 0.60606 | 0.84921 | 0.8692 | 70 | 0.92683 | 0.99588 | 1.0163 | 0.97674 | 1.0357 | 70 | 0.93023 | 1 | 1 | 0.96341 | 1 | 70 | 0.56911 | 0.93496 | 1.0333 | 0.975 | 1.0355
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 22 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 74 | 15 | 59 | 4 | 26 | 0.038462 | 0.039178 | 0.00066923 | 0.057395 | 71 | 0.05092 | 0.63628 | 0.62879 | 0.85714 | 0.8692 | 71 | 0.9187 | 0.99588 | 1.0244 | 0.96899 | 1.0357 | 71 | 0.93798 | 1 | 1.0083 | 0.96341 | 1 | 71 | 0.58537 | 0.94309 | 1.0333 | 0.95833 | 0.95855
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 23 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 75 | 20 | 59 | 6 | 28.3333 | 0.035294 | 0.036878 | 0.00062994 | 0.062546 | 72 | 0.053642 | 0.6467 | 0.63636 | 0.85714 | 0.8692 | 72 | 0.92683 | 1.0123 | 1.0244 | 0.9845 | 1.0476 | 72 | 0.94574 | 0.99187 | 1.0083 | 0.96341 | 0.97436 | 72 | 0.5935 | 0.95122 | 1.025 | 0.975 | 1.0355
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 24 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 76 | 20.3 | 60 | 6 | 28.7667 | 0.034762 | 0.035028 | 0.00059835 | 0.063503 | 73 | 0.055335 | 0.65017 | 0.65152 | 0.88095 | 0.87764 | 73 | 0.92683 | 1.0041 | 1.0163 | 0.9845 | 1.0476 | 73 | 0.94574 | 0.98374 | 1.0083 | 0.96341 | 1 | 73 | 0.60976 | 0.94309 | 1.0333 | 0.96667 | 1.0355
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 25 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 77 | 20.5 | 59 | 6 | 28.5 | 0.035088 | 0.034925 | 0.00059659 | 0.062914 | 74 | 0.057395 | 0.66927 | 0.65152 | 0.88095 | 0.87764 | 74 | 0.94309 | 1.0041 | 1.0244 | 0.9845 | 1.0476 | 74 | 0.94574 | 0.98374 | 1.0083 | 0.96341 | 0.97436 | 74 | 0.61789 | 0.93496 | 1.0333 | 0.95833 | 1.0355
16:16 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 26 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 78 | 24.9 | 62 | 6 | 30.9667 | 0.032293 | 0.03369 | 0.00057549 | 0.068359 | 75 | 0.062546 | 0.67795 | 0.66667 | 0.88095 | 0.90295 | 75 | 0.94309 | 0.99588 | 1.0325 | 0.9845 | 1.0476 | 75 | 0.95349 | 1 | 1 | 0.96341 | 1 | 75 | 0.62602 | 0.94309 | 1.0333 | 0.975 | 1.0355
16:17 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 27 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 79 | 24.6 | 64 | 6 | 31.5333 | 0.031712 | 0.032003 | 0.00054667 | 0.06961 | 76 | 0.063503 | 0.68142 | 0.66667 | 0.88889 | 0.89451 | 76 | 0.94309 | 1.0041 | 1.0244 | 0.99225 | 1.0476 | 76 | 0.96124 | 0.99187 | 1 | 0.97561 | 0.97436 | 76 | 0.63415 | 0.95122 | 1.0333 | 0.96667 | 1.0526
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 28 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 80 | 24.3 | 67 | 6 | 32.4333 | 0.030832 | 0.031272 | 0.00053419 | 0.071597 | 77 | 0.062914 | 0.69271 | 0.68182 | 0.88889 | 0.91139 | 77 | 0.95122 | 1.0041 | 1.0244 | 0.9845 | 1.0476 | 77 | 0.95349 | 0.99187 | 1.0083 | 0.97561 | 1 | 77 | 0.64228 | 0.95122 | 1.025 | 0.95833 | 1.044
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 29 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 81 | 24.3 | 68 | 5 | 32.4333 | 0.030832 | 0.030832 | 0.00052668 | 0.071597 | 78 | 0.068359 | 0.69531 | 0.68939 | 0.90476 | 0.91983 | 78 | 0.95122 | 1.0041 | 1.0244 | 0.97674 | 1.0476 | 78 | 0.96124 | 0.99187 | 1.0083 | 0.97561 | 0.97436 | 78 | 0.65854 | 0.95122 | 1.0333 | 0.975 | 1.044
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 30 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 82 | 28.9 | 70 | 5 | 34.6333 | 0.028874 | 0.029853 | 0.00050995 | 0.076453 | 79 | 0.06961 | 0.70747 | 0.68939 | 0.90476 | 0.91139 | 79 | 0.95122 | 0.99588 | 1.0163 | 0.9845 | 1.0476 | 79 | 0.95349 | 1.0081 | 1 | 0.96341 | 1 | 79 | 0.66667 | 0.95122 | 1.0333 | 0.96667 | 1.044
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 31 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 83 | 29.5 | 72 | 8 | 36.5 | 0.027397 | 0.028136 | 0.00048061 | 0.080574 | 80 | 0.071597 | 0.70833 | 0.70455 | 0.90476 | 0.91983 | 80 | 0.95935 | 1.0041 | 1.0163 | 0.9845 | 1.0476 | 80 | 0.96124 | 1.0081 | 1.0083 | 0.96341 | 1 | 80 | 0.66667 | 0.95935 | 1.0333 | 0.95833 | 1.0355
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 32 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 84 | 30.4 | 74 | 8 | 37.4667 | 0.02669 | 0.027044 | 0.00046196 | 0.082708 | 81 | 0.071597 | 0.72135 | 0.71212 | 0.92063 | 0.91139 | 81 | 0.95935 | 0.99588 | 1.0244 | 0.9845 | 1.0595 | 81 | 0.96124 | 1 | 1.0083 | 0.96341 | 1 | 81 | 0.6748 | 0.95935 | 1.0333 | 0.95833 | 0.95855
16:17 | 20.3 | uS | 10.2 | ppm | 22.7 | C | 33 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 85 | 34.5 | 75 | 9 | 39.5 | 0.025316 | 0.026003 | 0.00044419 | 0.087196 | 82 | 0.076453 | 0.73264 | 0.71212 | 0.92063 | 0.92827 | 82 | 0.95935 | 1.0123 | 1.0244 | 0.96899 | 1.0476 | 82 | 0.96124 | 0.98374 | 1.0083 | 0.96341 | 1 | 82 | 0.68293 | 0.95935 | 1.025 | 0.975 | 1.044
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 34 | 20.3 | 0.049261 | 0 | 0 | µS |  |  | 86 | 30 | 76 | 9 | 38.3333 | 0.026087 | 0.025702 | 0.00043903 | 0.084621 | 83 | 0.080574 | 0.74306 | 0.71212 | 0.92063 | 0.91983 | 83 | 0.95935 | 1.0041 | 1.0244 | 0.99225 | 1.0595 | 83 | 0.96124 | 0.99187 | 1.0083 | 0.97561 | 1 | 83 | 0.69106 | 0.95935 | 1.0417 | 0.98333 | 1.0355
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 35 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 87 | 30.4 | 78 | 12 | 40.1333 | 0.024917 | 0.025502 | 0.00043562 | 0.088595 | 84 | 0.082708 | 0.74479 | 0.71212 | 0.92857 | 0.91983 | 84 | 0.96748 | 1.0041 | 1.0163 | 0.9845 | 1.0476 | 84 | 0.97674 | 1 | 1 | 0.9878 | 1 | 84 | 0.69919 | 0.96748 | 1.0333 | 0.96667 | 1.044
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 36 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 88 | 31 | 80 | 12 | 41 | 0.02439 | 0.024654 | 0.00042113 | 0.090508 | 85 | 0.087196 | 0.75434 | 0.72727 | 0.92063 | 0.93671 | 85 | 0.96748 | 1.0041 | 1.0244 | 0.99225 | 1.0595 | 85 | 0.96899 | 0.99187 | 1 | 0.97561 |  | 85 | 0.71545 | 0.95935 | 1.0417 | 0.975 | 0.95855
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 37 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 89 | 35.3 | 83 | 12 | 43.4333 | 0.023024 | 0.023707 | 0.00040496 | 0.095879 | 86 | 0.084621 | 0.76649 | 0.73485 | 0.92857 | 0.92827 | 86 | 0.96748 | 1.0041 | 1.0163 | 0.99225 | 1.0476 | 86 | 0.96899 | 0.99187 | 0.99167 | 0.97561 |  | 86 | 0.71545 | 0.96748 | 1.0417 | 0.98333 | 1.044
16:17 | 20.4 | uS | 10.2 | ppm | 22.7 | C | 38 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 90 | 38.7 | 84 | 12 | 44.9 | 0.022272 | 0.022648 | 0.00038687 | 0.099117 | 87 | 0.088595 | 0.76823 | 0.73485 | 0.92857 | 0.94515 | 87 | 0.96748 | 1.0123 | 1.0244 | 0.97674 | 1.0476 | 87 | 0.97674 | 0.99187 | 1.0083 | 0.97561 |  | 87 | 0.71545 | 0.96748 | 1.0333 | 0.96667 | 1.0526
16:17 | 20.5 | uS | 10.3 | ppm | 22.7 | C | 39 | 20.4 | 0.04902 | 0.00024148 | 0.1 | µS |  |  | 91 | 38.2 | 84 | 11 | 44.4 | 0.022523 | 0.022397 | 0.00038259 | 0.098013 | 88 | 0.090508 | 0.77951 | 0.75 | 0.92857 | 0.93671 | 88 | 0.97561 | 1.0123 | 1.0081 | 0.9845 | 1.0476 | 88 | 0.97674 | 1 | 1.0083 | 0.97561 |  | 88 | 0.73984 | 0.96748 | 1.0333 | 0.98333 | 1.0355
16:17 | 20.5 | uS | 10.3 | ppm | 22.7 | C | 40 | 20.5 | 0.04878 | 0.0004806 | 0.2 | µS |  |  | 92 | 38.2 | 83 | 11 | 44.0667 | 0.022693 | 0.022608 | 0.00038618 | 0.097277 | 89 | 0.095879 | 0.78212 | 0.75758 | 0.93651 | 0.95359 | 89 | 0.96748 | 1.0041 | 1.0244 | 0.9845 | 1.0595 | 89 | 0.96899 | 0.98374 | 1.0083 | 0.97561 |  | 89 | 0.73984 | 0.96748 | 1.0333 | 0.98333 | 1.044
16:17 | 20.5 | uS | 10.3 | ppm | 22.7 | C | 41 | 20.5 | 0.04878 | 0.0004806 | 0.2 | µS |  |  | 93 | 41.7 | 84 | 14 | 46.5667 | 0.021475 | 0.022084 | 0.00037723 | 0.1028 | 90 | 0.099117 | 0.79948 | 0.75758 | 0.95238 | 0.93671 | 90 | 0.96748 | 1.0041 | 1.0244 | 0.97674 | 1.0595 | 90 | 0.97674 | 0.98374 | 1.0083 | 0.96341 |  | 90 | 0.74797 | 0.96748 | 1.0417 | 0.975 | 1.044
16:17 | 20.6 | uS | 10.3 | ppm | 22.7 | C | 42 | 20.5 | 0.04878 | 0.0004806 | 0.2 | µS |  |  | 94 | 41.3 | 85 | 14 | 46.7667 | 0.021383 | 0.021429 | 0.00036604 | 0.10324 | 91 | 0.098013 | 0.80295 | 0.75758 | 0.95238 | 0.94515 | 91 | 0.96748 | 1.0123 | 1.0163 | 0.9845 | 1.0595 | 91 | 0.97674 | 1 | 1 | 0.96341 |  | 91 | 0.76423 | 0.97561 | 1.0417 | 0.98333 | 1.044
16:17 | 20.7 | uS | 10.4 | ppm | 22.7 | C | 43 | 20.6 | 0.048544 | 0.00071739 | 0.3 | µS |  |  | 95 | 47.3 | 88 | 15 | 50.1 | 0.01996 | 0.020671 | 0.00035311 | 0.1106 | 92 | 0.097277 | 0.80382 | 0.75758 | 0.95238 | 0.94515 | 92 | 0.97561 | 1.0123 | 1.0163 | 0.97674 | 1.0476 | 92 | 0.96899 | 1 | 1 | 0.96341 |  | 92 | 0.76423 | 0.96748 | 1.0333 | 0.98333 | 1.0526
16:17 | 20.8 | uS | 10.4 | ppm | 22.7 | C | 44 | 20.7 | 0.048309 | 0.0009519 | 0.4 | µS |  |  | 96 | 46.8 | 90 | 15 | 50.6 | 0.019763 | 0.019861 | 0.00033927 | 0.1117 | 93 | 0.1028 | 0.8151 | 0.76515 | 0.96032 | 0.95359 | 93 | 0.97561 | 1.0041 | 1.0163 | 0.96124 | 1.0476 | 93 | 0.97674 | 1 | 1.0083 | 0.96341 |  | 93 | 0.77236 | 0.96748 | 1.025 | 0.96667 | 1.0355
16:17 | 20.8 | uS | 10.4 | ppm | 22.7 | C | 45 | 20.8 | 0.048077 | 0.0011842 | 0.5 | µS |  |  | 97 | 46.3 | 92 | 17 | 51.7667 | 0.019317 | 0.01954 | 0.00033378 | 0.11428 | 94 | 0.10324 | 0.81684 | 0.76515 | 0.96032 | 0.95359 | 94 | 0.97561 | 1.0041 | 1.0244 | 0.97674 | 1.0595 | 94 | 0.97674 | 0.99187 | 1.0083 | 0.96341 |  | 94 | 0.78049 | 0.96748 | 1.0333 | 0.98333 | 1.044
16:17 | 20.9 | uS | 10.5 | ppm | 22.7 | C | 46 | 20.8 | 0.048077 | 0.0011842 | 0.5 | µS |  |  | 98 | 50 | 93 | 17 | 53.3333 | 0.01875 | 0.019034 | 0.00032513 | 0.11773 | 95 | 0.1106 | 0.80816 | 0.77273 | 0.95238 | 0.95359 | 95 | 0.97561 | 1.0041 | 1.0163 | 0.9845 | 1.0476 | 95 | 0.96899 | 1 | 1 | 0.96341 |  | 95 | 0.78862 | 0.97561 | 1.025 | 0.98333 | 1.044
16:17 | 21.2 | uS | 10.6 | ppm | 22.7 | C | 47 | 20.9 | 0.047847 | 0.0014142 | 0.6 | µS |  |  | 99 | 50 | 93 | 16 | 53 | 0.018868 | 0.018809 | 0.00032129 | 0.117 | 96 | 0.1117 | 0.82986 | 0.78788 | 0.96032 | 0.95359 | 96 | 0.97561 | 1.0041 | 1.0163 | 0.97674 | 1.0595 | 96 | 0.97674 | 0.99187 | 1.0083 | 0.96341 |  | 96 | 0.78862 | 0.96748 | 1.0417 | 0.98333 | 1.044
16:17 | 21.3 | uS | 10.7 | ppm | 22.7 | C | 48 | 21.2 | 0.04717 | 0.0020913 | 0.9 | µS |  |  | 100 | 45.7 | 96 | 16 | 52.5667 | 0.019023 | 0.018946 | 0.00032363 | 0.11604 | 97 | 0.11428 | 0.83854 | 0.78788 | 0.95238 | 0.94515 | 97 | 0.97561 | 1.0041 | 1.0163 | 0.9845 | 1.0476 | 97 | 0.97674 | 0.98374 | 1.0083 | 0.96341 |  | 97 | 0.80488 | 0.97561 | 1.0417 | 0.975 | 1.0526
16:17 | 21.4 | uS | 10.7 | ppm | 22.7 | C | 49 | 21.3 | 0.046948 | 0.0023127 | 1 | µS |  |  | 101 | 45.6 | 97 | 18 | 53.5333 | 0.01868 | 0.018852 | 0.00032202 | 0.11818 | 98 | 0.11773 | 0.77734 | 0.7803 | 0.94444 | 0.95359 | 98 | 0.96748 | 0.99588 | 1.0244 | 0.99225 | 1.0595 | 98 | 0.96899 | 0.9878 | 1.0083 | 0.96341 |  | 98 | 0.80488 | 0.98374 | 1.0417 | 0.98333 | 1.044
16:17 | 21.6 | uS | 10.8 | ppm | 22.7 | C | 50 | 21.4 | 0.046729 | 0.0025321 | 1.1 | µS |  |  | 102 | 51.7 | 97 | 18 | 55.5667 | 0.017996 | 0.018338 | 0.00031325 | 0.12266 | 99 | 0.117 | 0.77865 | 0.78788 | 0.94444 | 0.95359 | 99 | 0.97561 | 0.99588 | 1.0163 | 0.99225 | 1.0476 | 99 | 0.96899 | 1 | 1.0125 | 0.97561 |  | 99 | 0.80488 | 0.97561 | 1.0333 | 0.98333 | 1.044
16:17 | 21.8 | uS | 10.9 | ppm | 22.7 | C | 51 | 21.6 | 0.046296 | 0.0029648 | 1.3 | µS |  |  | 103 | 52.1 | 99 | 20 | 57.0333 | 0.017534 | 0.017765 | 0.00030346 | 0.1259 | 100 | 0.11604 | 0.77995 | 0.79545 | 0.95238 | 0.96203 | 100 | 0.99187 | 1.0041 | 1.0163 | 0.99225 | 1.0476 | 100 | 0.97674 | 1 | 0.9875 | 0.97561 |  | 100 | 0.81301 | 0.96748 | 1.0417 | 0.975 | 1.0526
16:17 | 21.9 | uS | 11 | ppm | 22.7 | C | 52 | 21.8 | 0.045872 | 0.0033895 | 1.5 | µS |  |  | 104 | 51.9 | 98 | 20 | 56.6333 | 0.017657 | 0.017596 | 0.00030057 | 0.12502 | 101 | 0.11818 | 0.78385 | 0.80303 | 0.94444 | 0.96203 | 101 | 0.98374 | 1.0041 | 1.0163 | 0.99225 | 1.0476 | 101 | 0.96899 | 1 | 0.9875 | 0.97561 |  | 101 | 0.82114 | 0.97561 | 1.0417 | 0.98333 | 1.044
16:17 | 22.2 | uS | 11.1 | ppm | 22.7 | C | 53 | 21.9 | 0.045662 | 0.003599 | 1.6 | µS |  |  | 105 | 52.4 | 99 | 21 | 57.4667 | 0.017401 | 0.017529 | 0.00029944 | 0.12686 | 102 | 0.12266 | 0.79948 | 0.79545 | 0.94444 | 0.96203 | 102 | 0.97561 | 1.0041 | 1.0244 | 0.9845 | 1.0357 | 102 | 0.96899 | 1 | 1.0125 | 0.97561 |  | 102 | 0.81301 | 0.99187 | 1.0417 | 0.975 | 1.044
16:17 | 22.4 | uS | 11.2 | ppm | 22.7 | C | 54 | 22.2 | 0.045045 | 0.004216 | 1.9 | µS |  |  | 106 | 52.8 | 98 | 21 | 57.2667 | 0.017462 | 0.017432 | 0.00029777 | 0.12642 | 103 | 0.1259 | 0.80208 | 0.80303 | 0.95238 | 0.96203 | 103 | 0.97561 | 0.99588 | 1.0244 | 0.9845 | 1.0357 | 103 | 0.97674 | 1 | 1 | 0.97561 |  | 103 | 0.82927 | 0.97561 | 1.05 | 0.98333 | 1.044
16:17 | 22.7 | uS | 11.4 | ppm | 22.7 | C | 55 | 22.4 | 0.044643 | 0.0046182 | 2.1 | µS |  |  | 107 | 53.2 | 103 | 23 | 59.7333 | 0.016741 | 0.017102 | 0.00029213 | 0.13186 | 104 | 0.12502 | 0.81771 | 0.81061 | 0.94444 | 0.97046 | 104 | 0.98374 | 0.99588 | 1.0081 | 0.97674 | 1.0357 | 104 | 0.96899 | 0.9878 | 0.9875 | 0.97561 |  | 104 | 0.8374 | 0.96748 | 1.05 | 0.98333 | 1
16:17 | 23.1 | uS | 11.6 | ppm | 22.7 | C | 56 | 22.7 | 0.044053 | 0.0052082 | 2.4 | µS |  |  | 108 | 53.2 | 101 | 23 | 59.0667 | 0.01693 | 0.016836 | 0.00028758 | 0.13039 | 105 | 0.12686 | 0.81771 | 0.81061 | 0.94444 | 0.97046 | 105 | 0.98374 | 1.0041 | 1.0244 | 0.97674 | 1.0238 | 105 | 0.96899 | 1 | 1.0125 | 1 |  | 105 | 0.8374 | 0.97561 | 1.0417 | 0.975 | 1.0128
16:17 | 23.6 | uS | 11.8 | ppm | 22.7 | C | 57 | 23.1 | 0.04329 | 0.005971 | 2.8 | µS |  |  | 109 | 53.9 | 105 | 25 | 61.3 | 0.016313 | 0.016622 | 0.00028393 | 0.13532 | 106 | 0.12642 | 0.82161 | 0.81061 | 0.96032 | 0.96203 | 106 | 0.98374 | 0.99588 | 1.0244 | 0.9845 | 1.0357 | 106 | 0.97674 | 1 | 1.0125 | 1 |  | 106 | 0.8374 | 0.98374 | 1.0375 | 0.98333 | 0.98718
16:17 | 23.8 | uS | 11.9 | ppm | 22.7 | C | 58 | 23.6 | 0.042373 | 0.0068882 | 3.3 | µS |  |  | 110 | 58.1 | 105 | 25 | 62.7 | 0.015949 | 0.016131 | 0.00027555 | 0.13841 | 107 | 0.13186 | 0.82161 | 0.81061 | 0.96032 | 0.97046 | 107 | 0.99187 | 1.0041 | 1.0163 | 0.97674 | 1.0357 | 107 | 0.96899 | 0.9878 | 1 | 1 |  | 107 | 0.84553 | 0.98374 | 1.0375 | 0.98333 | 1
16:17 | 24 | uS | 12 | ppm | 22.7 | C | 59 | 23.8 | 0.042017 | 0.0072443 | 3.5 | µS |  |  | 111 | 57.7 | 109 | 29 | 65.2333 | 0.01533 | 0.015639 | 0.00026715 | 0.144 | 108 | 0.13039 | 0.82552 | 0.81061 | 0.96032 | 0.9789 | 108 | 0.99187 | 1.0206 | 1.0081 | 0.97674 | 1.0238 | 108 | 0.96899 | 0.9878 | 1 | 1 |  | 108 | 0.85366 | 0.97561 | 1.0375 | 0.975 | 1
16:17 | 24.5 | uS | 12.3 | ppm | 22.7 | C | 60 | 24 | 0.041667 | 0.0075944 | 3.7 | µS |  |  | 112 | 58.3 | 109 | 29 | 65.4333 | 0.015283 | 0.015306 | 0.00026146 | 0.14444 | 109 | 0.13532 | 0.82812 | 0.81818 | 0.96032 | 0.97046 | 109 | 0.97561 | 1.0041 | 1.0163 | 0.9845 | 1.0238 | 109 | 0.97674 | 0.9878 | 1.025 | 1 |  | 109 | 0.85366 | 0.98374 | 1.025 | 0.975 | 1
16:17 | 24.7 | uS | 12.4 | ppm | 22.7 | C | 61 | 24.5 | 0.040816 | 0.0084448 | 4.2 | µS |  |  | 113 | 63.6 | 112 | 28 | 67.8667 | 0.014735 | 0.015009 | 0.00025638 | 0.14982 | 110 | 0.13841 | 0.82812 | 0.83333 | 0.96032 | 0.97046 | 110 | 0.97561 | 1.0123 | 1.0244 | 0.9845 | 1.0238 | 110 | 0.97674 | 0.9878 | 1 | 1 |  | 110 | 0.85366 | 0.98374 | 1.025 | 0.98333 | 1
16:17 | 25.2 | uS | 12.6 | ppm | 22.7 | C | 62 | 24.7 | 0.040486 | 0.0087753 | 4.4 | µS |  |  | 114 | 63.8 | 112 | 28 | 67.9333 | 0.01472 | 0.014728 | 0.00025157 | 0.14996 | 111 | 0.144 | 0.84635 | 0.81818 | 0.95238 | 0.97046 | 111 | 0.97561 | 1.0041 | 1.0244 | 0.9845 | 1.0238 | 111 | 0.97674 | 0.9878 | 1 | 1 |  | 111 | 0.86179 | 0.98374 | 1.025 | 0.975 | 1
16:17 | 26 | uS | 13 | ppm | 22.7 | C | 63 | 25.2 | 0.039683 | 0.0095785 | 4.9 | µS |  |  | 115 | 63.6 | 115 | 29 | 69.2 | 0.014451 | 0.014586 | 0.00024915 | 0.15276 | 112 | 0.14444 | 0.84635 | 0.81818 | 0.95238 | 0.9789 | 112 | 0.97561 | 1.0123 | 1.0244 | 0.9845 | 1.0238 | 112 | 0.99225 | 0.9878 | 1 |  |  | 112 | 0.86179 | 0.98374 | 1.0375 | 0.98333 | 1
16:17 | 26.9 | uS | 13.5 | ppm | 22.7 | C | 64 | 26 | 0.038462 | 0.0108 | 5.7 | µS |  |  | 116 | 67.3 | 116 | 29 | 70.7667 | 0.014131 | 0.014291 | 0.00024412 | 0.15622 | 113 | 0.14982 | 0.84766 | 0.81818 | 0.95238 | 0.97046 | 113 | 0.98374 | 1.0041 | 1.0081 | 0.9845 | 1.0357 | 113 | 0.97674 | 1.0122 |  |  |  | 113 | 0.86992 | 0.98374 | 1.0375 | 0.98333 | 1
16:17 | 27.3 | uS | 13.7 | ppm | 22.7 | C | 65 | 26.9 | 0.037175 | 0.012086 | 6.6 | µS |  |  | 117 | 67.4 | 118 | 29 | 71.4667 | 0.013993 | 0.014062 | 0.0002402 | 0.15776 | 114 | 0.14996 | 0.85026 | 0.82576 | 0.94444 | 0.97468 | 114 | 0.99187 | 0.99588 | 1.0163 | 0.9845 | 1.0357 | 114 | 0.97674 | 1.0122 |  |  |  | 114 | 0.86992 | 0.98374 | 1.025 | 0.975 | 1
16:17 | 27.5 | uS | 13.8 | ppm | 22.7 | C | 66 | 27.3 | 0.03663 | 0.012631 | 7 | µS |  |  | 118 | 71.9 | 120 | 29 | 73.6333 | 0.013581 | 0.013787 | 0.0002355 | 0.16255 | 115 | 0.15276 | 0.85156 | 0.83333 | 0.96032 | 0.98734 | 115 | 1 | 0.99588 | 1.0163 | 0.96899 | 1.0238 | 115 | 0.97674 | 0.9878 |  |  |  | 115 | 0.86992 | 0.97561 | 1.025 | 0.98333 | 1
16:17 | 28.3 | uS | 14.2 | ppm | 22.7 | C | 67 | 27.5 | 0.036364 | 0.012897 | 7.2 | µS |  |  | 119 | 71.8 | 122 | 30 | 74.6 | 0.013405 | 0.013493 | 0.00023048 | 0.16468 | 116 | 0.15622 | 0.86849 | 0.83333 | 0.96825 | 0.97468 | 116 | 0.99187 | 1.0041 | 1.0081 | 0.97674 | 1.0238 | 116 | 0.97674 | 0.9878 |  |  |  | 116 | 0.86992 | 0.97561 | 1 | 0.975 | 1
16:17 | 28.8 | uS | 14.4 | ppm | 22.7 | C | 68 | 28.3 | 0.035336 | 0.013925 | 8 | µS |  |  | 120 | 71.2 | 124 | 30 | 75.0667 | 0.013321 | 0.013363 | 0.00022827 | 0.16571 | 117 | 0.15776 | 0.86849 | 0.82576 | 0.96032 | 0.98734 | 117 | 0.99187 | 1.0041 | 1.0244 | 0.97674 | 1.0357 | 117 | 0.97674 | 1 |  |  |  | 117 | 0.87805 | 0.98374 |  | 0.975 | 1
16:17 | 29.2 | uS | 14.6 | ppm | 22.7 | C | 69 | 28.8 | 0.034722 | 0.014539 | 8.5 | µS |  |  | 121 | 70.7 | 128 | 33 | 77.2333 | 0.012948 | 0.013135 | 0.00022436 | 0.17049 | 118 | 0.16255 | 0.87109 | 0.84091 | 0.96825 | 0.97468 | 118 | 0.98374 | 1.0041 | 1.0163 | 0.9845 | 1.0476 | 118 | 0.97674 | 0.9878 |  |  |  | 118 | 0.88618 | 0.98374 |  | 0.98333 | 1
16:17 | 33.5 | uS | 16.8 | ppm | 22.7 | C | 70 | 29.2 | 0.034247 | 0.015015 | 8.9 | µS |  |  | 122 | 70.7 | 129 | 33 | 77.5667 | 0.012892 | 0.01292 | 0.0002207 | 0.17123 | 119 | 0.16468 | 0.88542 | 0.84848 | 0.96825 | 0.98734 | 119 | 0.98374 | 0.99588 | 1.0244 | 0.9845 | 1.0238 | 119 | 0.9845 | 0.9878 |  |  |  | 119 | 0.88618 | 0.99187 |  | 0.98333 | 1
16:17 | 34.2 | uS | 17.1 | ppm | 22.7 | C | 71 | 33.5 | 0.029851 | 0.01941 | 13.2 | µS |  |  | 123 | 74.8 | 131 | 34 | 79.9333 | 0.01251 | 0.012701 | 0.00021696 | 0.17645 | 120 | 0.16571 | 0.88672 | 0.84091 | 0.97619 | 0.97468 | 120 | 0.99187 | 1.0041 | 1.0163 | 0.9845 | 1.0476 | 120 | 0.97674 | 0.9878 |  |  |  | 120 | 0.88618 | 0.98374 |  | 0.975 | 1
16:17 | 34.5 | uS | 17.3 | ppm | 22.7 | C | 72 | 34.2 | 0.02924 | 0.020021 | 13.9 | µS |  |  | 124 | 78.5 | 133 | 34 | 81.8333 | 0.01222 | 0.012365 | 0.00021122 | 0.18065 | 121 | 0.17049 | 0.89062 | 0.84848 | 0.96825 | 0.98734 | 121 | 0.99187 | 1.0041 | 1.0163 | 0.9845 | 1.0357 | 121 | 0.96899 | 1 |  |  |  | 121 | 0.89431 | 0.98374 |  | 0.98333
16:17 | 35.3 | uS | 17.7 | ppm | 22.7 | C | 73 | 34.5 | 0.028986 | 0.020276 | 14.2 | µS |  |  | 125 | 74.4 | 135 | 35 | 81.4667 | 0.012275 | 0.012247 | 0.00020921 | 0.17984 | 122 | 0.17123 | 0.89062 | 0.84848 | 0.97619 | 1 | 122 | 0.99187 | 1.0041 | 1.0081 | 0.9845 | 1.0357 | 122 | 0.97674 | 1 |  |  |  | 122 | 0.90244 | 0.98374 |  | 0.98333
16:17 | 40.3 | uS | 20.2 | ppm | 22.7 | C | 74 | 35.3 | 0.028329 | 0.020932 | 15 | µS |  |  | 126 | 78.5 | 138 | 35 | 83.8333 | 0.011928 | 0.012102 | 0.00020672 | 0.18506 | 123 | 0.17645 | 0.89323 | 0.84848 | 0.98413 | 0.97468 | 123 | 0.98374 | 1.0123 | 1.0244 | 0.9845 | 1.0238 | 123 | 0.97674 | 1 |  |  |  | 123 | 0.91057 | 0.98374 |  | 1
16:17 | 40.6 | uS | 20.3 | ppm | 22.7 | C | 75 | 40.3 | 0.024814 | 0.024447 | 20 | µS |  |  | 127 | 86.2 | 140 | 33 | 86.4 | 0.011574 | 0.011751 | 0.00020073 | 0.19073 | 124 | 0.18065 | 0.89583 | 0.84848 | 0.97619 | 0.97468 | 124 | 0.98374 | 1.0041 | 1.0244 | 0.9845 | 1.0238 | 124 | 0.97674 | 1 |  |  |  | 124 | 0.91057 | 0.97561 |  | 1
16:17 | 40.8 | uS | 20.4 | ppm | 22.7 | C | 76 | 40.6 | 0.024631 | 0.024631 | 20.3 | µS |  |  | 128 | 85.9 | 143 | 33 | 87.3 | 0.011455 | 0.011514 | 0.00019669 | 0.19272 | 125 | 0.17984 | 0.91016 | 0.84848 | 0.97619 | 0.98734 | 125 | 0.99187 | 0.99588 | 1.0325 | 0.9845 | 1.0238 | 125 | 0.97674 | 1 |  |  |  | 125 | 0.91057 | 0.98374 |  | 1.0125
16:17 | 45.2 | uS | 22.6 | ppm | 22.7 | C | 77 | 40.8 | 0.02451 | 0.024751 | 20.5 | µS |  |  | 129 | 85 | 143 | 35 | 87.6667 | 0.011407 | 0.011431 | 0.00019526 | 0.19352 | 126 | 0.18506 | 0.90104 | 0.84848 | 0.97619 | 1 | 126 | 0.98374 | 0.99588 | 1.0244 | 0.97674 | 1.0238 | 126 | 0.99225 | 1 |  |  |  | 126 | 0.91057 | 0.98374 |  | 1
16:17 | 44.9 | uS | 22.5 | ppm | 22.7 | C | 78 | 45.2 | 0.022124 | 0.027137 | 24.9 | µS |  |  | 130 | 85.1 | 145 | 35 | 88.3667 | 0.011316 | 0.011362 | 0.00019408 | 0.19507 | 127 | 0.19073 | 0.89974 | 0.84848 | 0.97619 | 0.98734 | 127 | 0.98374 | 0.99588 | 1.0163 | 0.99225 | 1.0238 | 127 | 0.97674 | 1 |  |  |  | 127 | 0.91057 | 0.98374 |  | 1
16:17 | 44.6 | uS | 22.3 | ppm | 22.7 | C | 79 | 44.9 | 0.022272 | 0.026989 | 24.6 | µS |  |  | 131 | 91 | 144 | 35 | 90 | 0.011111 | 0.011214 | 0.00019155 | 0.19868 | 128 | 0.19272 | 0.90234 | 0.85606 | 0.96825 | 1 | 128 | 0.98374 | 1.0041 | 1.0163 | 0.9845 | 1.0357 | 128 | 0.9845 | 1 |  |  |  | 128 | 0.91057 | 0.98374 |  | 1
16:17 | 44.6 | uS | 22.3 | ppm | 22.7 | C | 80 | 44.6 | 0.022422 | 0.02684 | 24.3 | µS |  |  | 132 | 91.4 | 144 | 35 | 90.1333 | 0.011095 | 0.011103 | 0.00018966 | 0.19897 | 129 | 0.19352 | 0.91667 | 0.85606 | 0.96032 | 1 | 129 | 0.99187 | 0.99588 | 1.0163 | 0.9845 | 1.0357 | 129 | 0.9845 | 0.9878 |  |  |  | 129 | 0.91057 | 0.97561 |  | 1
16:17 | 49.2 | uS | 24.6 | ppm | 22.7 | C | 81 | 44.6 | 0.022422 | 0.02684 | 24.3 | µS |  |  | 133 | 87.5 | 144 | 37 | 89.5 | 0.011173 | 0.011134 | 0.00019019 | 0.19757 | 130 | 0.19507 | 0.91667 | 0.85606 | 0.96032 | 0.98734 | 130 | 0.99187 | 1.0041 | 1.0244 | 0.97674 | 1.0357 | 130 | 0.9845 | 0.9878 |  |  |  | 130 | 0.91057 | 0.99187 |  | 0.975
16:17 | 49.8 | uS | 24.9 | ppm | 22.7 | C | 82 | 49.2 | 0.020325 | 0.028936 | 28.9 | µS |  |  | 134 | 92.2 | 145 | 37 | 91.4 | 0.010941 | 0.011057 | 0.00018888 | 0.20177 | 131 | 0.19868 | 0.92057 | 0.86364 | 0.96032 | 0.98734 | 131 | 0.98374 | 0.99588 | 1.0325 | 0.97674 | 1.0357 | 131 | 0.9845 | 0.9878 |  |  |  | 131 | 0.91057 | 0.99187 |  | 1
16:17 | 50.7 | uS | 25.4 | ppm | 22.7 | C | 83 | 49.8 | 0.02008 | 0.029181 | 29.5 | µS |  |  | 135 | 92.4 | 146 | 38 | 92.1333 | 0.010854 | 0.010897 | 0.00018615 | 0.20338 | 132 | 0.19897 | 0.93229 | 0.87121 | 0.96825 | 0.97468 | 132 | 0.98374 | 0.99588 | 1.0244 | 0.9845 | 1.0357 | 132 | 0.97674 | 1.0122 |  |  |  | 132 | 0.92683 | 0.99187 |  | 1
16:17 | 54.8 | uS | 27.4 | ppm | 22.7 | C | 84 | 50.7 | 0.019724 | 0.029537 | 30.4 | µS |  |  | 136 | 92.3 | 148 | 38 | 92.7667 | 0.01078 | 0.010817 | 0.00018477 | 0.20478 | 133 | 0.19757 | 0.93359 | 0.85606 | 0.96032 | 1 | 133 | 0.99187 | 0.99588 | 1.0244 | 0.9845 | 1.0595 | 133 | 0.9845 | 0.9878 |  |  |  | 133 | 0.92683 | 0.98374 |  | 1
16:17 | 50.3 | uS | 25.2 | ppm | 22.7 | C | 85 | 54.8 | 0.018248 | 0.031013 | 34.5 | µS |  |  | 137 | 96.9 | 149 | 38 | 94.6333 | 0.010567 | 0.010673 | 0.00018232 | 0.2089 | 134 | 0.20177 | 0.9349 | 0.86364 | 0.96032 |  | 134 | 0.99187 | 0.99588 | 1.0244 | 0.97674 | 1.0476 | 134 | 0.9845 | 0.9878 |  |  |  | 134 | 0.91057 | 0.99187 |  | 1
16:18 | 50.7 | uS | 25.4 | ppm | 22.7 | C | 86 | 50.3 | 0.019881 | 0.02938 | 30 | µS |  |  | 138 | 101 | 148 | 38 | 95.6667 | 0.010453 | 0.01051 | 0.00017953 | 0.21118 | 135 | 0.20338 | 0.9349 | 0.87121 | 0.96825 |  | 135 | 0.99187 | 1.0041 | 1.0081 | 0.9845 | 1.0595 | 135 | 0.9845 | 1 |  |  |  | 135 | 0.92683 | 0.99187 |  | 1
16:18 | 51.3 | uS | 25.7 | ppm | 22.7 | C | 87 | 50.7 | 0.019724 | 0.029537 | 30.4 | µS |  |  | 139 | 101.2 | 148 | 40 | 96.4 | 0.010373 | 0.010413 | 0.00017788 | 0.2128 | 136 | 0.20478 | 0.9388 | 0.87121 | 0.96825 |  | 136 | 0.98374 | 1.0041 | 1.0244 | 0.96899 | 1.0595 | 136 | 0.9845 | 1 |  |  |  | 136 | 0.93496 | 0.98374 |  | 1
16:18 | 55.6 | uS | 27.8 | ppm | 22.7 | C | 88 | 51.3 | 0.019493 | 0.029768 | 31 | µS |  |  | 140 | 101.8 | 148 | 40 | 96.6 | 0.010352 | 0.010363 | 0.00017701 | 0.21325 | 137 | 0.2089 | 0.95052 | 0.87121 | 0.96825 |  | 137 | 0.99187 | 1.0041 | 1.0366 | 0.97674 | 1.0476 | 137 | 0.9845 | 0.9878 |  |  |  | 137 | 0.93496 | 0.99187 |  | 1
16:18 | 59 | uS | 29.5 | ppm | 22.7 | C | 89 | 55.6 | 0.017986 | 0.031275 | 35.3 | µS |  |  | 141 | 98 | 147 | 39 | 94.6667 | 0.010563 | 0.010458 | 0.00017864 | 0.20898 | 138 | 0.21118 | 0.95312 | 0.87121 | 0.97619 |  | 138 | 0.99187 | 1.0041 | 1.0122 | 0.96512 | 1.0357 | 138 | 0.9845 | 0.9878 |  |  |  | 138 | 0.94309 | 0.98374 |  | 1
16:18 | 58.5 | uS | 29.3 | ppm | 22.7 | C | 90 | 59 | 0.016949 | 0.032312 | 38.7 | µS |  |  | 142 | 103.1 | 150 | 39 | 97.3667 | 0.01027 | 0.010417 | 0.00017794 | 0.21494 | 139 | 0.2128 | 0.94141 | 0.86364 | 0.97619 |  | 139 | 1.0244 | 1 | 1.0244 | 0.97674 | 1.0476 | 139 | 0.9845 | 1 |  |  |  | 139 | 0.94309 | 0.99187 |  | 1
16:18 | 58.5 | uS | 29.3 | ppm | 22.7 | C | 91 | 58.5 | 0.017094 | 0.032167 | 38.2 | µS |  |  | 143 | 108.7 | 149 | 42 | 99.9 | 0.01001 | 0.01014 | 0.00017321 | 0.22053 | 140 | 0.21325 | 0.94141 | 0.87121 | 0.99206 |  | 140 | 1.0244 | 1 | 1.0244 | 0.97674 | 1.0357 | 140 | 0.99225 | 1 |  |  |  | 140 | 0.93496 | 0.99187 |  | 1
16:18 | 62 | uS | 31 | ppm | 22.7 | C | 92 | 58.5 | 0.017094 | 0.032167 | 38.2 | µS |  |  | 144 | 108.8 | 154 | 42 | 101.6 | 0.0098425 | 0.0099263 | 0.00016956 | 0.22428 | 141 | 0.20898 | 0.94531 | 0.86364 | 0.97619 |  | 141 | 1.0244 | 1 | 1.0244 | 0.97674 | 1.0357 | 141 | 0.97674 | 0.9878 |  |  |  | 141 | 0.94309 | 0.98374 |  | 1
16:18 | 61.6 | uS | 30.8 | ppm | 22.7 | C | 93 | 62 | 0.016129 | 0.033132 | 41.7 | µS |  |  | 145 | 109.1 | 152 | 43 | 101.3667 | 0.0098652 | 0.0098538 | 0.00016832 | 0.22377 | 142 | 0.21494 | 0.94401 | 0.87121 | 0.97619 |  | 142 | 1.0122 | 1 | 1.0244 | 0.97674 | 1.0476 | 142 | 0.99225 | 0.9878 |  |  |  | 142 | 0.94309 | 0.99187 |  | 1
16:18 | 67.6 | uS | 33.8 | ppm | 22.7 | C | 94 | 61.6 | 0.016234 | 0.033027 | 41.3 | µS |  |  | 146 | 109.1 | 156 | 43 | 102.7 | 0.0097371 | 0.0098011 | 0.00016742 | 0.22671 | 143 | 0.22053 | 0.95833 | 0.87879 | 0.98413 |  | 143 | 1.0244 | 1 | 1.0244 | 0.97674 | 1.0357 | 143 | 0.9845 | 0.9878 |  |  |  | 143 | 0.94309 | 1 |  | 1.025
16:18 | 67.1 | uS | 33.6 | ppm | 22.7 | C | 95 | 67.6 | 0.014793 | 0.034468 | 47.3 | µS |  |  | 147 | 108.7 | 156 | 45 | 103.2333 | 0.0096868 | 0.0097119 | 0.0001659 | 0.22789 | 144 | 0.22428 | 0.94661 | 0.87879 | 0.97619 |  | 144 | 1.0122 |  | 1.0244 | 0.97674 | 1.0238 | 144 | 0.97674 | 1 |  |  |  | 144 | 0.95122 | 0.97561 |  | 1
16:18 | 66.6 | uS | 33.3 | ppm | 22.7 | C | 96 | 67.1 | 0.014903 | 0.034358 | 46.8 | µS |  |  | 148 | 104.5 | 159 | 45 | 102.8333 | 0.0097245 | 0.0097056 | 0.00016579 | 0.22701 | 145 | 0.22377 | 0.95964 | 0.87879 | 0.98413 |  | 145 | 1 |  | 1.0244 | 0.97674 | 1 | 145 | 0.97674 | 1 |  |  |  | 145 | 0.95122 | 1 |  | 1
16:18 | 70.3 | uS | 35.2 | ppm | 22.7 | C | 97 | 66.6 | 0.015015 | 0.034246 | 46.3 | µS |  |  | 149 | 109.9 | 159 | 47 | 105.3 | 0.0094967 | 0.0096106 | 0.00016417 | 0.23245 | 146 | 0.22671 | 0.96354 | 0.87879 | 0.97619 |  | 146 | 1.0122 |  | 1.0244 | 0.97674 | 1 | 146 | 0.97674 | 1 |  |  |  | 146 | 0.95122 | 1 |  | 1.025
16:18 | 70.3 | uS | 35.2 | ppm | 22.7 | C | 98 | 70.3 | 0.014225 | 0.035036 | 50 | µS |  |  | 150 | 110.5 | 162 | 47 | 106.5 | 0.0093897 | 0.0094432 | 0.00016131 | 0.2351 | 147 | 0.22789 | 0.96224 | 0.87879 | 0.96825 |  | 147 | 1.0244 |  | 1 | 1 | 1 | 147 | 0.9845 | 1 |  |  |  | 147 | 0.95122 | 0.97561 |  | 1.025
16:18 | 66 | uS | 33 | ppm | 22.7 | C | 99 | 70.3 | 0.014225 | 0.035036 | 50 | µS |  |  | 151 | 110.8 | 162 | 45 | 105.9333 | 0.0094399 | 0.0094148 | 0.00016082 | 0.23385 | 148 | 0.22701 | 1 | 0.87879 | 0.98413 |  | 148 | 1.0244 |  |  | 0.98413 |  | 148 | 0.99225 | 1 |  |  |  | 148 | 0.95122 | 0.97561 |  | 1
16:18 | 65.9 | uS | 33 | ppm | 22.7 | C | 100 | 66 | 0.015152 | 0.03411 | 45.7 | µS |  |  | 152 | 111.3 | 164 | 45 | 106.7667 | 0.0093662 | 0.0094031 | 0.00016062 | 0.23569 | 149 | 0.23245 | 1 | 0.87879 | 0.97619 |  | 149 | 1.0244 |  |  | 0.97619 |  | 149 | 0.9845 | 1 |  |  |  | 149 | 0.95122 | 0.97561 |  | 1.025
16:18 | 72 | uS | 36 | ppm | 22.7 | C | 101 | 65.9 | 0.015175 | 0.034087 | 45.6 | µS |  |  | 153 | 122.1 | 166 | 46 | 111.3667 | 0.0089793 | 0.0091728 | 0.00015669 | 0.24584 | 150 | 0.2351 | 1 | 0.87879 | 0.96825 |  | 150 | 1.0122 |  |  | 0.96825 |  | 150 | 0.9845 | 1 |  |  |  | 150 | 0.95122 | 1 |  | 1.025
16:18 | 72.4 | uS | 36.2 | ppm | 22.7 | C | 102 | 72 | 0.013889 | 0.035372 | 51.7 | µS |  |  | 154 | 121.5 | 168 | 46 | 111.8333 | 0.0089419 | 0.0089606 | 0.00015306 | 0.24687 | 151 | 0.23385 |  | 0.87879 | 0.97619 |  | 151 | 1 |  |  | 0.97619 |  | 151 | 0.9845 |  |  |  |  | 151 | 0.95122 | 1 |  | 1
16:18 | 72.2 | uS | 36.1 | ppm | 22.7 | C | 103 | 72.4 | 0.013812 | 0.035449 | 52.1 | µS |  |  | 155 | 117.2 | 169 | 46 | 110.7333 | 0.0090307 | 0.0089863 | 0.0001535 | 0.24444 | 152 | 0.23569 |  | 0.88636 | 0.96825 |  | 152 | 1 |  |  | 0.96825 |  | 152 | 0.9845 |  |  |  |  | 152 | 0.95935 | 0.97561 |  | 1.025
16:18 | 72.7 | uS | 36.4 | ppm | 22.7 | C | 104 | 72.2 | 0.01385 | 0.035411 | 51.9 | µS |  |  | 156 | 117.6 | 171 | 46 | 111.5333 | 0.0089659 | 0.0089983 | 0.00015371 | 0.24621 | 153 | 0.24584 |  | 0.88636 | 0.96825 |  | 153 | 1 |  |  | 0.96825 |  | 153 | 0.97674 |  |  |  |  | 153 | 0.96748 | 1 |  | 1.025
16:18 | 73.1 | uS | 36.6 | ppm | 22.7 | C | 105 | 72.7 | 0.013755 | 0.035506 | 52.4 | µS |  |  | 157 | 121.9 | 172 | 50 | 114.6333 | 0.0087235 | 0.0088447 | 0.00015108 | 0.25305 | 154 | 0.24687 |  | 0.87879 | 0.96825 |  | 154 | 1 |  |  | 0.96825 |  | 154 | 0.97674 |  |  |  |  | 154 | 0.96748 |  |  | 1
16:18 | 73.5 | uS | 36.8 | ppm | 22.7 | C | 106 | 73.1 | 0.01368 | 0.035581 | 52.8 | µS |  |  | 158 | 128.8 | 173 | 50 | 117.2667 | 0.0085276 | 0.0086255 | 0.00014734 | 0.25887 | 155 | 0.24444 |  | 0.87879 | 0.97619 | [5 empty] | 0.97619 |  | 155 | 0.97674 |  |  |  |  | 155 | 0.96748 |  |  | 1.025
16:18 | 73.5 | uS | 36.8 | ppm | 22.7 | C | 107 | 73.5 | 0.013605 | 0.035656 | 53.2 | µS |  |  | 159 | 124.7 | 174 | 50 | 116.2333 | 0.0086034 | 0.0085655 | 0.00014631 | 0.25659 | 156 | 0.24621 |  | 0.87879 | 0.97619 | [5 empty] | 0.97619 |  | 156 | 0.97674 |  |  |  |  | 156 | 0.97561 |  |  | 1
16:18 | 74.2 | uS | 37.1 | ppm | 22.7 | C | 108 | 73.5 | 0.013605 | 0.035656 | 53.2 | µS |  |  | 160 | 124.6 | 175 | 50 | 116.5333 | 0.0085812 | 0.0085923 | 0.00014677 | 0.25725 | 157 | 0.25305 |  | 0.88636 | 0.97619 | [5 empty] | 0.97619 |  | 157 | 1 |  |  |  |  | 157 | 0.97561 |  |  | 1
16:18 | 78.4 | uS | 39.2 | ppm | 22.7 | C | 109 | 74.2 | 0.013477 | 0.035784 | 53.9 | µS |  |  | 161 | 124 | 176 | 52 | 117.3333 | 0.0085227 | 0.008552 | 0.00014608 | 0.25901 | 158 | 0.25887 |  | 0.89394 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 158 | 0.96748 |  |  | 1
16:18 | 78 | uS | 39 | ppm | 22.7 | C | 110 | 78.4 | 0.012755 | 0.036506 | 58.1 | µS |  |  | 162 | 123.9 | 178 | 52 | 117.9667 | 0.008477 | 0.0084998 | 0.00014519 | 0.26041 | 159 | 0.25659 |  | 0.88636 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 159 | 0.97561 |  |  | 1
16:18 | 78.6 | uS | 39.3 | ppm | 22.7 | C | 111 | 78 | 0.012821 | 0.036441 | 57.7 | µS |  |  | 163 | 129.7 | 181 | 54 | 121.5667 | 0.0082259 | 0.0083515 | 0.00014266 | 0.26836 | 160 | 0.25725 |  | 0.87879 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 160 | 0.97561 |  |  | 1
16:18 | 83.9 | uS | 42 | ppm | 22.7 | C | 112 | 78.6 | 0.012723 | 0.036538 | 58.3 | µS |  |  | 164 | 125.4 | 182 | 54 | 120.4667 | 0.0083011 | 0.0082635 | 0.00014116 | 0.26593 | 161 | 0.25901 |  | 0.87879 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 161 | 0.96748 |  |  | 1
16:18 | 84.1 | uS | 42.1 | ppm | 22.7 | C | 113 | 83.9 | 0.011919 | 0.037342 | 63.6 | µS |  |  | 165 | 129.5 | 182 | 56 | 122.5 | 0.0081633 | 0.0082322 | 0.00014062 | 0.27042 | 162 | 0.26041 |  | 0.90152 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 162 | 0.97561 |  |  | 1
16:18 | 83.9 | uS | 42 | ppm | 22.7 | C | 114 | 84.1 | 0.011891 | 0.03737 | 63.8 | µS |  |  | 166 | 130.1 | 183 | 56 | 123.0333 | 0.0081279 | 0.0081456 | 0.00013914 | 0.2716 | 163 | 0.26836 |  | 0.89394 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 163 | 0.97561 |  |  | 1
16:18 | 87.6 | uS | 43.8 | ppm | 22.7 | C | 115 | 83.9 | 0.011919 | 0.037342 | 63.6 | µS |  |  | 167 | 134.8 | 186 | 59 | 126.6 | 0.0078989 | 0.0080134 | 0.00013688 | 0.27947 | 164 | 0.26593 |  | 0.88636 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 164 | 0.96748 |  |  | 1
16:18 | 87.7 | uS | 43.9 | ppm | 22.7 | C | 116 | 87.6 | 0.011416 | 0.037846 | 67.3 | µS |  |  | 168 | 135.7 | 187 | 59 | 127.2333 | 0.0078596 | 0.0078792 | 0.00013459 | 0.28087 | 165 | 0.27042 |  | 0.88636 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 165 | 0.97561 |  |  | 1
16:18 | 92.2 | uS | 46.1 | ppm | 22.7 | C | 117 | 87.7 | 0.011403 | 0.037859 | 67.4 | µS |  |  | 169 | 140.3 | 189 | 62 | 130.4333 | 0.0076668 | 0.0077632 | 0.00013261 | 0.28793 | 166 | 0.2716 |  | 0.89394 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 166 | 0.97561 |  |  | 1
16:18 | 92.1 | uS | 46.1 | ppm | 22.7 | C | 118 | 92.2 | 0.010846 | 0.038415 | 71.9 | µS |  |  | 170 | 140.9 | 191 | 62 | 131.3 | 0.0076161 | 0.0076414 | 0.00013053 | 0.28985 | 167 | 0.27947 |  | 0.88636 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 167 | 0.96748 |  |  | 1
16:18 | 91.5 | uS | 45.8 | ppm | 22.7 | C | 119 | 92.1 | 0.010858 | 0.038403 | 71.8 | µS |  |  | 171 | 138.8 | 194 | 66 | 132.9333 | 0.0075226 | 0.0075694 | 0.0001293 | 0.29345 | 168 | 0.28087 |  | 0.89394 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 168 | 0.97561 |  |  | 1
16:18 | 91 | uS | 45.5 | ppm | 22.7 | C | 120 | 91.5 | 0.010929 | 0.038332 | 71.2 | µS |  |  | 172 | 139.3 | 195 | 66 | 133.4333 | 0.0074944 | 0.0075085 | 0.00012826 | 0.29455 | 169 | 0.28793 |  | 0.90152 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 169 | 0.97561 |  |  | 1
16:18 | 91 | uS | 45.5 | ppm | 22.7 | C | 121 | 91 | 0.010989 | 0.038272 | 70.7 | µS |  |  | 173 | 139.4 | 197 | 66 | 134.1333 | 0.0074553 | 0.0074748 | 0.00012768 | 0.2961 | 170 | 0.28985 |  | 0.90152 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 170 | 0.97561 |  |  | 1
16:18 | 95.1 | uS | 47.6 | ppm | 22.7 | C | 122 | 91 | 0.010989 | 0.038272 | 70.7 | µS |  |  | 174 | 139.7 | 198 | 66 | 134.5667 | 0.0074313 | 0.0074433 | 0.00012715 | 0.29706 | 171 | 0.29345 |  | 0.89394 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 171 | 0.97561 |  |  | 1
16:18 | 98.8 | uS | 49.4 | ppm | 22.7 | C | 123 | 95.1 | 0.010515 | 0.038746 | 74.8 | µS |  |  | 175 | 139.8 | 201 | 68 | 136.2667 | 0.0073386 | 0.0073849 | 0.00012615 | 0.30081 | 172 | 0.29455 |  | 0.89394 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 172 | 0.97561 |  |  | 1
16:18 | 94.7 | uS | 47.4 | ppm | 22.7 | C | 124 | 98.8 | 0.010121 | 0.03914 | 78.5 | µS |  |  | 176 | 144.4 | 199 | 68 | 137.1333 | 0.0072922 | 0.0073154 | 0.00012496 | 0.30272 | 173 | 0.2961 |  | 0.90152 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 173 | 0.97561 |  |  | 1
16:18 | 98.8 | uS | 49.4 | ppm | 22.7 | C | 125 | 94.7 | 0.01056 | 0.038701 | 74.4 | µS |  |  | 177 | 152.9 | 200 | 68 | 140.3 | 0.0071276 | 0.0072099 | 0.00012316 | 0.30971 | 174 | 0.29706 |  | 0.89394 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 174 | 0.97561
16:18 | 106.5 | uS | 53.3 | ppm | 22.7 | C | 126 | 98.8 | 0.010121 | 0.03914 | 78.5 | µS |  |  | 178 | 153.1 | 197 | 68 | 139.3667 | 0.0071753 | 0.0071515 | 0.00012216 | 0.30765 | 175 | 0.30081 |  | 0.90152 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 175 | 0.97561
16:18 | 106.2 | uS | 53.1 | ppm | 22.7 | C | 127 | 106.5 | 0.0093897 | 0.039871 | 86.2 | µS |  |  | 179 | 149 | 199 | 69 | 139 | 0.0071942 | 0.0071848 | 0.00012273 | 0.30684 | 176 | 0.30272 |  | 0.90152 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 176 | 0.98374
16:18 | 105.3 | uS | 52.7 | ppm | 22.7 | C | 128 | 106.2 | 0.0094162 | 0.039845 | 85.9 | µS |  |  | 180 | 149.2 | 199 | 69 | 139.0667 | 0.0071908 | 0.0071925 | 0.00012286 | 0.30699 | 177 | 0.30971 |  | 0.90152 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 177 | 0.97561
16:18 | 105.4 | uS | 52.7 | ppm | 22.7 | C | 129 | 105.3 | 0.0094967 | 0.039764 | 85 | µS |  |  | 181 | 153.5 | 204 | 72 | 143.1667 | 0.0069849 | 0.0070878 | 0.00012107 | 0.31604 | 178 | 0.30765 |  | 0.90152 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 178 | 0.97561
16:18 | 111.3 | uS | 55.7 | ppm | 22.7 | C | 130 | 105.4 | 0.0094877 | 0.039773 | 85.1 | µS |  |  | 182 | 153.1 | 205 | 72 | 143.3667 | 0.0069751 | 0.00698 | 0.00011923 | 0.31648 | 179 | 0.30684 |  | 0.90152 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 179 | 0.98374
16:18 | 111.7 | uS | 55.9 | ppm | 22.7 | C | 131 | 111.3 | 0.0089847 | 0.040276 | 91 | µS |  |  | 183 | 153.2 | 208 | 75 | 145.4 | 0.0068776 | 0.0069264 | 0.00011832 | 0.32097 | 180 | 0.30699 |  | 0.89394 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 180 | 0.97561
16:18 | 107.8 | uS | 53.9 | ppm | 22.7 | C | 132 | 111.7 | 0.0089526 | 0.040309 | 91.4 | µS |  |  | 184 | 153.4 | 210 | 75 | 146.1333 | 0.0068431 | 0.0068603 | 0.00011719 | 0.32259 | 181 | 0.31604 |  | 0.90152 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 181 | 0.99187
16:18 | 112.5 | uS | 56.3 | ppm | 22.7 | C | 133 | 107.8 | 0.0092764 | 0.039985 | 87.5 | µS |  |  | 185 | 158 | 214 | 78 | 150 | 0.0066667 | 0.0067549 | 0.00011539 | 0.33113 | 182 | 0.31648 |  | 0.90152 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 182 | 0.99187
16:18 | 112.7 | uS | 56.4 | ppm | 22.7 | C | 134 | 112.5 | 0.0088889 | 0.040372 | 92.2 | µS |  |  | 186 | 165.8 | 216 | 78 | 153.2667 | 0.0065246 | 0.0065956 | 0.00011267 | 0.33834 | 183 | 0.32097 |  | 0.89394 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 183 | 0.98374
16:18 | 112.6 | uS | 56.3 | ppm | 22.7 | C | 135 | 112.7 | 0.0088731 | 0.040388 | 92.4 | µS |  |  | 187 | 165.9 | 218 | 77 | 153.6333 | 0.006509 | 0.0065168 | 0.00011132 | 0.33915 | 184 | 0.32259 |  | 0.90152 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 184 | 0.98374
16:18 | 117.2 | uS | 58.6 | ppm | 22.7 | C | 136 | 112.6 | 0.008881 | 0.04038 | 92.3 | µS |  |  | 188 | 165.5 | 219 | 77 | 153.8333 | 0.0065005 | 0.0065048 | 0.00011111 | 0.33959 | 185 | 0.33113 |  | 0.90152 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 185 | 0.98374
16:18 | 121.3 | uS | 60.7 | ppm | 22.7 | C | 137 | 117.2 | 0.0085324 | 0.040729 | 96.9 | µS |  |  | 189 | 161.5 | 220 | 79 | 153.5 | 0.0065147 | 0.0065076 | 0.00011116 | 0.33885 | 186 | 0.33834 |  | 0.90152 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 186 | 0.99187
16:18 | 121.5 | uS | 60.8 | ppm | 22.7 | C | 138 | 121.3 | 0.008244 | 0.041017 | 101 | µS |  |  | 190 | 165.8 | 222 | 79 | 155.6 | 0.0064267 | 0.0064707 | 0.00011053 | 0.34349 | 187 | 0.33915 |  | 0.90152 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 187 | 0.98374
16:18 | 122.1 | uS | 61.1 | ppm | 22.7 | C | 139 | 121.5 | 0.0082305 | 0.041031 | 101.2 | µS |  |  | 191 | 165.6 | 222 | 79 | 155.5333 | 0.0064295 | 0.0064281 | 0.0001098 | 0.34334 | 188 | 0.33959 |  | 0.90152 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 188 | 0.98374
16:18 | 118.3 | uS | 59.2 | ppm | 22.7 | C | 140 | 122.1 | 0.00819 | 0.041071 | 101.8 | µS |  |  | 192 | 165.1 | 223 | 79 | 155.7 | 0.0064226 | 0.006426 | 0.00010977 | 0.34371 | 189 | 0.33885 |  | 0.90152 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 189 | 0.98374
16:18 | 123.4 | uS | 61.7 | ppm | 22.7 | C | 141 | 118.3 | 0.0084531 | 0.040808 | 98 | µS |  |  | 193 | 164.9 | 223 | 81 | 156.3 | 0.006398 | 0.0064103 | 0.0001095 | 0.34503 | 190 | 0.34349 |  | 0.90152 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 190 | 0.99187
16:18 | 129 | uS | 64.5 | ppm | 22.7 | C | 142 | 123.4 | 0.0081037 | 0.041157 | 103.1 | µS |  |  | 194 | 164.8 | 224 | 81 | 156.6 | 0.0063857 | 0.0063918 | 0.00010918 | 0.3457 | 191 | 0.34334 |  | 0.90152 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 191 | 0.98374
16:18 | 129.1 | uS | 64.6 | ppm | 22.7 | C | 143 | 129 | 0.0077519 | 0.041509 | 108.7 | µS |  |  | 195 | 171.3 | 224 | 81 | 158.7667 | 0.0062986 | 0.0063421 | 0.00010834 | 0.35048 | 192 | 0.34371 |  | 0.90152 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 192 | 0.98374
16:18 | 129.4 | uS | 64.7 | ppm | 22.7 | C | 144 | 129.1 | 0.0077459 | 0.041515 | 108.8 | µS |  |  | 196 | 175.9 | 224 | 81 | 160.3 | 0.0062383 | 0.0062684 | 0.00010708 | 0.35386 | 193 | 0.34503 |  | 0.89394 | 0.96032 | [5 empty] | 0.96032 | [7 empty] | 193 | 0.99187
16:18 | 129.4 | uS | 64.7 | ppm | 22.7 | C | 145 | 129.4 | 0.007728 | 0.041533 | 109.1 | µS |  |  | 197 | 175.7 | 224 | 82 | 160.5667 | 0.0062279 | 0.0062331 | 0.00010647 | 0.35445 | 194 | 0.3457 |  | 0.89394 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 194 | 0.99187
16:18 | 129 | uS | 64.5 | ppm | 22.7 | C | 146 | 129.4 | 0.007728 | 0.041533 | 109.1 | µS |  |  | 198 | 175.6 | 226 | 82 | 161.2 | 0.0062035 | 0.0062157 | 0.00010618 | 0.35585 | 195 | 0.35048 |  | 0.90909 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 195 | 0.99187
16:19 | 124.8 | uS | 62.4 | ppm | 22.7 | C | 147 | 129 | 0.0077519 | 0.041509 | 108.7 | µS |  |  | 199 | 175.6 | 228 | 85 | 162.8667 | 0.00614 | 0.0061717 | 0.00010542 | 0.35953 | 196 | 0.35386 |  | 0.90909 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 196 | 0.99187
16:19 | 130.2 | uS | 65.1 | ppm | 22.7 | C | 148 | 124.8 | 0.0080128 | 0.041248 | 104.5 | µS |  |  | 200 | 174.8 | 230 | 85 | 163.2667 | 0.0061249 | 0.0061325 | 0.00010475 | 0.36041 | 197 | 0.35445 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 197 | 0.99187
16:19 | 130.8 | uS | 65.4 | ppm | 22.7 | C | 149 | 130.2 | 0.0076805 | 0.041581 | 109.9 | µS |  |  | 201 | 175.5 | 230 | 85 | 163.5 | 0.0061162 | 0.0061206 | 0.00010455 | 0.36093 | 198 | 0.35585 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 198 | 0.99187
16:19 | 131.1 | uS | 65.6 | ppm | 22.7 | C | 150 | 130.8 | 0.0076453 | 0.041616 | 110.5 | µS |  |  | 202 | 179.5 | 231 | 85 | 165.1667 | 0.0060545 | 0.0060853 | 0.00010395 | 0.36461 | 199 | 0.35953 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 199 | 0.99187
16:19 | 131.6 | uS | 65.8 | ppm | 22.7 | C | 151 | 131.1 | 0.0076278 | 0.041633 | 110.8 | µS |  |  | 203 | 179.6 | 229 | 88 | 165.5333 | 0.0060411 | 0.0060478 | 0.00010331 | 0.36542 | 200 | 0.36041 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 200 | 0.99187
16:19 | 142.4 | uS | 71.2 | ppm | 22.7 | C | 152 | 131.6 | 0.0075988 | 0.041662 | 111.3 | µS |  |  | 204 | 185.7 | 231 | 88 | 168.2333 | 0.0059441 | 0.0059926 | 0.00010237 | 0.37138 | 201 | 0.36093 |  | 0.90152 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 201 | 0.99187
16:19 | 141.8 | uS | 70.9 | ppm | 22.7 | C | 153 | 142.4 | 0.0070225 | 0.042239 | 122.1 | µS |  |  | 205 | 186.7 | 231 | 90 | 169.2333 | 0.005909 | 0.0059266 | 0.00010124 | 0.37358 | 202 | 0.36461 |  | 0.90909 | 0.99206 | [5 empty] | 0.99206 | [7 empty] | 202 | 0.99187
16:19 | 137.5 | uS | 68.8 | ppm | 22.7 | C | 154 | 141.8 | 0.0070522 | 0.042209 | 121.5 | µS |  |  | 206 | 187.7 | 235 | 90 | 170.9 | 0.0058514 | 0.0058802 | 0.00010044 | 0.37726 | 203 | 0.36542 |  | 0.90909 | 1 | [5 empty] | 1 | [7 empty] | 203 | 0.98374
16:19 | 137.9 | uS | 69 | ppm | 22.7 | C | 155 | 137.5 | 0.0072727 | 0.041988 | 117.2 | µS |  |  | 207 | 188.7 | 234 | 94 | 172.2333 | 0.0058061 | 0.0058287 | 9.9566e-05 | 0.38021 | 204 | 0.37138 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 204 | 0.99187
16:19 | 142.2 | uS | 71.1 | ppm | 22.7 | C | 156 | 137.9 | 0.0072516 | 0.042009 | 117.6 | µS |  |  | 208 | 191.7 | 237 | 94 | 174.2333 | 0.0057394 | 0.0057728 | 9.861e-05 | 0.38462 | 205 | 0.37358 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 205 | 0.99187
16:19 | 149.1 | uS | 74.6 | ppm | 22.7 | C | 157 | 142.2 | 0.0070323 | 0.042229 | 121.9 | µS |  |  | 209 | 191.7 | 236 | 97 | 174.9 | 0.0057176 | 0.0057285 | 9.7854e-05 | 0.38609 | 206 | 0.37726 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413 | [7 empty] | 206 | 0.9878
16:19 | 145 | uS | 72.5 | ppm | 22.7 | C | 158 | 149.1 | 0.0067069 | 0.042554 | 128.8 | µS |  |  | 210 | 194.7 | 237 | 97 | 176.2333 | 0.0056743 | 0.0056959 | 9.7297e-05 | 0.38904 | 207 | 0.38021 |  | 0.90152 | 0.97619 | [5 empty] | 0.97619 | [7 empty] | 207 | 0.9878
16:19 | 144.9 | uS | 72.5 | ppm | 22.7 | C | 159 | 145 | 0.0068966 | 0.042365 | 124.7 | µS |  |  | 211 | 196.7 | 238 | 97 | 177.2333 | 0.0056423 | 0.0056583 | 9.6654e-05 | 0.39124 | 208 | 0.38462 |  | 0.90909 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 208 | 0.97561
16:19 | 144.3 | uS | 72.2 | ppm | 22.7 | C | 160 | 144.9 | 0.0069013 | 0.04236 | 124.6 | µS |  |  | 212 | 197.7 | 239 | 97 | 177.9 | 0.0056211 | 0.0056317 | 9.62e-05 | 0.39272 | 209 | 0.38609 |  | 0.90909 | 0.96825 | [5 empty] | 0.96825 | [7 empty] | 209 | 1
16:19 | 144.2 | uS | 72.1 | ppm | 22.7 | C | 161 | 144.3 | 0.00693 | 0.042331 | 124 | µS |  |  | 213 | 196.7 | 239 | 98 | 177.9 | 0.0056211 | 0.0056211 | 9.602e-05 | 0.39272 | 210 | 0.38904 |  | 0.90909 | 0.97619 | [5 empty] | 0.97619
16:19 | 150 | uS | 75 | ppm | 22.7 | C | 162 | 144.2 | 0.0069348 | 0.042326 | 123.9 | µS |  |  | 214 | 198.7 | 240 | 98 | 178.9 | 0.0055897 | 0.0056054 | 9.5751e-05 | 0.39492 | 211 | 0.39124 |  | 0.90909 | 0.97619 | [5 empty] | 0.97619
16:19 | 145.7 | uS | 72.9 | ppm | 22.7 | C | 163 | 150 | 0.0066667 | 0.042594 | 129.7 | µS |  |  | 215 | 201.7 | 241 | 96 | 179.5667 | 0.005569 | 0.0055793 | 9.5306e-05 | 0.39639 | 212 | 0.39272 |  | 0.90909 | 0.96032 | [5 empty] | 0.96032
16:19 | 149.8 | uS | 74.9 | ppm | 22.7 | C | 164 | 145.7 | 0.0068634 | 0.042398 | 125.4 | µS |  |  | 216 | 201.7 | 241 | 96 | 179.5667 | 0.005569 | 0.005569 | 9.5128e-05 | 0.39639 | 213 | 0.39272 |  | 0.90909 | 0.97619 | [5 empty] | 0.97619
16:19 | 150.4 | uS | 75.2 | ppm | 22.7 | C | 165 | 149.8 | 0.0066756 | 0.042586 | 129.5 | µS |  |  | 217 | 203.7 | 242 | 101 | 182.2333 | 0.0054875 | 0.0055282 | 9.4432e-05 | 0.40228 | 214 | 0.39492 |  | 0.90909 | 0.97619 | [5 empty] | 0.97619
16:19 | 155.1 | uS | 77.6 | ppm | 22.7 | C | 166 | 150.4 | 0.0066489 | 0.042612 | 130.1 | µS |  |  | 218 | 202.7 | 244 | 101 | 182.5667 | 0.0054775 | 0.0054825 | 9.3651e-05 | 0.40302 | 215 | 0.39639 |  | 0.90909 | 0.96825 | [5 empty] | 0.96825
16:19 | 156 | uS | 78 | ppm | 22.7 | C | 167 | 155.1 | 0.0064475 | 0.042814 | 134.8 | µS |  |  | 219 | 203.7 | 247 | 103 | 184.5667 | 0.0054181 | 0.0054478 | 9.3058e-05 | 0.40743 | 216 | 0.39639 |  | 0.91667 | 0.97619 | [5 empty] | 0.97619
16:19 | 160.6 | uS | 80.3 | ppm | 22.7 | C | 168 | 156 | 0.0064103 | 0.042851 | 135.7 | µS |  |  | 220 | 204.7 | 248 | 103 | 185.2333 | 0.0053986 | 0.0054083 | 9.2385e-05 | 0.4089 | 217 | 0.40228 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413
16:19 | 161.2 | uS | 80.6 | ppm | 22.7 | C | 169 | 160.6 | 0.0062267 | 0.043034 | 140.3 | µS |  |  | 221 | 206.7 | 251 | 104 | 187.2333 | 0.0053409 | 0.0053698 | 9.1726e-05 | 0.41332 | 218 | 0.40302 |  | 0.90909 | 0.96825 | [5 empty] | 0.96825
16:19 | 159.1 | uS | 79.6 | ppm | 22.7 | C | 170 | 161.2 | 0.0062035 | 0.043058 | 140.9 | µS |  |  | 222 | 208.7 | 251 | 104 | 187.9 | 0.005322 | 0.0053315 | 9.1071e-05 | 0.41479 | 219 | 0.40743 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413
16:19 | 159.6 | uS | 79.8 | ppm | 22.7 | C | 171 | 159.1 | 0.0062854 | 0.042976 | 138.8 | µS |  |  | 223 | 207.7 | 253 | 108 | 189.5667 | 0.0052752 | 0.0052986 | 9.051e-05 | 0.41847 | 220 | 0.4089 |  | 0.90909 | 0.98413 | [5 empty] | 0.98413
16:19 | 159.7 | uS | 79.9 | ppm | 22.7 | C | 172 | 159.6 | 0.0062657 | 0.042995 | 139.3 | µS |  |  | 224 | 210.7 | 254 | 108 | 190.9 | 0.0052383 | 0.0052568 | 8.9796e-05 | 0.42141 | 221 | 0.41332 |  | 0.90909 | 0.96825 | [5 empty] | 0.96825
16:19 | 160 | uS | 80 | ppm | 22.7 | C | 173 | 159.7 | 0.0062617 | 0.042999 | 139.4 | µS |  |  | 225 | 209.7 | 255 | 109 | 191.2333 | 0.0052292 | 0.0052338 | 8.9403e-05 | 0.42215 | 222 | 0.41479 |  | 0.90909 | 0.97619 | [5 empty] | 0.97619
16:19 | 160.1 | uS | 80.1 | ppm | 22.7 | C | 174 | 160 | 0.00625 | 0.043011 | 139.7 | µS |  |  | 226 | 209.7 | 257 | 109 | 191.9 | 0.005211 | 0.0052201 | 8.917e-05 | 0.42362 | 223 | 0.41847 |  | 0.90909 | 1 | [5 empty] | 1
16:19 | 164.7 | uS | 82.4 | ppm | 22.7 | C | 175 | 160.1 | 0.0062461 | 0.043015 | 139.8 | µS |  |  | 227 | 211.7 | 258 | 111 | 193.5667 | 0.0051662 | 0.0051886 | 8.8631e-05 | 0.4273 | 224 | 0.42141 |  | 0.90909 | 1 | [5 empty] | 1
16:19 | 173.2 | uS | 86.6 | ppm | 22.7 | C | 176 | 164.7 | 0.0060716 | 0.043189 | 144.4 | µS |  |  | 228 | 213.7 | 257 | 111 | 193.9 | 0.0051573 | 0.0051617 | 8.8172e-05 | 0.42804 | 225 | 0.42215 |  | 0.90909 | 0.9881 | [5 empty] | 0.9881
16:19 | 173.4 | uS | 86.7 | ppm | 22.7 | C | 177 | 173.2 | 0.0057737 | 0.043487 | 152.9 | µS |  |  | 229 | 213.7 | 259 | 110 | 194.2333 | 0.0051484 | 0.0051529 | 8.8021e-05 | 0.42877 | 226 | 0.42362 |  | 0.90909 | 0.9881 | [5 empty] | 0.9881
16:19 | 169.3 | uS | 84.7 | ppm | 22.7 | C | 178 | 173.4 | 0.005767 | 0.043494 | 153.1 | µS |  |  | 230 | 213.7 | 259 | 110 | 194.2333 | 0.0051484 | 0.0051484 | 8.7945e-05 | 0.42877 | 227 | 0.4273 |  | 0.90909 | 1 | [5 empty] | 1
16:19 | 169.5 | uS | 84.8 | ppm | 22.7 | C | 179 | 169.3 | 0.0059067 | 0.043354 | 149 | µS |  |  | 231 | 215.7 | 262 | 112 | 196.5667 | 0.0050873 | 0.0051179 | 8.7423e-05 | 0.43392 | 228 | 0.42804 |  | 0.92424 | 1 | [5 empty] | 1
16:19 | 173.8 | uS | 86.9 | ppm | 22.7 | C | 180 | 169.5 | 0.0058997 | 0.043361 | 149.2 | µS |  |  | 232 | 214.7 | 260 | 112 | 195.5667 | 0.0051133 | 0.0051003 | 8.7124e-05 | 0.43171 | 229 | 0.42877 |  | 0.92424 | 1.0238 | [5 empty] | 1.0238
16:19 | 173.4 | uS | 86.7 | ppm | 22.7 | C | 181 | 173.8 | 0.0057537 | 0.043507 | 153.5 | µS |  |  | 233 | 218.7 | 261 | 114 | 197.9 | 0.0050531 | 0.0050832 | 8.6831e-05 | 0.43687 | 230 | 0.42877 |  | 0.90909 | 1.0238 | [5 empty] | 1.0238
16:19 | 173.5 | uS | 86.8 | ppm | 22.7 | C | 182 | 173.4 | 0.005767 | 0.043494 | 153.1 | µS |  |  | 234 | 218.7 | 261 | 114 | 197.9 | 0.0050531 | 0.0050531 | 8.6316e-05 | 0.43687 | 231 | 0.43392 |  | 0.90909 | 1.0238 | [5 empty] | 1.0238
16:19 | 173.7 | uS | 86.9 | ppm | 22.7 | C | 183 | 173.5 | 0.0057637 | 0.043497 | 153.2 | µS |  |  | 235 | 220.7 | 262 | 118 | 200.2333 | 0.0049942 | 0.0050236 | 8.5813e-05 | 0.44202 | 232 | 0.43171 |  | 0.88636 | 1 | [5 empty] | 1
16:19 | 178.3 | uS | 89.2 | ppm | 22.7 | C | 184 | 173.7 | 0.0057571 | 0.043504 | 153.4 | µS |  |  | 236 | 221.7 | 263 | 118 | 200.9 | 0.0049776 | 0.0049859 | 8.5168e-05 | 0.44349 | 233 | 0.43687 |  | 0.88636 | 1 | [5 empty] | 1
16:19 | 186.1 | uS | 93.1 | ppm | 22.7 | C | 185 | 178.3 | 0.0056085 | 0.043653 | 158 | µS |  |  | 237 | 223.7 | 266 | 120 | 203.2333 | 0.0049205 | 0.004949 | 8.4539e-05 | 0.44864 | 234 | 0.43687 |  | 0.89773 | 1 | [5 empty] | 1
16:19 | 186.2 | uS | 93.1 | ppm | 22.7 | C | 186 | 186.1 | 0.0053735 | 0.043888 | 165.8 | µS |  |  | 238 | 224.7 | 267 | 120 | 203.9 | 0.0049044 | 0.0049124 | 8.3913e-05 | 0.45011 | 235 | 0.44202 |  | 0.88636
16:19 | 185.8 | uS | 92.9 | ppm | 22.7 | C | 187 | 186.2 | 0.0053706 | 0.043891 | 165.9 | µS |  |  | 239 | 223.7 | 269 | 120 | 204.2333 | 0.0048964 | 0.0049004 | 8.3708e-05 | 0.45085 | 236 | 0.44349 |  | 0.88636
16:19 | 181.8 | uS | 90.9 | ppm | 22.7 | C | 188 | 185.8 | 0.0053821 | 0.043879 | 165.5 | µS |  |  | 240 | 225.7 | 269 | 120 | 204.9 | 0.0048804 | 0.0048884 | 8.3503e-05 | 0.45232 | 237 | 0.44864 |  | 0.88636
16:19 | 186.1 | uS | 93.1 | ppm | 22.7 | C | 189 | 181.8 | 0.0055006 | 0.043761 | 161.5 | µS |  |  | 241 | 224.7 | 270 | 122 | 205.5667 | 0.0048646 | 0.0048725 | 8.3232e-05 | 0.45379 | 238 | 0.45011 |  | 0.88636
16:19 | 185.9 | uS | 93 | ppm | 22.7 | C | 190 | 186.1 | 0.0053735 | 0.043888 | 165.8 | µS |  |  | 242 | 226.7 | 271 | 122 | 206.5667 | 0.0048411 | 0.0048528 | 8.2896e-05 | 0.456 | 239 | 0.45085 |  | 0.88636
16:19 | 185.4 | uS | 92.7 | ppm | 22.7 | C | 191 | 185.9 | 0.0053792 | 0.043882 | 165.6 | µS |  |  | 243 | 228.7 | 271 | 122 | 207.2333 | 0.0048255 | 0.0048333 | 8.2561e-05 | 0.45747 | 240 | 0.45232 |  | 0.89773
16:19 | 185.2 | uS | 92.6 | ppm | 22.7 | C | 192 | 185.4 | 0.0053937 | 0.043867 | 165.1 | µS |  |  | 244 | 229.7 | 273 | 122 | 208.2333 | 0.0048023 | 0.0048139 | 8.223e-05 | 0.45968 | 241 | 0.45379 |  | 0.89773
16:19 | 185.1 | uS | 92.6 | ppm | 22.7 | C | 193 | 185.2 | 0.0053996 | 0.043862 | 164.9 | µS |  |  | 245 | 229.7 | 273 | 126 | 209.5667 | 0.0047718 | 0.004787 | 8.1772e-05 | 0.46262 | 242 | 0.456 |  | 0.88636
16:19 | 191.6 | uS | 95.8 | ppm | 22.7 | C | 194 | 185.1 | 0.0054025 | 0.043859 | 164.8 | µS |  |  | 246 | 231.7 | 275 | 126 | 210.9 | 0.0047416 | 0.0047567 | 8.1253e-05 | 0.46556 | 243 | 0.45747 |  | 0.88636
16:19 | 196.2 | uS | 98.1 | ppm | 22.7 | C | 195 | 191.6 | 0.0052192 | 0.044042 | 171.3 | µS |  |  | 247 | 233.7 | 276 | 128 | 212.5667 | 0.0047044 | 0.004723 | 8.0678e-05 | 0.46924 | 244 | 0.45968 |  | 0.89773
16:19 | 196 | uS | 98 | ppm | 22.7 | C | 196 | 196.2 | 0.0050968 | 0.044164 | 175.9 | µS |  |  | 248 | 231.7 | 277 | 128 | 212.2333 | 0.0047118 | 0.0047081 | 8.0423e-05 | 0.46851 | 245 | 0.46262 |  | 0.89773
16:19 | 195.9 | uS | 98 | ppm | 22.7 | C | 197 | 196 | 0.005102 | 0.044159 | 175.7 | µS |  |  | 249 | 233.7 | 278 | 132 | 214.5667 | 0.0046606 | 0.0046862 | 8.0049e-05 | 0.47366 | 246 | 0.46556 |  | 0.89773
16:19 | 195.9 | uS | 98 | ppm | 22.7 | C | 198 | 195.9 | 0.0051046 | 0.044156 | 175.6 | µS |  |  | 250 | 233.7 | 280 | 132 | 215.2333 | 0.0046461 | 0.0046533 | 7.9488e-05 | 0.47513 | 247 | 0.46924 |  | 0.97727
16:19 | 195.1 | uS | 97.6 | ppm | 22.7 | C | 199 | 195.9 | 0.0051046 | 0.044156 | 175.6 | µS |  |  | 251 | 234.7 | 281 | 133 | 216.2333 | 0.0046246 | 0.0046354 | 7.9181e-05 | 0.47734 | 248 | 0.46851 |  | 0.97727
16:19 | 195.8 | uS | 97.9 | ppm | 22.7 | C | 200 | 195.1 | 0.0051256 | 0.044136 | 174.8 | µS |  |  | 252 | 234.7 | 283 | 133 | 216.9 | 0.0046104 | 0.0046175 | 7.8876e-05 | 0.47881 | 249 | 0.47366 |  | 1
16:19 | 199.8 | uS | 99.9 | ppm | 22.7 | C | 201 | 195.8 | 0.0051073 | 0.044154 | 175.5 | µS |  |  | 253 | 236.7 | 283 | 134 | 217.9 | 0.0045893 | 0.0045998 | 7.8574e-05 | 0.48102 | 250 | 0.47513 |  | 1
16:19 | 199.9 | uS | 100 | ppm | 22.7 | C | 202 | 199.8 | 0.005005 | 0.044256 | 179.5 | µS |  |  | 254 | 236.7 | 286 | 134 | 218.9 | 0.0045683 | 0.0045788 | 7.8214e-05 | 0.48322 | 251 | 0.47734 |  | 1
16:19 | 206 | uS | 103 | ppm | 22.7 | C | 203 | 199.9 | 0.0050025 | 0.044259 | 179.6 | µS |  |  | 255 | 237.7 | 284 | 134 | 218.5667 | 0.0045753 | 0.0045718 | 7.8095e-05 | 0.48249 | 252 | 0.47881 |  | 0.97727
16:19 | 207 | uS | 104 | ppm | 22.7 | C | 204 | 206 | 0.0048544 | 0.044407 | 185.7 | µS |  |  | 256 | 240.7 | 286 | 134 | 220.2333 | 0.0045406 | 0.004558 | 7.7858e-05 | 0.48617 | 253 | 0.48102 |  | 0.97727
16:19 | 208 | uS | 104 | ppm | 22.7 | C | 205 | 207 | 0.0048309 | 0.04443 | 186.7 | µS |  |  | 257 | 239.7 | 284 | 132 | 218.5667 | 0.0045753 | 0.004558 | 7.7858e-05 | 0.48249 | 254 | 0.48322 |  | 0.97727
16:20 | 209 | uS | 105 | ppm | 22.7 | C | 206 | 208 | 0.0048077 | 0.044453 | 187.7 | µS |  |  | 258 | 242.7 | 286 | 132 | 220.2333 | 0.0045406 | 0.004558 | 7.7858e-05 | 0.48617 | 255 | 0.48249 |  | 1
16:20 | 212 | uS | 106 | ppm | 22.7 | C | 207 | 209 | 0.0047847 | 0.044476 | 188.7 | µS |  |  | 259 | 242.7 | 287 | 136 | 221.9 | 0.0045065 | 0.0045236 | 7.7271e-05 | 0.48985 | 256 | 0.48617 |  | 1
16:20 | 212 | uS | 106 | ppm | 22.7 | C | 208 | 212 | 0.004717 | 0.044544 | 191.7 | µS |  |  | 260 | 244.7 | 290 | 136 | 223.5667 | 0.0044729 | 0.0044897 | 7.6693e-05 | 0.49352 | 257 | 0.48249 |  | 1
16:20 | 215 | uS | 108 | ppm | 22.7 | C | 209 | 212 | 0.004717 | 0.044544 | 191.7 | µS |  |  | 261 | 244.7 | 290 | 138 | 224.2333 | 0.0044596 | 0.0044663 | 7.6293e-05 | 0.495 | 258 | 0.48617 |  | 0.97727
16:20 | 217 | uS | 109 | ppm | 22.7 | C | 210 | 215 | 0.0046512 | 0.04461 | 194.7 | µS |  |  | 262 | 245.7 | 292 | 138 | 225.2333 | 0.0044398 | 0.0044497 | 7.601e-05 | 0.4972 | 259 | 0.48985 |  | 0.97727
16:20 | 218 | uS | 109 | ppm | 22.7 | C | 211 | 217 | 0.0046083 | 0.044653 | 196.7 | µS |  |  | 263 | 246.7 | 293 | 140 | 226.5667 | 0.0044137 | 0.0044268 | 7.5618e-05 | 0.50015 | 260 | 0.49352 |  | 1
16:20 | 217 | uS | 109 | ppm | 22.7 | C | 212 | 218 | 0.0045872 | 0.044674 | 197.7 | µS |  |  | 264 | 245.7 | 293 | 140 | 226.2333 | 0.0044202 | 0.004417 | 7.545e-05 | 0.49941 | 261 | 0.495 |  | 1
16:20 | 219 | uS | 110 | ppm | 22.7 | C | 213 | 217 | 0.0046083 | 0.044653 | 196.7 | µS |  |  | 265 | 247.7 | 295 | 142 | 228.2333 | 0.0043815 | 0.0044008 | 7.5175e-05 | 0.50383 | 262 | 0.4972 |  | 1
16:20 | 222 | uS | 111 | ppm | 22.7 | C | 214 | 219 | 0.0045662 | 0.044695 | 198.7 | µS |  |  | 266 | 244.7 | 297 | 142 | 227.9 | 0.0043879 | 0.0043847 | 7.4899e-05 | 0.50309 | 263 | 0.50015 |  | 1
16:20 | 222 | uS | 111 | ppm | 22.7 | C | 215 | 222 | 0.0045045 | 0.044757 | 201.7 | µS |  |  | 267 | 247.7 | 299 | 143 | 229.9 | 0.0043497 | 0.0043688 | 7.4627e-05 | 0.50751 | 264 | 0.49941 |  | 1
16:20 | 224 | uS | 112 | ppm | 22.7 | C | 216 | 222 | 0.0045045 | 0.044757 | 201.7 | µS |  |  | 268 | 248.7 | 300 | 143 | 230.5667 | 0.0043371 | 0.0043434 | 7.4194e-05 | 0.50898 | 265 | 0.50383 |  | 1
16:20 | 223 | uS | 112 | ppm | 22.7 | C | 217 | 224 | 0.0044643 | 0.044797 | 203.7 | µS |  |  | 269 | 249.7 | 302 | 144 | 231.9 | 0.0043122 | 0.0043247 | 7.3874e-05 | 0.51192 | 266 | 0.50309
16:20 | 224 | uS | 112 | ppm | 22.7 | C | 218 | 223 | 0.0044843 | 0.044777 | 202.7 | µS |  |  | 270 | 248.7 | 302 | 144 | 231.5667 | 0.0043184 | 0.0043153 | 7.3714e-05 | 0.51118 | 267 | 0.50751
16:20 | 225 | uS | 113 | ppm | 22.7 | C | 219 | 224 | 0.0044643 | 0.044797 | 203.7 | µS |  |  | 271 | 249.7 | 302 | 142 | 231.2333 | 0.0043246 | 0.0043215 | 7.382e-05 | 0.51045 | 268 | 0.50898
16:20 | 227 | uS | 114 | ppm | 22.7 | C | 220 | 225 | 0.0044444 | 0.044817 | 204.7 | µS |  |  | 272 | 250.7 | 302 | 142 | 231.5667 | 0.0043184 | 0.0043215 | 7.382e-05 | 0.51118 | 269 | 0.51192
16:20 | 229 | uS | 115 | ppm | 22.7 | C | 221 | 227 | 0.0044053 | 0.044856 | 206.7 | µS |  |  | 273 | 250.7 | 304 | 145 | 233.2333 | 0.0042876 | 0.004303 | 7.3503e-05 | 0.51486 | 270 | 0.51118
16:20 | 228 | uS | 114 | ppm | 22.7 | C | 222 | 229 | 0.0043668 | 0.044894 | 208.7 | µS |  |  | 274 | 253.7 | 305 | 145 | 234.5667 | 0.0042632 | 0.0042754 | 7.3031e-05 | 0.51781 | 271 | 0.51045
16:20 | 231 | uS | 116 | ppm | 22.7 | C | 223 | 228 | 0.004386 | 0.044875 | 207.7 | µS |  |  | 275 | 252.7 | 308 | 145 | 235.2333 | 0.0042511 | 0.0042571 | 7.272e-05 | 0.51928 | 272 | 0.51118
16:20 | 230 | uS | 115 | ppm | 22.7 | C | 224 | 231 | 0.004329 | 0.044932 | 210.7 | µS |  |  | 276 | 254.7 | 306 | 145 | 235.2333 | 0.0042511 | 0.0042511 | 7.2617e-05 | 0.51928 | 273 | 0.51486
16:20 | 230 | uS | 115 | ppm | 22.7 | C | 225 | 230 | 0.0043478 | 0.044913 | 209.7 | µS |  |  | 277 | 257.7 | 310 | 149 | 238.9 | 0.0041859 | 0.0042185 | 7.206e-05 | 0.52737 | 274 | 0.51781
16:20 | 232 | uS | 116 | ppm | 22.7 | C | 226 | 230 | 0.0043478 | 0.044913 | 209.7 | µS |  |  | 278 | 258.7 | 306 | 149 | 237.9 | 0.0042034 | 0.0041946 | 7.1653e-05 | 0.52517 | 275 | 0.51928
16:20 | 234 | uS | 117 | ppm | 22.7 | C | 227 | 232 | 0.0043103 | 0.044951 | 211.7 | µS |  |  | 279 | 258.7 | 308 | 151 | 239.2333 | 0.00418 | 0.0041917 | 7.1603e-05 | 0.52811 | 276 | 0.51928
16:20 | 234 | uS | 117 | ppm | 22.7 | C | 228 | 234 | 0.0042735 | 0.044988 | 213.7 | µS |  |  | 280 | 259.7 | 309 | 151 | 239.9 | 0.0041684 | 0.0041742 | 7.1303e-05 | 0.52958 | 277 | 0.52737
16:20 | 234 | uS | 117 | ppm | 22.7 | C | 229 | 234 | 0.0042735 | 0.044988 | 213.7 | µS |  |  | 281 | 260.7 | 311 | 151 | 240.9 | 0.0041511 | 0.0041598 | 7.1056e-05 | 0.53179 | 278 | 0.52517
16:20 | 236 | uS | 118 | ppm | 22.7 | C | 230 | 234 | 0.0042735 | 0.044988 | 213.7 | µS |  |  | 282 | 259.7 | 311 | 151 | 240.5667 | 0.0041569 | 0.004154 | 7.0958e-05 | 0.53105 | 279 | 0.52811
16:20 | 235 | uS | 118 | ppm | 22.7 | C | 231 | 236 | 0.0042373 | 0.045024 | 215.7 | µS |  |  | 283 | 262.7 | 313 | 154 | 243.2333 | 0.0041113 | 0.0041341 | 7.0618e-05 | 0.53694 | 280 | 0.52958
16:20 | 239 | uS | 120 | ppm | 22.7 | C | 232 | 235 | 0.0042553 | 0.045006 | 214.7 | µS |  |  | 284 | 262.7 | 314 | 154 | 243.5667 | 0.0041057 | 0.0041085 | 7.018e-05 | 0.53767 | 281 | 0.53179
16:20 | 239 | uS | 120 | ppm | 22.7 | C | 233 | 239 | 0.0041841 | 0.045077 | 218.7 | µS |  |  | 285 | 264.7 | 316 | 156 | 245.5667 | 0.0040722 | 0.0040889 | 6.9847e-05 | 0.54209 | 282 | 0.53105
16:20 | 241 | uS | 121 | ppm | 22.7 | C | 234 | 239 | 0.0041841 | 0.045077 | 218.7 | µS |  |  | 286 | 266.7 | 318 | 156 | 246.9 | 0.0040502 | 0.0040612 | 6.9373e-05 | 0.54503 | 283 | 0.53694
16:20 | 242 | uS | 121 | ppm | 22.7 | C | 235 | 241 | 0.0041494 | 0.045112 | 220.7 | µS |  |  | 287 | 266.7 | 318 | 160 | 248.2333 | 0.0040285 | 0.0040393 | 6.9e-05 | 0.54798 | 284 | 0.53767
16:20 | 244 | uS | 122 | ppm | 22.7 | C | 236 | 242 | 0.0041322 | 0.045129 | 221.7 | µS |  |  | 288 | 266.7 | 320 | 160 | 248.9 | 0.0040177 | 0.0040231 | 6.8722e-05 | 0.54945 | 285 | 0.54209
16:20 | 245 | uS | 123 | ppm | 22.7 | C | 237 | 244 | 0.0040984 | 0.045163 | 223.7 | µS |  |  | 289 | 266.7 | 319 | 159 | 248.2333 | 0.0040285 | 0.0040231 | 6.8722e-05 | 0.54798 | 286 | 0.54503
16:20 | 244 | uS | 122 | ppm | 22.7 | C | 238 | 245 | 0.0040816 | 0.045179 | 224.7 | µS |  |  | 290 | 268.7 | 321 | 159 | 249.5667 | 0.0040069 | 0.0040177 | 6.863e-05 | 0.55092 | 287 | 0.54798
16:20 | 246 | uS | 123 | ppm | 22.7 | C | 239 | 244 | 0.0040984 | 0.045163 | 223.7 | µS |  |  | 291 | 266.7 | 322 | 160 | 249.5667 | 0.0040069 | 0.0040069 | 6.8446e-05 | 0.55092 | 288 | 0.54945
16:20 | 245 | uS | 123 | ppm | 22.7 | C | 240 | 246 | 0.004065 | 0.045196 | 225.7 | µS |  |  | 292 | 269.7 | 322 | 160 | 250.5667 | 0.003991 | 0.0039989 | 6.831e-05 | 0.55313 | 289 | 0.54798
16:20 | 247 | uS | 124 | ppm | 22.7 | C | 241 | 245 | 0.0040816 | 0.045179 | 224.7 | µS |  |  | 293 | 270.7 | 323 | 161 | 251.5667 | 0.0039751 | 0.003983 | 6.8038e-05 | 0.55533 | 290 | 0.55092
16:20 | 249 | uS | 125 | ppm | 22.7 | C | 242 | 247 | 0.0040486 | 0.045213 | 226.7 | µS |  |  | 294 | 271.7 | 323 | 161 | 251.9 | 0.0039698 | 0.0039725 | 6.7857e-05 | 0.55607 | 291 | 0.55092
16:20 | 250 | uS | 125 | ppm | 22.7 | C | 243 | 249 | 0.0040161 | 0.045245 | 228.7 | µS |  |  | 295 | 272.7 | 323 | 161 | 252.2333 | 0.0039646 | 0.0039672 | 6.7767e-05 | 0.55681 | 292 | 0.55313
16:20 | 250 | uS | 125 | ppm | 22.7 | C | 244 | 250 | 0.004 | 0.045261 | 229.7 | µS |  |  | 296 | 274.7 | 325 | 161 | 253.5667 | 0.0039437 | 0.0039542 | 6.7545e-05 | 0.55975 | 293 | 0.55533
16:20 | 252 | uS | 126 | ppm | 22.7 | C | 245 | 250 | 0.004 | 0.045261 | 229.7 | µS |  |  | 297 | 274.7 | 323 | 162 | 253.2333 | 0.0039489 | 0.0039463 | 6.7411e-05 | 0.55901 | 294 | 0.55607
16:20 | 254 | uS | 127 | ppm | 22.7 | C | 246 | 252 | 0.0039683 | 0.045293 | 231.7 | µS |  |  | 298 | 274.7 | 326 | 162 | 254.2333 | 0.0039334 | 0.0039412 | 6.7323e-05 | 0.56122 | 295 | 0.55681
16:20 | 252 | uS | 126 | ppm | 22.7 | C | 247 | 254 | 0.003937 | 0.045324 | 233.7 | µS |  |  | 299 | 276.7 | 324 | 161 | 253.9 | 0.0039386 | 0.003936 | 6.7234e-05 | 0.56049 | 296 | 0.55975
16:20 | 254 | uS | 127 | ppm | 22.7 | C | 248 | 252 | 0.0039683 | 0.045293 | 231.7 | µS |  |  | 300 | 273.7 | 327 | 161 | 253.9 | 0.0039386 | 0.0039386 | 6.7278e-05 | 0.56049 | 297 | 0.55901
16:20 | 254 | uS | 127 | ppm | 22.7 | C | 249 | 254 | 0.003937 | 0.045324 | 233.7 | µS |  |  | 301 | 275.7 | 326 | 164 | 255.2333 | 0.003918 | 0.0039283 | 6.7102e-05 | 0.56343 | 298 | 0.56122
16:20 | 255 | uS | 128 | ppm | 22.7 | C | 250 | 254 | 0.003937 | 0.045324 | 233.7 | µS |  |  | 302 | 275.7 | 328 | 164 | 255.9 | 0.0039078 | 0.0039129 | 6.6839e-05 | 0.5649 | 299 | 0.56049
16:20 | 255 | uS | 128 | ppm | 22.7 | C | 251 | 255 | 0.0039216 | 0.04534 | 234.7 | µS |  |  | 303 | 276.7 | 328 | 165 | 256.5667 | 0.0038976 | 0.0039027 | 6.6666e-05 | 0.56637 | 300 | 0.56049
16:20 | 257 | uS | 129 | ppm | 22.7 | C | 252 | 255 | 0.0039216 | 0.04534 | 234.7 | µS |  |  | 304 | 277.7 | 330 | 165 | 257.5667 | 0.0038825 | 0.0038901 | 6.645e-05 | 0.56858 | 301 | 0.56343
16:20 | 257 | uS | 129 | ppm | 22.7 | C | 253 | 257 | 0.0038911 | 0.04537 | 236.7 | µS |  |  | 305 | 279.7 | 330 | 168 | 259.2333 | 0.0038575 | 0.00387 | 6.6107e-05 | 0.57226 | 302 | 0.5649
16:20 | 258 | uS | 129 | ppm | 22.7 | C | 254 | 257 | 0.0038911 | 0.04537 | 236.7 | µS |  |  | 306 | 279.7 | 332 | 168 | 259.9 | 0.0038476 | 0.0038526 | 6.5809e-05 | 0.57373 | 303 | 0.56637
16:20 | 261 | uS | 131 | ppm | 22.7 | C | 255 | 258 | 0.003876 | 0.045385 | 237.7 | µS |  |  | 307 | 279.7 | 333 | 169 | 260.5667 | 0.0038378 | 0.0038427 | 6.5641e-05 | 0.5752 | 304 | 0.56858
16:20 | 260 | uS | 130 | ppm | 22.7 | C | 256 | 261 | 0.0038314 | 0.04543 | 240.7 | µS |  |  | 308 | 280.7 | 333 | 169 | 260.9 | 0.0038329 | 0.0038353 | 6.5515e-05 | 0.57594 | 305 | 0.57226
16:20 | 263 | uS | 132 | ppm | 22.7 | C | 257 | 260 | 0.0038462 | 0.045415 | 239.7 | µS |  |  | 309 | 278.7 | 334 | 169 | 260.5667 | 0.0038378 | 0.0038353 | 6.5515e-05 | 0.5752 | 306 | 0.57373
16:20 | 263 | uS | 132 | ppm | 22.7 | C | 258 | 263 | 0.0038023 | 0.045459 | 242.7 | µS |  |  | 310 | 280.7 | 333 | 169 | 260.9 | 0.0038329 | 0.0038353 | 6.5515e-05 | 0.57594 | 307 | 0.5752
16:20 | 265 | uS | 133 | ppm | 22.7 | C | 259 | 263 | 0.0038023 | 0.045459 | 242.7 | µS |  |  | 311 | 281.7 | 333 | 172 | 262.2333 | 0.0038134 | 0.0038231 | 6.5307e-05 | 0.57888 | 308 | 0.57594
16:20 | 265 | uS | 133 | ppm | 22.7 | C | 260 | 265 | 0.0037736 | 0.045487 | 244.7 | µS |  |  | 312 | 283.7 | 334 | 172 | 263.2333 | 0.0037989 | 0.0038062 | 6.5016e-05 | 0.58109 | 309 | 0.5752
16:20 | 266 | uS | 133 | ppm | 22.7 | C | 261 | 265 | 0.0037736 | 0.045487 | 244.7 | µS |  |  | 313 | 285.7 | 336 | 174 | 265.2333 | 0.0037703 | 0.0037846 | 6.4648e-05 | 0.5855 | 310 | 0.57594
16:20 | 267 | uS | 134 | ppm | 22.7 | C | 262 | 266 | 0.0037594 | 0.045502 | 245.7 | µS |  |  | 314 | 287.7 | 337 | 174 | 266.2333 | 0.0037561 | 0.0037632 | 6.4282e-05 | 0.58771 | 311 | 0.57888
16:20 | 266 | uS | 133 | ppm | 22.7 | C | 263 | 267 | 0.0037453 | 0.045516 | 246.7 | µS |  |  | 315 | 287.7 | 338 | 179 | 268.2333 | 0.0037281 | 0.0037421 | 6.3922e-05 | 0.59213 | 312 | 0.58109
16:20 | 268 | uS | 134 | ppm | 22.7 | C | 264 | 266 | 0.0037594 | 0.045502 | 245.7 | µS |  |  | 316 | 288.7 | 337 | 179 | 268.2333 | 0.0037281 | 0.0037281 | 6.3683e-05 | 0.59213 | 313 | 0.5855
16:20 | 265 | uS | 133 | ppm | 22.7 | C | 265 | 268 | 0.0037313 | 0.04553 | 247.7 | µS |  |  | 317 | 290.7 | 339 | 180 | 269.9 | 0.0037051 | 0.0037166 | 6.3486e-05 | 0.59581 | 314 | 0.58771
16:20 | 268 | uS | 134 | ppm | 22.7 | C | 266 | 265 | 0.0037736 | 0.045487 | 244.7 | µS |  |  | 318 | 288.7 | 338 | 180 | 268.9 | 0.0037189 | 0.003712 | 6.3407e-05 | 0.5936 | 315 | 0.59213
16:21 | 269 | uS | 135 | ppm | 22.7 | C | 267 | 268 | 0.0037313 | 0.04553 | 247.7 | µS |  |  | 319 | 291.7 | 342 | 183 | 272.2333 | 0.0036733 | 0.0036961 | 6.3136e-05 | 0.60096 | 316 | 0.59213
16:21 | 270 | uS | 135 | ppm | 22.7 | C | 268 | 269 | 0.0037175 | 0.045544 | 248.7 | µS |  |  | 320 | 291.7 | 340 | 183 | 271.5667 | 0.0036823 | 0.0036778 | 6.2824e-05 | 0.59948 | 317 | 0.59581
16:21 | 269 | uS | 135 | ppm | 22.7 | C | 269 | 270 | 0.0037037 | 0.045557 | 249.7 | µS |  |  | 321 | 292.7 | 342 | 185 | 273.2333 | 0.0036599 | 0.0036711 | 6.2709e-05 | 0.60316 | 318 | 0.5936
16:21 | 270 | uS | 135 | ppm | 22.7 | C | 270 | 269 | 0.0037175 | 0.045544 | 248.7 | µS |  |  | 322 | 292.7 | 342 | 185 | 273.2333 | 0.0036599 | 0.0036599 | 6.2518e-05 | 0.60316 | 319 | 0.60096
16:21 | 271 | uS | 136 | ppm | 22.7 | C | 271 | 270 | 0.0037037 | 0.045557 | 249.7 | µS |  |  | 323 | 293.7 | 344 | 185 | 274.2333 | 0.0036465 | 0.0036532 | 6.2404e-05 | 0.60537 | 320 | 0.59948
16:21 | 271 | uS | 136 | ppm | 22.7 | C | 272 | 271 | 0.00369 | 0.045571 | 250.7 | µS |  |  | 324 | 294.7 | 344 | 185 | 274.5667 | 0.0036421 | 0.0036443 | 6.2252e-05 | 0.60611 | 321 | 0.60316
16:21 | 274 | uS | 137 | ppm | 22.7 | C | 273 | 271 | 0.00369 | 0.045571 | 250.7 | µS |  |  | 325 | 294.7 | 345 | 187 | 275.5667 | 0.0036289 | 0.0036355 | 6.2101e-05 | 0.60831 | 322 | 0.60316
16:21 | 273 | uS | 137 | ppm | 22.7 | C | 274 | 274 | 0.0036496 | 0.045611 | 253.7 | µS |  |  | 326 | 296.7 | 345 | 187 | 276.2333 | 0.0036201 | 0.0036245 | 6.1913e-05 | 0.60979 | 323 | 0.60537
16:21 | 275 | uS | 138 | ppm | 22.7 | C | 275 | 273 | 0.003663 | 0.045598 | 252.7 | µS |  |  | 327 | 294.7 | 346 | 188 | 276.2333 | 0.0036201 | 0.0036201 | 6.1839e-05 | 0.60979 | 324 | 0.60611
16:21 | 278 | uS | 139 | ppm | 22.7 | C | 276 | 275 | 0.0036364 | 0.045625 | 254.7 | µS |  |  | 328 | 297.7 | 347 | 188 | 277.5667 | 0.0036027 | 0.0036114 | 6.169e-05 | 0.61273 | 325 | 0.60831
16:21 | 279 | uS | 140 | ppm | 22.7 | C | 277 | 278 | 0.0035971 | 0.045664 | 257.7 | µS |  |  | 329 | 297.7 | 346 | 193 | 278.9 | 0.0035855 | 0.0035941 | 6.1395e-05 | 0.61567 | 326 | 0.60979
16:21 | 279 | uS | 140 | ppm | 22.7 | C | 278 | 279 | 0.0035842 | 0.045677 | 258.7 | µS |  |  | 330 | 297.7 | 347 | 193 | 279.2333 | 0.0035812 | 0.0035834 | 6.1211e-05 | 0.61641 | 327 | 0.60979
16:21 | 280 | uS | 140 | ppm | 22.7 | C | 279 | 279 | 0.0035842 | 0.045677 | 258.7 | µS |  |  | 331 | 299.7 | 347 | 194 | 280.2333 | 0.0035685 | 0.0035748 | 6.1065e-05 | 0.61862 | 328 | 0.61273
16:21 | 281 | uS | 141 | ppm | 22.7 | C | 280 | 280 | 0.0035714 | 0.04569 | 259.7 | µS |  |  | 332 | 301.7 | 348 | 194 | 281.2333 | 0.0035558 | 0.0035621 | 6.0848e-05 | 0.62082 | 329 | 0.61567
16:21 | 280 | uS | 140 | ppm | 22.7 | C | 281 | 281 | 0.0035587 | 0.045702 | 260.7 | µS |  |  | 333 | 301.7 | 349 | 195 | 281.9 | 0.0035474 | 0.0035516 | 6.0667e-05 | 0.6223 | 330 | 0.61641
16:21 | 283 | uS | 142 | ppm | 22.7 | C | 282 | 280 | 0.0035714 | 0.04569 | 259.7 | µS |  |  | 334 | 301.7 | 350 | 195 | 282.2333 | 0.0035432 | 0.0035453 | 6.056e-05 | 0.62303 | 331 | 0.61862
16:21 | 283 | uS | 142 | ppm | 22.7 | C | 283 | 283 | 0.0035336 | 0.045728 | 262.7 | µS |  |  | 335 | 302.7 | 350 | 196 | 282.9 | 0.0035348 | 0.003539 | 6.0453e-05 | 0.6245 | 332 | 0.62082
16:21 | 285 | uS | 143 | ppm | 22.7 | C | 284 | 283 | 0.0035336 | 0.045728 | 262.7 | µS |  |  | 336 | 301.7 | 351 | 196 | 282.9 | 0.0035348 | 0.0035348 | 6.0381e-05 | 0.6245 | 333 | 0.6223
16:21 | 287 | uS | 144 | ppm | 22.7 | C | 285 | 285 | 0.0035088 | 0.045752 | 264.7 | µS |  |  | 337 | 303.7 | 350 | 198 | 283.9 | 0.0035224 | 0.0035286 | 6.0275e-05 | 0.62671 | 334 | 0.62303
16:21 | 287 | uS | 144 | ppm | 22.7 | C | 286 | 287 | 0.0034843 | 0.045777 | 266.7 | µS |  |  | 338 | 303.7 | 353 | 198 | 284.9 | 0.00351 | 0.0035162 | 6.0063e-05 | 0.62892 | 335 | 0.6245
16:21 | 287 | uS | 144 | ppm | 22.7 | C | 287 | 287 | 0.0034843 | 0.045777 | 266.7 | µS |  |  | 339 | 305.7 | 351 | 202 | 286.2333 | 0.0034937 | 0.0035018 | 5.9818e-05 | 0.63186 | 336 | 0.6245
16:21 | 287 | uS | 144 | ppm | 22.7 | C | 288 | 287 | 0.0034843 | 0.045777 | 266.7 | µS |  |  | 340 | 306.7 | 354 | 202 | 287.5667 | 0.0034775 | 0.0034856 | 5.954e-05 | 0.63481 | 337 | 0.62671
16:21 | 289 | uS | 145 | ppm | 22.7 | C | 289 | 287 | 0.0034843 | 0.045777 | 266.7 | µS |  |  | 341 | 307.7 | 354 | 201 | 287.5667 | 0.0034775 | 0.0034775 | 5.9402e-05 | 0.63481 | 338 | 0.62892
16:21 | 287 | uS | 144 | ppm | 22.7 | C | 290 | 289 | 0.0034602 | 0.045801 | 268.7 | µS |  |  | 342 | 307.7 | 356 | 201 | 288.2333 | 0.0034694 | 0.0034734 | 5.9333e-05 | 0.63628 | 339 | 0.63186
16:21 | 290 | uS | 145 | ppm | 22.7 | C | 291 | 287 | 0.0034843 | 0.045777 | 266.7 | µS |  |  | 343 | 308.7 | 356 | 204 | 289.5667 | 0.0034534 | 0.0034614 | 5.9128e-05 | 0.63922 | 340 | 0.63481
16:21 | 291 | uS | 146 | ppm | 22.7 | C | 292 | 290 | 0.0034483 | 0.045813 | 269.7 | µS |  |  | 344 | 309.7 | 357 | 204 | 290.2333 | 0.0034455 | 0.0034495 | 5.8924e-05 | 0.64069 | 341 | 0.63481
16:21 | 292 | uS | 146 | ppm | 22.7 | C | 293 | 291 | 0.0034364 | 0.045825 | 270.7 | µS |  |  | 345 | 309.7 | 358 | 207 | 291.5667 | 0.0034297 | 0.0034376 | 5.8721e-05 | 0.64364 | 342 | 0.63628
16:21 | 293 | uS | 147 | ppm | 22.7 | C | 294 | 292 | 0.0034247 | 0.045836 | 271.7 | µS |  |  | 346 | 314.7 | 359 | 207 | 293.5667 | 0.0034064 | 0.0034181 | 5.8387e-05 | 0.64805 | 343 | 0.63922
16:21 | 295 | uS | 148 | ppm | 22.7 | C | 295 | 293 | 0.003413 | 0.045848 | 272.7 | µS |  |  | 347 | 313.7 | 361 | 209 | 294.5667 | 0.0033948 | 0.0034006 | 5.8089e-05 | 0.65026 | 344 | 0.64069
16:21 | 295 | uS | 148 | ppm | 22.7 | C | 296 | 295 | 0.0033898 | 0.045871 | 274.7 | µS |  |  | 348 | 314.7 | 361 | 209 | 294.9 | 0.003391 | 0.0033929 | 5.7957e-05 | 0.65099 | 345 | 0.64364
16:21 | 295 | uS | 148 | ppm | 22.7 | C | 297 | 295 | 0.0033898 | 0.045871 | 274.7 | µS |  |  | 349 | 315.7 | 363 | 210 | 296.2333 | 0.0033757 | 0.0033833 | 5.7794e-05 | 0.65394 | 346 | 0.64805
16:21 | 297 | uS | 149 | ppm | 22.7 | C | 298 | 295 | 0.0033898 | 0.045871 | 274.7 | µS |  |  | 350 | 314.7 | 362 | 210 | 295.5667 | 0.0033833 | 0.0033795 | 5.7729e-05 | 0.65247 | 347 | 0.65026
16:21 | 294 | uS | 147 | ppm | 22.7 | C | 299 | 297 | 0.003367 | 0.045894 | 276.7 | µS |  |  | 351 | 314.7 | 363 | 209 | 295.5667 | 0.0033833 | 0.0033833 | 5.7794e-05 | 0.65247 | 348 | 0.65099
16:21 | 296 | uS | 148 | ppm | 22.7 | C | 300 | 294 | 0.0034014 | 0.04586 | 273.7 | µS |  |  | 352 | 316.7 | 364 | 209 | 296.5667 | 0.0033719 | 0.0033776 | 5.7696e-05 | 0.65467 | 349 | 0.65394
16:21 | 296 | uS | 148 | ppm | 22.7 | C | 301 | 296 | 0.0033784 | 0.045883 | 275.7 | µS |  |  | 353 | 317.7 | 365 | 212 | 298.2333 | 0.0033531 | 0.0033625 | 5.7438e-05 | 0.65835 | 350 | 0.65247
16:21 | 297 | uS | 149 | ppm | 22.7 | C | 302 | 296 | 0.0033784 | 0.045883 | 275.7 | µS |  |  | 354 | 318.7 | 364 | 212 | 298.2333 | 0.0033531 | 0.0033531 | 5.7277e-05 | 0.65835 | 351 | 0.65247
16:21 | 298 | uS | 149 | ppm | 22.7 | C | 303 | 297 | 0.003367 | 0.045894 | 276.7 | µS |  |  | 355 | 320.7 | 366 | 213 | 299.9 | 0.0033344 | 0.0033438 | 5.7118e-05 | 0.66203 | 352 | 0.65467
16:21 | 300 | uS | 150 | ppm | 22.7 | C | 304 | 298 | 0.0033557 | 0.045905 | 277.7 | µS |  |  | 356 | 318.7 | 364 | 213 | 298.5667 | 0.0033493 | 0.0033419 | 5.7086e-05 | 0.65909 | 353 | 0.65835
16:21 | 300 | uS | 150 | ppm | 22.7 | C | 305 | 300 | 0.0033333 | 0.045928 | 279.7 | µS |  |  | 357 | 320.7 | 367 | 217 | 301.5667 | 0.003316 | 0.0033327 | 5.6928e-05 | 0.66571 | 354 | 0.65835
16:21 | 300 | uS | 150 | ppm | 22.7 | C | 306 | 300 | 0.0033333 | 0.045928 | 279.7 | µS |  |  | 358 | 321.7 | 365 | 217 | 301.2333 | 0.0033197 | 0.0033179 | 5.6675e-05 | 0.66497 | 355 | 0.66203
16:21 | 301 | uS | 151 | ppm | 22.7 | C | 307 | 300 | 0.0033333 | 0.045928 | 279.7 | µS |  |  | 359 | 322.7 | 367 | 217 | 302.2333 | 0.0033087 | 0.0033142 | 5.6613e-05 | 0.66718 | 356 | 0.65909
16:21 | 299 | uS | 150 | ppm | 22.7 | C | 308 | 301 | 0.0033223 | 0.045939 | 280.7 | µS |  |  | 360 | 322.7 | 367 | 217 | 302.2333 | 0.0033087 | 0.0033087 | 5.6519e-05 | 0.66718 | 357 | 0.66571
16:21 | 301 | uS | 151 | ppm | 22.7 | C | 309 | 299 | 0.0033445 | 0.045917 | 278.7 | µS |  |  | 361 | 323.7 | 369 | 219 | 303.9 | 0.0032906 | 0.0032996 | 5.6364e-05 | 0.67086 | 358 | 0.66497
16:21 | 302 | uS | 151 | ppm | 22.7 | C | 310 | 301 | 0.0033223 | 0.045939 | 280.7 | µS |  |  | 362 | 323.7 | 370 | 219 | 304.2333 | 0.003287 | 0.0032888 | 5.6178e-05 | 0.6716 | 359 | 0.66718
16:21 | 304 | uS | 152 | ppm | 22.7 | C | 311 | 302 | 0.0033113 | 0.04595 | 281.7 | µS |  |  | 363 | 325.7 | 371 | 221 | 305.9 | 0.003269 | 0.003278 | 5.5994e-05 | 0.67528 | 360 | 0.66718
16:21 | 306 | uS | 153 | ppm | 22.7 | C | 312 | 304 | 0.0032895 | 0.045972 | 283.7 | µS |  |  | 364 | 327.7 | 371 | 221 | 306.5667 | 0.0032619 | 0.0032655 | 5.5781e-05 | 0.67675 | 361 | 0.67086
16:21 | 308 | uS | 154 | ppm | 22.7 | C | 313 | 306 | 0.003268 | 0.045993 | 285.7 | µS |  |  | 365 | 326.7 | 371 | 220 | 305.9 | 0.003269 | 0.0032655 | 5.5781e-05 | 0.67528 | 362 | 0.6716
16:21 | 308 | uS | 154 | ppm | 22.7 | C | 314 | 308 | 0.0032468 | 0.046014 | 287.7 | µS |  |  | 366 | 329.7 | 372 | 220 | 307.2333 | 0.0032549 | 0.0032619 | 5.572e-05 | 0.67822 | 363 | 0.67528
16:21 | 309 | uS | 155 | ppm | 22.7 | C | 315 | 308 | 0.0032468 | 0.046014 | 287.7 | µS |  |  | 367 | 328.7 | 372 | 221 | 307.2333 | 0.0032549 | 0.0032549 | 5.5599e-05 | 0.67822 | 364 | 0.67675
16:21 | 311 | uS | 156 | ppm | 22.7 | C | 316 | 309 | 0.0032362 | 0.046025 | 288.7 | µS |  |  | 368 | 329.7 | 373 | 221 | 307.9 | 0.0032478 | 0.0032513 | 5.5539e-05 | 0.67969 | 365 | 0.67528
16:21 | 309 | uS | 155 | ppm | 22.7 | C | 317 | 311 | 0.0032154 | 0.046046 | 290.7 | µS |  |  | 369 | 330.7 | 374 | 219 | 307.9 | 0.0032478 | 0.0032478 | 5.5479e-05 | 0.67969 | 366 | 0.67822
16:21 | 312 | uS | 156 | ppm | 22.7 | C | 318 | 309 | 0.0032362 | 0.046025 | 288.7 | µS |  |  | 370 | 330.7 | 374 | 219 | 307.9 | 0.0032478 | 0.0032478 | 5.5479e-05 | 0.67969 | 367 | 0.67822
16:21 | 312 | uS | 156 | ppm | 22.7 | C | 319 | 312 | 0.0032051 | 0.046056 | 291.7 | µS |  |  | 371 | 330.7 | 374 | 223 | 309.2333 | 0.0032338 | 0.0032408 | 5.5359e-05 | 0.68263 | 368 | 0.67969
16:21 | 313 | uS | 157 | ppm | 22.7 | C | 320 | 312 | 0.0032051 | 0.046056 | 291.7 | µS |  |  | 372 | 331.7 | 376 | 223 | 310.2333 | 0.0032234 | 0.0032286 | 5.5151e-05 | 0.68484 | 369 | 0.67969
16:21 | 313 | uS | 157 | ppm | 22.7 | C | 321 | 313 | 0.0031949 | 0.046066 | 292.7 | µS |  |  | 373 | 332.7 | 375 | 225 | 310.9 | 0.0032165 | 0.0032199 | 5.5002e-05 | 0.68631 | 370 | 0.67969
16:21 | 314 | uS | 157 | ppm | 22.7 | C | 322 | 313 | 0.0031949 | 0.046066 | 292.7 | µS |  |  | 374 | 331.7 | 377 | 225 | 311.2333 | 0.003213 | 0.0032147 | 5.4914e-05 | 0.68705 | 371 | 0.68263
16:21 | 315 | uS | 158 | ppm | 22.7 | C | 323 | 314 | 0.0031847 | 0.046076 | 293.7 | µS |  |  | 375 | 334.7 | 375 | 225 | 311.5667 | 0.0032096 | 0.0032113 | 5.4855e-05 | 0.68779 | 372 | 0.68484
16:22 | 315 | uS | 158 | ppm | 22.7 | C | 324 | 315 | 0.0031746 | 0.046086 | 294.7 | µS |  |  | 376 | 333.7 | 378 | 225 | 312.2333 | 0.0032027 | 0.0032062 | 5.4767e-05 | 0.68926 | 373 | 0.68631
16:22 | 317 | uS | 159 | ppm | 22.7 | C | 325 | 315 | 0.0031746 | 0.046086 | 294.7 | µS |  |  | 377 | 334.7 | 376 | 225 | 311.9 | 0.0032062 | 0.0032044 | 5.4738e-05 | 0.68852 | 374 | 0.68705
16:22 | 315 | uS | 158 | ppm | 22.7 | C | 326 | 317 | 0.0031546 | 0.046107 | 296.7 | µS |  |  | 378 | 335.7 | 378 | 225 | 312.9 | 0.0031959 | 0.003201 | 5.468e-05 | 0.69073 | 375 | 0.68779
16:22 | 318 | uS | 159 | ppm | 22.7 | C | 327 | 315 | 0.0031746 | 0.046086 | 294.7 | µS |  |  | 379 | 336.7 | 378 | 224 | 312.9 | 0.0031959 | 0.0031959 | 5.4592e-05 | 0.69073 | 376 | 0.68926
16:22 | 318 | uS | 159 | ppm | 22.7 | C | 328 | 318 | 0.0031447 | 0.046116 | 297.7 | µS |  |  | 380 | 337.7 | 380 | 224 | 313.9 | 0.0031857 | 0.0031908 | 5.4505e-05 | 0.69294 | 377 | 0.68852
16:22 | 318 | uS | 159 | ppm | 22.7 | C | 329 | 318 | 0.0031447 | 0.046116 | 297.7 | µS |  |  | 381 | 339.7 | 380 | 228 | 315.9 | 0.0031656 | 0.0031756 | 5.4246e-05 | 0.69735 | 378 | 0.69073
16:22 | 320 | uS | 160 | ppm | 22.7 | C | 330 | 318 | 0.0031447 | 0.046116 | 297.7 | µS |  |  | 382 | 340.7 | 381 | 228 | 316.5667 | 0.0031589 | 0.0031622 | 5.4017e-05 | 0.69882 | 379 | 0.69073
16:22 | 322 | uS | 161 | ppm | 22.7 | C | 331 | 320 | 0.003125 | 0.046136 | 299.7 | µS |  |  | 383 | 341.7 | 382 | 226 | 316.5667 | 0.0031589 | 0.0031589 | 5.396e-05 | 0.69882 | 380 | 0.69294
16:22 | 322 | uS | 161 | ppm | 22.7 | C | 332 | 322 | 0.0031056 | 0.046155 | 301.7 | µS |  |  | 384 | 342.7 | 381 | 226 | 316.5667 | 0.0031589 | 0.0031589 | 5.396e-05 | 0.69882 | 381 | 0.69735
16:22 | 322 | uS | 161 | ppm | 22.7 | C | 333 | 322 | 0.0031056 | 0.046155 | 301.7 | µS |  |  | 385 | 341.7 | 382 | 230 | 317.9 | 0.0031456 | 0.0031523 | 5.3847e-05 | 0.70177 | 382 | 0.69882
16:22 | 323 | uS | 162 | ppm | 22.7 | C | 334 | 322 | 0.0031056 | 0.046155 | 301.7 | µS |  |  | 386 | 344.7 | 382 | 230 | 318.9 | 0.0031358 | 0.0031407 | 5.3649e-05 | 0.70397 | 383 | 0.69882
16:22 | 322 | uS | 161 | ppm | 22.7 | C | 335 | 323 | 0.003096 | 0.046165 | 302.7 | µS |  |  | 387 | 345.7 | 383 | 231 | 319.9 | 0.003126 | 0.0031309 | 5.3481e-05 | 0.70618 | 384 | 0.69882
16:22 | 324 | uS | 162 | ppm | 22.7 | C | 336 | 322 | 0.0031056 | 0.046155 | 301.7 | µS |  |  | 388 | 345.7 | 384 | 231 | 320.2333 | 0.0031227 | 0.0031243 | 5.337e-05 | 0.70692 | 385 | 0.70177
16:22 | 324 | uS | 162 | ppm | 22.7 | C | 337 | 324 | 0.0030864 | 0.046175 | 303.7 | µS |  |  | 389 | 346.7 | 385 | 233 | 321.5667 | 0.0031098 | 0.0031162 | 5.3231e-05 | 0.70986 | 386 | 0.70397
16:22 | 326 | uS | 163 | ppm | 22.7 | C | 338 | 324 | 0.0030864 | 0.046175 | 303.7 | µS |  |  | 390 | 347.7 | 385 | 233 | 321.9 | 0.0031066 | 0.0031082 | 5.3093e-05 | 0.7106 | 387 | 0.70618
16:22 | 327 | uS | 164 | ppm | 22.7 | C | 339 | 326 | 0.0030675 | 0.046194 | 305.7 | µS |  |  | 391 | 347.7 | 387 | 235 | 323.2333 | 0.0030937 | 0.0031001 | 5.2956e-05 | 0.71354 | 388 | 0.70692
16:22 | 328 | uS | 164 | ppm | 22.7 | C | 340 | 327 | 0.0030581 | 0.046203 | 306.7 | µS |  |  | 392 | 347.7 | 385 | 235 | 322.5667 | 0.0031001 | 0.0030969 | 5.2902e-05 | 0.71207 | 389 | 0.70986
16:22 | 328 | uS | 164 | ppm | 22.7 | C | 341 | 328 | 0.0030488 | 0.046212 | 307.7 | µS |  |  | 393 | 349.7 | 387 | 237 | 324.5667 | 0.003081 | 0.0030906 | 5.2793e-05 | 0.71648 | 390 | 0.7106
16:22 | 329 | uS | 165 | ppm | 22.7 | C | 342 | 328 | 0.0030488 | 0.046212 | 307.7 | µS |  |  | 394 | 348.7 | 385 | 237 | 323.5667 | 0.0030906 | 0.0030858 | 5.2711e-05 | 0.71428 | 391 | 0.71354
16:22 | 330 | uS | 165 | ppm | 22.7 | C | 343 | 329 | 0.0030395 | 0.046222 | 308.7 | µS |  |  | 395 | 351.7 | 386 | 239 | 325.5667 | 0.0030716 | 0.0030811 | 5.263e-05 | 0.71869 | 392 | 0.71207
16:22 | 330 | uS | 165 | ppm | 22.7 | C | 344 | 330 | 0.0030303 | 0.046231 | 309.7 | µS |  |  | 396 | 349.7 | 386 | 239 | 324.9 | 0.0030779 | 0.0030747 | 5.2522e-05 | 0.71722 | 393 | 0.71648
16:22 | 335 | uS | 168 | ppm | 22.7 | C | 345 | 330 | 0.0030303 | 0.046231 | 309.7 | µS |  |  | 397 | 351.7 | 387 | 237 | 325.2333 | 0.0030747 | 0.0030763 | 5.2549e-05 | 0.71795 | 394 | 0.71428
16:22 | 334 | uS | 167 | ppm | 22.7 | C | 346 | 335 | 0.0029851 | 0.046276 | 314.7 | µS |  |  | 398 | 352.7 | 387 | 237 | 325.5667 | 0.0030716 | 0.0030731 | 5.2495e-05 | 0.71869 | 395 | 0.71869
16:22 | 335 | uS | 168 | ppm | 22.7 | C | 347 | 334 | 0.002994 | 0.046267 | 313.7 | µS |  |  | 399 | 353.7 | 388 | 239 | 326.9 | 0.003059 | 0.0030653 | 5.2361e-05 | 0.72163 | 396 | 0.71722
16:22 | 336 | uS | 168 | ppm | 22.7 | C | 348 | 335 | 0.0029851 | 0.046276 | 314.7 | µS |  |  | 400 | 353.7 | 388 | 239 | 326.9 | 0.003059 | 0.003059 | 5.2254e-05 | 0.72163 | 397 | 0.71795
16:22 | 335 | uS | 168 | ppm | 22.7 | C | 349 | 336 | 0.0029762 | 0.046285 | 315.7 | µS |  |  | 401 | 354.7 | 388 | 239 | 327.2333 | 0.0030559 | 0.0030575 | 5.2228e-05 | 0.72237 | 398 | 0.71869
16:22 | 335 | uS | 168 | ppm | 22.7 | C | 350 | 335 | 0.0029851 | 0.046276 | 314.7 | µS |  |  | 402 | 354.7 | 389 | 239 | 327.5667 | 0.0030528 | 0.0030544 | 5.2174e-05 | 0.72311 | 399 | 0.72163
16:22 | 337 | uS | 169 | ppm | 22.7 | C | 351 | 335 | 0.0029851 | 0.046276 | 314.7 | µS |  |  | 403 | 354.7 | 389 | 239 | 327.5667 | 0.0030528 | 0.0030528 | 5.2148e-05 | 0.72311 | 400 | 0.72163
16:22 | 338 | uS | 169 | ppm | 22.7 | C | 352 | 337 | 0.0029674 | 0.046294 | 316.7 | µS |  |  | 404 | 354.7 | 389 | 239 | 327.5667 | 0.0030528 | 0.0030528 | 5.2148e-05 | 0.72311 | 401 | 0.72237
16:22 | 339 | uS | 170 | ppm | 22.7 | C | 353 | 338 | 0.0029586 | 0.046303 | 317.7 | µS |  |  | 405 | 353.7 | 390 | 241 | 328.2333 | 0.0030466 | 0.0030497 | 5.2095e-05 | 0.72458 | 402 | 0.72311
16:22 | 341 | uS | 171 | ppm | 22.7 | C | 354 | 339 | 0.0029499 | 0.046311 | 318.7 | µS |  |  | 406 | 355.7 | 390 | 241 | 328.9 | 0.0030404 | 0.0030435 | 5.1989e-05 | 0.72605 | 403 | 0.72311
16:22 | 339 | uS | 170 | ppm | 22.7 | C | 355 | 341 | 0.0029326 | 0.046329 | 320.7 | µS |  |  | 407 | 354.7 | 390 | 243 | 329.2333 | 0.0030374 | 0.0030389 | 5.191e-05 | 0.72678 | 404 | 0.72311
16:22 | 341 | uS | 171 | ppm | 22.7 | C | 356 | 339 | 0.0029499 | 0.046311 | 318.7 | µS |  |  | 408 | 355.7 | 392 | 243 | 330.2333 | 0.0030282 | 0.0030328 | 5.1805e-05 | 0.72899 | 405 | 0.72458
16:22 | 342 | uS | 171 | ppm | 22.7 | C | 357 | 341 | 0.0029326 | 0.046329 | 320.7 | µS |  |  | 409 | 356.7 | 390 | 248 | 331.5667 | 0.003016 | 0.0030221 | 5.1623e-05 | 0.73194 | 406 | 0.72605
16:22 | 343 | uS | 172 | ppm | 22.7 | C | 358 | 342 | 0.002924 | 0.046337 | 321.7 | µS |  |  | 410 | 356.7 | 392 | 248 | 332.2333 | 0.0030099 | 0.003013 | 5.1467e-05 | 0.73341 | 407 | 0.72678
16:22 | 343 | uS | 172 | ppm | 22.7 | C | 359 | 343 | 0.0029155 | 0.046346 | 322.7 | µS |  |  | 411 | 356.7 | 391 | 248 | 331.9 | 0.003013 | 0.0030114 | 5.1441e-05 | 0.73267 | 408 | 0.72899
16:22 | 344 | uS | 172 | ppm | 22.7 | C | 360 | 343 | 0.0029155 | 0.046346 | 322.7 | µS |  |  | 412 | 358.7 | 392 | 248 | 332.9 | 0.0030039 | 0.0030084 | 5.139e-05 | 0.73488 | 409 | 0.73194
16:22 | 344 | uS | 172 | ppm | 22.7 | C | 361 | 344 | 0.002907 | 0.046354 | 323.7 | µS |  |  | 413 | 359.7 | 392 | 249 | 333.5667 | 0.0029979 | 0.0030009 | 5.1261e-05 | 0.73635 | 410 | 0.73341
16:22 | 346 | uS | 173 | ppm | 22.7 | C | 362 | 344 | 0.002907 | 0.046354 | 323.7 | µS |  |  | 414 | 358.7 | 393 | 249 | 333.5667 | 0.0029979 | 0.0029979 | 5.121e-05 | 0.73635 | 411 | 0.73267
16:22 | 348 | uS | 174 | ppm | 22.7 | C | 363 | 346 | 0.0028902 | 0.046371 | 325.7 | µS |  |  | 415 | 360.7 | 393 | 250 | 334.5667 | 0.0029889 | 0.0029934 | 5.1133e-05 | 0.73856 | 412 | 0.73488
16:22 | 347 | uS | 174 | ppm | 22.7 | C | 364 | 348 | 0.0028736 | 0.046388 | 327.7 | µS |  |  | 416 | 357.7 | 394 | 250 | 333.9 | 0.0029949 | 0.0029919 | 5.1108e-05 | 0.73709 | 413 | 0.73635
16:22 | 350 | uS | 175 | ppm | 22.7 | C | 365 | 347 | 0.0028818 | 0.046379 | 326.7 | µS |  |  | 417 | 360.7 | 394 | 250 | 334.9 | 0.002986 | 0.0029904 | 5.1082e-05 | 0.73929 | 414 | 0.73635
16:22 | 349 | uS | 175 | ppm | 22.7 | C | 366 | 350 | 0.0028571 | 0.046404 | 329.7 | µS |  |  | 418 | 359.7 | 393 | 250 | 334.2333 | 0.0029919 | 0.0029889 | 5.1057e-05 | 0.73782 | 415 | 0.73856
16:22 | 350 | uS | 175 | ppm | 22.7 | C | 367 | 349 | 0.0028653 | 0.046396 | 328.7 | µS |  |  | 419 | 360.7 | 395 | 252 | 335.9 | 0.0029771 | 0.0029845 | 5.0981e-05 | 0.7415 | 416 | 0.73709
16:22 | 351 | uS | 176 | ppm | 22.7 | C | 368 | 350 | 0.0028571 | 0.046404 | 329.7 | µS |  |  | 420 | 361.7 | 395 | 252 | 336.2333 | 0.0029741 | 0.0029756 | 5.0829e-05 | 0.74224 | 417 | 0.73929
16:22 | 351 | uS | 176 | ppm | 22.7 | C | 369 | 351 | 0.002849 | 0.046412 | 330.7 | µS |  |  | 421 | 362.7 | 396 | 252 | 336.9 | 0.0029682 | 0.0029712 | 5.0753e-05 | 0.74371 | 418 | 0.73782
16:22 | 351 | uS | 176 | ppm | 22.7 | C | 370 | 351 | 0.002849 | 0.046412 | 330.7 | µS |  |  | 422 | 361.7 | 397 | 252 | 336.9 | 0.0029682 | 0.0029682 | 5.0703e-05 | 0.74371 | 419 | 0.7415
16:22 | 352 | uS | 176 | ppm | 22.7 | C | 371 | 351 | 0.002849 | 0.046412 | 330.7 | µS |  |  | 423 | 362.7 | 397 | 257 | 338.9 | 0.0029507 | 0.0029595 | 5.0554e-05 | 0.74812 | 420 | 0.74224
16:22 | 353 | uS | 177 | ppm | 22.7 | C | 372 | 352 | 0.0028409 | 0.04642 | 331.7 | µS |  |  | 424 | 363.7 | 396 | 257 | 338.9 | 0.0029507 | 0.0029507 | 5.0404e-05 | 0.74812 | 421 | 0.74371
16:22 | 352 | uS | 176 | ppm | 22.7 | C | 373 | 353 | 0.0028329 | 0.046428 | 332.7 | µS |  |  | 425 | 363.7 | 398 | 258 | 339.9 | 0.002942 | 0.0029464 | 5.033e-05 | 0.75033 | 422 | 0.74371
16:22 | 355 | uS | 178 | ppm | 22.7 | C | 374 | 352 | 0.0028409 | 0.04642 | 331.7 | µS |  |  | 426 | 365.7 | 396 | 258 | 339.9 | 0.002942 | 0.002942 | 5.0256e-05 | 0.75033 | 423 | 0.74812
16:22 | 354 | uS | 177 | ppm | 22.7 | C | 375 | 355 | 0.0028169 | 0.046444 | 334.7 | µS |  |  | 427 | 363.7 | 398 | 261 | 340.9 | 0.0029334 | 0.0029377 | 5.0182e-05 | 0.75254 | 424 | 0.74812
16:22 | 355 | uS | 178 | ppm | 22.7 | C | 376 | 354 | 0.0028249 | 0.046436 | 333.7 | µS |  |  | 428 | 366.7 | 397 | 261 | 341.5667 | 0.0029277 | 0.0029305 | 5.0059e-05 | 0.75401 | 425 | 0.75033
16:22 | 356 | uS | 178 | ppm | 22.7 | C | 377 | 355 | 0.0028169 | 0.046444 | 334.7 | µS |  |  | 429 | 366.7 | 398 | 261 | 341.9 | 0.0029248 | 0.0029263 | 4.9986e-05 | 0.75475 | 426 | 0.75033
16:22 | 357 | uS | 179 | ppm | 22.7 | C | 378 | 356 | 0.002809 | 0.046452 | 335.7 | µS |  |  | 430 | 366.7 | 398 | 261 | 341.9 | 0.0029248 | 0.0029248 | 4.9962e-05 | 0.75475 | 427 | 0.75254
16:22 | 358 | uS | 179 | ppm | 22.7 | C | 379 | 357 | 0.0028011 | 0.04646 | 336.7 | µS |  |  | 431 | 367.7 | 399 | 261 | 342.5667 | 0.0029191 | 0.002922 | 4.9913e-05 | 0.75622 | 428 | 0.75401
16:22 | 360 | uS | 180 | ppm | 22.7 | C | 380 | 358 | 0.0027933 | 0.046468 | 337.7 | µS |  |  | 432 | 368.7 | 400 | 261 | 343.2333 | 0.0029135 | 0.0029163 | 4.9816e-05 | 0.75769 | 429 | 0.75475
16:22 | 361 | uS | 181 | ppm | 22.7 | C | 381 | 360 | 0.0027778 | 0.046483 | 339.7 | µS |  |  | 433 | 367.7 | 400 | 263 | 343.5667 | 0.0029106 | 0.0029121 | 4.9743e-05 | 0.75843 | 430 | 0.75475
16:22 | 362 | uS | 181 | ppm | 22.7 | C | 382 | 361 | 0.0027701 | 0.046491 | 340.7 | µS |  |  | 434 | 368.7 | 401 | 263 | 344.2333 | 0.002905 | 0.0029078 | 4.9671e-05 | 0.7599 | 431 | 0.75622
16:22 | 363 | uS | 182 | ppm | 22.7 | C | 383 | 362 | 0.0027624 | 0.046499 | 341.7 | µS |  |  | 435 | 369.7 | 401 | 263 | 344.5667 | 0.0029022 | 0.0029036 | 4.9599e-05 | 0.76063 | 432 | 0.75769
16:22 | 362 | uS | 181 | ppm | 22.7 | C | 384 | 363 | 0.0027548 | 0.046506 | 342.7 | µS |  |  | 436 | 369.7 | 402 | 263 | 344.9 | 0.0028994 | 0.0029008 | 4.9551e-05 | 0.76137 | 433 | 0.75843
16:23 | 365 | uS | 183 | ppm | 22.7 | C | 385 | 362 | 0.0027624 | 0.046499 | 341.7 | µS |  |  | 437 | 371.7 | 402 | 267 | 346.9 | 0.0028827 | 0.002891 | 4.9384e-05 | 0.76578 | 434 | 0.7599
16:23 | 366 | uS | 183 | ppm | 22.7 | C | 386 | 365 | 0.0027397 | 0.046521 | 344.7 | µS |  |  | 438 | 369.7 | 403 | 267 | 346.5667 | 0.0028854 | 0.0028841 | 4.9265e-05 | 0.76505 | 435 | 0.76063
16:23 | 366 | uS | 183 | ppm | 22.7 | C | 387 | 366 | 0.0027322 | 0.046529 | 345.7 | µS |  |  | 439 | 371.7 | 402 | 268 | 347.2333 | 0.0028799 | 0.0028827 | 4.9242e-05 | 0.76652 | 436 | 0.76137
16:23 | 367 | uS | 184 | ppm | 22.7 | C | 388 | 366 | 0.0027322 | 0.046529 | 345.7 | µS |  |  | 440 | 371.7 | 403 | 268 | 347.5667 | 0.0028771 | 0.0028785 | 4.9171e-05 | 0.76726 | 437 | 0.76578
16:23 | 368 | uS | 184 | ppm | 22.7 | C | 389 | 367 | 0.0027248 | 0.046536 | 346.7 | µS |  |  | 441 | 371.7 | 401 | 270 | 347.5667 | 0.0028771 | 0.0028771 | 4.9147e-05 | 0.76726 | 438 | 0.76505
16:23 | 368 | uS | 184 | ppm | 22.7 | C | 390 | 368 | 0.0027174 | 0.046544 | 347.7 | µS |  |  | 442 | 372.7 | 404 | 270 | 348.9 | 0.0028662 | 0.0028716 | 4.9053e-05 | 0.7702 | 439 | 0.76652
16:23 | 368 | uS | 184 | ppm | 22.7 | C | 391 | 368 | 0.0027174 | 0.046544 | 347.7 | µS |  |  | 443 | 373.7 | 403 | 269 | 348.5667 | 0.0028689 | 0.0028675 | 4.8983e-05 | 0.76946 | 440 | 0.76726
16:23 | 370 | uS | 185 | ppm | 22.7 | C | 392 | 368 | 0.0027174 | 0.046544 | 347.7 | µS |  |  | 444 | 372.7 | 405 | 269 | 348.9 | 0.0028662 | 0.0028675 | 4.8983e-05 | 0.7702 | 441 | 0.76726
16:23 | 369 | uS | 185 | ppm | 22.7 | C | 393 | 370 | 0.0027027 | 0.046558 | 349.7 | µS |  |  | 445 | 373.7 | 405 | 270 | 349.5667 | 0.0028607 | 0.0028634 | 4.8913e-05 | 0.77167 | 442 | 0.7702
16:23 | 372 | uS | 186 | ppm | 22.7 | C | 394 | 369 | 0.00271 | 0.046551 | 348.7 | µS |  |  | 446 | 374.7 | 406 | 269 | 349.9 | 0.002858 | 0.0028593 | 4.8843e-05 | 0.77241 | 443 | 0.76946
16:23 | 370 | uS | 185 | ppm | 22.7 | C | 395 | 372 | 0.0026882 | 0.046573 | 351.7 | µS |  |  | 447 | 375.7 | 405 | 269 | 349.9 | 0.002858 | 0.002858 | 4.8819e-05 | 0.77241 | 444 | 0.7702
16:23 | 372 | uS | 186 | ppm | 22.7 | C | 396 | 370 | 0.0027027 | 0.046558 | 349.7 | µS |  |  | 448 | 377.7 | 405 | 273 | 351.9 | 0.0028417 | 0.0028498 | 4.8681e-05 | 0.77682 | 445 | 0.77167
16:23 | 373 | uS | 187 | ppm | 22.7 | C | 397 | 372 | 0.0026882 | 0.046573 | 351.7 | µS |  |  | 449 | 374.7 | 405 | 273 | 350.9 | 0.0028498 | 0.0028458 | 4.8611e-05 | 0.77461 | 446 | 0.77241
16:23 | 374 | uS | 187 | ppm | 22.7 | C | 398 | 373 | 0.002681 | 0.04658 | 352.7 | µS |  |  | 450 | 376.7 | 405 | 273 | 351.5667 | 0.0028444 | 0.0028471 | 4.8634e-05 | 0.77609 | 447 | 0.77241
16:23 | 374 | uS | 187 | ppm | 22.7 | C | 399 | 374 | 0.0026738 | 0.046587 | 353.7 | µS |  |  | 451 | 376.7 | 406 | 273 | 351.9 | 0.0028417 | 0.0028431 | 4.8565e-05 | 0.77682 | 448 | 0.77682
16:23 | 375 | uS | 188 | ppm | 22.7 | C | 400 | 374 | 0.0026738 | 0.046587 | 353.7 | µS |  |  | 452 | 376.7 | 406 | 275 | 352.5667 | 0.0028363 | 0.002839 | 4.8496e-05 | 0.77829 | 449 | 0.77461
16:23 | 375 | uS | 188 | ppm | 22.7 | C | 401 | 375 | 0.0026667 | 0.046594 | 354.7 | µS |  |  | 453 | 377.7 | 407 | 275 | 353.2333 | 0.002831 | 0.0028337 | 4.8404e-05 | 0.77976 | 450 | 0.77609
16:23 | 375 | uS | 188 | ppm | 22.7 | C | 402 | 375 | 0.0026667 | 0.046594 | 354.7 | µS |  |  | 454 | 377.7 | 407 | 277 | 353.9 | 0.0028257 | 0.0028283 | 4.8313e-05 | 0.78124 | 451 | 0.77682
16:23 | 375 | uS | 188 | ppm | 22.7 | C | 403 | 375 | 0.0026667 | 0.046594 | 354.7 | µS |  |  | 455 | 377.7 | 408 | 277 | 354.2333 | 0.002823 | 0.0028243 | 4.8245e-05 | 0.78197 | 452 | 0.77829
16:23 | 374 | uS | 187 | ppm | 22.7 | C | 404 | 375 | 0.0026667 | 0.046594 | 354.7 | µS |  |  | 456 | 379.7 | 406 | 278 | 354.5667 | 0.0028203 | 0.0028217 | 4.82e-05 | 0.78271 | 453 | 0.77976
16:23 | 376 | uS | 188 | ppm | 22.7 | C | 405 | 374 | 0.0026738 | 0.046587 | 353.7 | µS |  |  | 457 | 379.7 | 408 | 278 | 355.2333 | 0.0028151 | 0.0028177 | 4.8132e-05 | 0.78418 | 454 | 0.78124
16:23 | 375 | uS | 188 | ppm | 22.7 | C | 406 | 376 | 0.0026596 | 0.046602 | 355.7 | µS |  |  | 458 | 379.7 | 406 | 282 | 355.9 | 0.0028098 | 0.0028124 | 4.8041e-05 | 0.78565 | 455 | 0.78197
16:23 | 376 | uS | 188 | ppm | 22.7 | C | 407 | 375 | 0.0026667 | 0.046594 | 354.7 | µS |  |  | 459 | 380.7 | 408 | 282 | 356.9 | 0.0028019 | 0.0028058 | 4.7929e-05 | 0.78786 | 456 | 0.78271
16:23 | 377 | uS | 189 | ppm | 22.7 | C | 408 | 376 | 0.0026596 | 0.046602 | 355.7 | µS |  |  | 460 | 379.7 | 407 | 281 | 355.9 | 0.0028098 | 0.0028058 | 4.7929e-05 | 0.78565 | 457 | 0.78418
16:23 | 377 | uS | 189 | ppm | 22.7 | C | 409 | 377 | 0.0026525 | 0.046609 | 356.7 | µS |  |  | 461 | 381.7 | 408 | 281 | 356.9 | 0.0028019 | 0.0028058 | 4.7929e-05 | 0.78786 | 458 | 0.78565
16:23 | 377 | uS | 189 | ppm | 22.7 | C | 410 | 377 | 0.0026525 | 0.046609 | 356.7 | µS |  |  | 462 | 380.7 | 408 | 285 | 357.9 | 0.0027941 | 0.002798 | 4.7795e-05 | 0.79007 | 459 | 0.78786
16:23 | 379 | uS | 190 | ppm | 22.7 | C | 411 | 377 | 0.0026525 | 0.046609 | 356.7 | µS |  |  | 463 | 380.7 | 408 | 285 | 357.9 | 0.0027941 | 0.0027941 | 4.7728e-05 | 0.79007 | 460 | 0.78565
16:23 | 380 | uS | 190 | ppm | 22.7 | C | 412 | 379 | 0.0026385 | 0.046623 | 358.7 | µS |  |  | 464 | 381.7 | 409 | 286 | 358.9 | 0.0027863 | 0.0027902 | 4.7662e-05 | 0.79227 | 461 | 0.78786
16:23 | 379 | uS | 190 | ppm | 22.7 | C | 413 | 380 | 0.0026316 | 0.04663 | 359.7 | µS |  |  | 465 | 382.7 | 409 | 286 | 359.2333 | 0.0027837 | 0.002785 | 4.7573e-05 | 0.79301 | 462 | 0.79007
16:23 | 381 | uS | 191 | ppm | 22.7 | C | 414 | 379 | 0.0026385 | 0.046623 | 358.7 | µS |  |  | 466 | 382.7 | 410 | 288 | 360.2333 | 0.002776 | 0.0027798 | 4.7485e-05 | 0.79522 | 463 | 0.79007
16:23 | 378 | uS | 189 | ppm | 22.7 | C | 415 | 381 | 0.0026247 | 0.046636 | 360.7 | µS |  |  | 467 | 383.7 | 410 | 288 | 360.5667 | 0.0027734 | 0.0027747 | 4.7397e-05 | 0.79595 | 464 | 0.79227
16:23 | 381 | uS | 191 | ppm | 22.7 | C | 416 | 378 | 0.0026455 | 0.046616 | 357.7 | µS |  |  | 468 | 382.7 | 410 | 290 | 360.9 | 0.0027709 | 0.0027721 | 4.7353e-05 | 0.79669 | 465 | 0.79301
16:23 | 380 | uS | 190 | ppm | 22.7 | C | 417 | 381 | 0.0026247 | 0.046636 | 360.7 | µS |  |  | 469 | 382.7 | 411 | 290 | 361.2333 | 0.0027683 | 0.0027696 | 4.731e-05 | 0.79742 | 466 | 0.79522
16:23 | 381 | uS | 191 | ppm | 22.7 | C | 418 | 380 | 0.0026316 | 0.04663 | 359.7 | µS |  |  | 470 | 384.7 | 413 | 290 | 362.5667 | 0.0027581 | 0.0027632 | 4.7201e-05 | 0.80037 | 467 | 0.79595
16:23 | 382 | uS | 191 | ppm | 22.7 | C | 419 | 381 | 0.0026247 | 0.046636 | 360.7 | µS |  |  | 471 | 382.7 | 411 | 290 | 361.2333 | 0.0027683 | 0.0027632 | 4.7201e-05 | 0.79742 | 468 | 0.79669
16:23 | 383 | uS | 192 | ppm | 22.7 | C | 420 | 382 | 0.0026178 | 0.046643 | 361.7 | µS |  |  | 472 | 385.7 | 413 | 291 | 363.2333 | 0.0027531 | 0.0027607 | 4.7158e-05 | 0.80184 | 469 | 0.79742
16:23 | 382 | uS | 191 | ppm | 22.7 | C | 421 | 383 | 0.002611 | 0.04665 | 362.7 | µS |  |  | 473 | 383.7 | 411 | 291 | 361.9 | 0.0027632 | 0.0027581 | 4.7114e-05 | 0.7989 | 470 | 0.80037
16:23 | 383 | uS | 192 | ppm | 22.7 | C | 422 | 382 | 0.0026178 | 0.046643 | 361.7 | µS |  |  | 474 | 384.7 | 413 | 288 | 361.9 | 0.0027632 | 0.0027632 | 4.7201e-05 | 0.7989 | 471 | 0.79742
16:23 | 384 | uS | 192 | ppm | 22.7 | C | 423 | 383 | 0.002611 | 0.04665 | 362.7 | µS |  |  | 475 | 385.7 | 413 | 288 | 362.2333 | 0.0027607 | 0.0027619 | 4.7179e-05 | 0.79963 | 472 | 0.80184
16:23 | 384 | uS | 192 | ppm | 22.7 | C | 424 | 384 | 0.0026042 | 0.046657 | 363.7 | µS |  |  | 476 | 386.7 | 414 | 289 | 363.2333 | 0.0027531 | 0.0027569 | 4.7092e-05 | 0.80184 | 473 | 0.7989
16:23 | 386 | uS | 193 | ppm | 22.7 | C | 425 | 384 | 0.0026042 | 0.046657 | 363.7 | µS |  |  | 477 | 387.7 | 413 | 289 | 363.2333 | 0.0027531 | 0.0027531 | 4.7027e-05 | 0.80184 | 474 | 0.7989
16:23 | 384 | uS | 192 | ppm | 22.7 | C | 426 | 386 | 0.0025907 | 0.04667 | 365.7 | µS |  |  | 478 | 388.7 | 413 | 291 | 364.2333 | 0.0027455 | 0.0027493 | 4.6963e-05 | 0.80405 | 475 | 0.79963
16:23 | 387 | uS | 194 | ppm | 22.7 | C | 427 | 384 | 0.0026042 | 0.046657 | 363.7 | µS |  |  | 479 | 387.7 | 414 | 291 | 364.2333 | 0.0027455 | 0.0027455 | 4.6898e-05 | 0.80405 | 476 | 0.80184
16:23 | 387 | uS | 194 | ppm | 22.7 | C | 428 | 387 | 0.002584 | 0.046677 | 366.7 | µS |  |  | 480 | 388.7 | 414 | 292 | 364.9 | 0.0027405 | 0.002743 | 4.6855e-05 | 0.80552 | 477 | 0.80184
16:23 | 387 | uS | 194 | ppm | 22.7 | C | 429 | 387 | 0.002584 | 0.046677 | 366.7 | µS |  |  | 481 | 389.7 | 414 | 292 | 365.2333 | 0.002738 | 0.0027392 | 4.6791e-05 | 0.80625 | 478 | 0.80405
16:23 | 388 | uS | 194 | ppm | 22.7 | C | 430 | 387 | 0.002584 | 0.046677 | 366.7 | µS |  |  | 482 | 388.7 | 415 | 294 | 365.9 | 0.002733 | 0.0027355 | 4.6727e-05 | 0.80773 | 479 | 0.80405
16:23 | 389 | uS | 195 | ppm | 22.7 | C | 431 | 388 | 0.0025773 | 0.046684 | 367.7 | µS |  |  | 483 | 391.7 | 415 | 294 | 366.9 | 0.0027255 | 0.0027293 | 4.6621e-05 | 0.80993 | 480 | 0.80552
16:23 | 388 | uS | 194 | ppm | 22.7 | C | 432 | 389 | 0.0025707 | 0.04669 | 368.7 | µS |  |  | 484 | 388.7 | 415 | 296 | 366.5667 | 0.002728 | 0.0027268 | 4.6579e-05 | 0.8092 | 481 | 0.80625
16:23 | 389 | uS | 195 | ppm | 22.7 | C | 433 | 388 | 0.0025773 | 0.046684 | 367.7 | µS |  |  | 485 | 390.7 | 416 | 296 | 367.5667 | 0.0027206 | 0.0027243 | 4.6536e-05 | 0.81141 | 482 | 0.80773
16:23 | 390 | uS | 195 | ppm | 22.7 | C | 434 | 389 | 0.0025707 | 0.04669 | 368.7 | µS |  |  | 486 | 390.7 | 414 | 300 | 368.2333 | 0.0027157 | 0.0027181 | 4.6431e-05 | 0.81288 | 483 | 0.80993
16:23 | 390 | uS | 195 | ppm | 22.7 | C | 435 | 390 | 0.0025641 | 0.046697 | 369.7 | µS |  |  | 487 | 391.7 | 416 | 300 | 369.2333 | 0.0027083 | 0.002712 | 4.6326e-05 | 0.81508 | 484 | 0.8092
16:23 | 392 | uS | 196 | ppm | 22.7 | C | 436 | 390 | 0.0025641 | 0.046697 | 369.7 | µS |  |  | 488 | 391.7 | 415 | 298 | 368.2333 | 0.0027157 | 0.002712 | 4.6326e-05 | 0.81288 | 485 | 0.81141
16:23 | 390 | uS | 195 | ppm | 22.7 | C | 437 | 392 | 0.002551 | 0.04671 | 371.7 | µS |  |  | 489 | 391.7 | 416 | 298 | 368.5667 | 0.0027132 | 0.0027144 | 4.6368e-05 | 0.81361 | 486 | 0.81288
16:23 | 392 | uS | 196 | ppm | 22.7 | C | 438 | 390 | 0.0025641 | 0.046697 | 369.7 | µS |  |  | 490 | 391.7 | 416 | 301 | 369.5667 | 0.0027059 | 0.0027095 | 4.6284e-05 | 0.81582 | 487 | 0.81508
16:23 | 392 | uS | 196 | ppm | 22.7 | C | 439 | 392 | 0.002551 | 0.04671 | 371.7 | µS |  |  | 491 | 392.7 | 417 | 301 | 370.2333 | 0.002701 | 0.0027034 | 4.618e-05 | 0.81729 | 488 | 0.81288
16:23 | 392 | uS | 196 | ppm | 22.7 | C | 440 | 392 | 0.002551 | 0.04671 | 371.7 | µS |  |  | 492 | 392.7 | 417 | 301 | 370.2333 | 0.002701 | 0.002701 | 4.6138e-05 | 0.81729 | 489 | 0.81361
16:23 | 393 | uS | 197 | ppm | 22.7 | C | 441 | 392 | 0.002551 | 0.04671 | 371.7 | µS |  |  | 493 | 392.7 | 417 | 301 | 370.2333 | 0.002701 | 0.002701 | 4.6138e-05 | 0.81729 | 490 | 0.81582
16:23 | 394 | uS | 197 | ppm | 22.7 | C | 442 | 393 | 0.0025445 | 0.046717 | 372.7 | µS |  |  | 494 | 394.7 | 418 | 301 | 371.2333 | 0.0026937 | 0.0026974 | 4.6076e-05 | 0.8195 | 491 | 0.81729
16:23 | 393 | uS | 197 | ppm | 22.7 | C | 443 | 394 | 0.0025381 | 0.046723 | 373.7 | µS |  |  | 495 | 393.7 | 417 | 301 | 370.5667 | 0.0026986 | 0.0026961 | 4.6055e-05 | 0.81803 | 492 | 0.81729
16:24 | 394 | uS | 197 | ppm | 22.7 | C | 444 | 393 | 0.0025445 | 0.046717 | 372.7 | µS |  |  | 496 | 395.7 | 418 | 303 | 372.2333 | 0.0026865 | 0.0026925 | 4.5994e-05 | 0.82171 | 493 | 0.81729
16:24 | 395 | uS | 198 | ppm | 22.7 | C | 445 | 394 | 0.0025381 | 0.046723 | 373.7 | µS |  |  | 497 | 394.7 | 418 | 303 | 371.9 | 0.0026889 | 0.0026877 | 4.5911e-05 | 0.82097 | 494 | 0.8195
16:24 | 396 | uS | 198 | ppm | 22.7 | C | 446 | 395 | 0.0025316 | 0.046729 | 374.7 | µS |  |  | 498 | 395.7 | 419 | 303 | 372.5667 | 0.0026841 | 0.0026865 | 4.589e-05 | 0.82244 | 495 | 0.81803
16:24 | 398 | uS | 199 | ppm | 22.7 | C | 447 | 396 | 0.0025253 | 0.046736 | 375.7 | µS |  |  | 499 | 395.7 | 418 | 303 | 372.2333 | 0.0026865 | 0.0026853 | 4.587e-05 | 0.82171 | 496 | 0.82171
16:24 | 395 | uS | 198 | ppm | 22.7 | C | 448 | 398 | 0.0025126 | 0.046749 | 377.7 | µS |  |  | 500 | 395.7 | 420 | 304 | 373.2333 | 0.0026793 | 0.0026829 | 4.5829e-05 | 0.82391 | 497 | 0.82097
16:24 | 397 | uS | 199 | ppm | 22.7 | C | 449 | 395 | 0.0025316 | 0.046729 | 374.7 | µS |  |  | 501 | 396.7 | 417 | 304 | 372.5667 | 0.0026841 | 0.0026817 | 4.5808e-05 | 0.82244 | 498 | 0.82244
16:24 | 397 | uS | 199 | ppm | 22.7 | C | 450 | 397 | 0.0025189 | 0.046742 | 376.7 | µS |  |  | 502 | 397.7 | 419 | 305 | 373.9 | 0.0026745 | 0.0026793 | 4.5768e-05 | 0.82539 | 499 | 0.82171
16:24 | 397 | uS | 199 | ppm | 22.7 | C | 451 | 397 | 0.0025189 | 0.046742 | 376.7 | µS |  |  | 503 | 396.7 | 419 | 305 | 373.5667 | 0.0026769 | 0.0026757 | 4.5706e-05 | 0.82465 | 500 | 0.82391
16:24 | 398 | uS | 199 | ppm | 22.7 | C | 452 | 397 | 0.0025189 | 0.046742 | 376.7 | µS |  |  | 504 | 397.7 | 420 | 306 | 374.5667 | 0.0026698 | 0.0026733 | 4.5665e-05 | 0.82686 | 501 | 0.82244
16:24 | 398 | uS | 199 | ppm | 22.7 | C | 453 | 398 | 0.0025126 | 0.046749 | 377.7 | µS |  |  | 505 | 398.7 | 419 | 306 | 374.5667 | 0.0026698 | 0.0026698 | 4.5604e-05 | 0.82686 | 502 | 0.82539
16:24 | 398 | uS | 199 | ppm | 22.7 | C | 454 | 398 | 0.0025126 | 0.046749 | 377.7 | µS |  |  | 506 | 398.7 | 420 | 307 | 375.2333 | 0.002665 | 0.0026674 | 4.5564e-05 | 0.82833 | 503 | 0.82465
16:24 | 400 | uS | 200 | ppm | 22.7 | C | 455 | 398 | 0.0025126 | 0.046749 | 377.7 | µS |  |  | 507 | 400.7 | 421 | 307 | 376.2333 | 0.0026579 | 0.0026615 | 4.5463e-05 | 0.83054 | 504 | 0.82686
16:24 | 400 | uS | 200 | ppm | 22.7 | C | 456 | 400 | 0.0025 | 0.046761 | 379.7 | µS |  |  | 508 | 397.7 | 420 | 308 | 375.2333 | 0.002665 | 0.0026615 | 4.5463e-05 | 0.82833 | 505 | 0.82686
16:24 | 400 | uS | 200 | ppm | 22.7 | C | 457 | 400 | 0.0025 | 0.046761 | 379.7 | µS |  |  | 509 | 399.7 | 421 | 309 | 376.5667 | 0.0026556 | 0.0026603 | 4.5443e-05 | 0.83127 | 506 | 0.82833
16:24 | 401 | uS | 201 | ppm | 22.7 | C | 458 | 400 | 0.0025 | 0.046761 | 379.7 | µS |  |  | 510 | 398.7 | 421 | 309 | 376.2333 | 0.0026579 | 0.0026567 | 4.5382e-05 | 0.83054 | 507 | 0.83054
16:24 | 400 | uS | 200 | ppm | 22.7 | C | 459 | 401 | 0.0024938 | 0.046767 | 380.7 | µS |  |  | 511 | 399.7 | 421 | 307 | 375.9 | 0.0026603 | 0.0026591 | 4.5423e-05 | 0.8298 | 508 | 0.82833
16:24 | 402 | uS | 201 | ppm | 22.7 | C | 460 | 400 | 0.0025 | 0.046761 | 379.7 | µS |  |  | 512 | 400.7 | 421 | 311 | 377.5667 | 0.0026485 | 0.0026544 | 4.5342e-05 | 0.83348 | 509 | 0.83127
16:24 | 401 | uS | 201 | ppm | 22.7 | C | 461 | 402 | 0.0024876 | 0.046774 | 381.7 | µS |  |  | 513 | 401.7 | 422 | 311 | 378.2333 | 0.0026439 | 0.0026462 | 4.5202e-05 | 0.83495 | 510 | 0.83054
16:24 | 401 | uS | 201 | ppm | 22.7 | C | 462 | 401 | 0.0024938 | 0.046767 | 380.7 | µS |  |  | 514 | 400.7 | 420 | 311 | 377.2333 | 0.0026509 | 0.0026474 | 4.5222e-05 | 0.83274 | 511 | 0.8298
16:24 | 402 | uS | 201 | ppm | 22.7 | C | 463 | 401 | 0.0024938 | 0.046767 | 380.7 | µS |  |  | 515 | 401.7 | 422 | 312 | 378.5667 | 0.0026415 | 0.0026462 | 4.5202e-05 | 0.83569 | 512 | 0.83348
16:24 | 403 | uS | 202 | ppm | 22.7 | C | 464 | 402 | 0.0024876 | 0.046774 | 381.7 | µS |  |  | 516 | 400.7 | 421 | 312 | 377.9 | 0.0026462 | 0.0026439 | 4.5162e-05 | 0.83422 | 513 | 0.83495
16:24 | 403 | uS | 202 | ppm | 22.7 | C | 465 | 403 | 0.0024814 | 0.04678 | 382.7 | µS |  |  | 517 | 401.7 | 422 | 314 | 379.2333 | 0.0026369 | 0.0026416 | 4.5123e-05 | 0.83716 | 514 | 0.83274
16:24 | 404 | uS | 202 | ppm | 22.7 | C | 466 | 403 | 0.0024814 | 0.04678 | 382.7 | µS |  |  | 518 | 402.7 | 422 | 314 | 379.5667 | 0.0026346 | 0.0026357 | 4.5023e-05 | 0.8379 | 515 | 0.83569
16:24 | 403 | uS | 202 | ppm | 22.7 | C | 467 | 404 | 0.0024752 | 0.046786 | 383.7 | µS |  |  | 519 | 401.7 | 423 | 315 | 379.9 | 0.0026323 | 0.0026334 | 4.4984e-05 | 0.83863 | 516 | 0.83422
16:24 | 403 | uS | 202 | ppm | 22.7 | C | 468 | 403 | 0.0024814 | 0.04678 | 382.7 | µS |  |  | 520 | 403.7 | 423 | 315 | 380.5667 | 0.0026277 | 0.00263 | 4.4925e-05 | 0.8401 | 517 | 0.83716
16:24 | 405 | uS | 203 | ppm | 22.7 | C | 469 | 403 | 0.0024814 | 0.04678 | 382.7 | µS |  |  | 521 | 401.7 | 423 | 317 | 380.5667 | 0.0026277 | 0.0026277 | 4.4885e-05 | 0.8401 | 518 | 0.8379
16:24 | 403 | uS | 202 | ppm | 22.7 | C | 470 | 405 | 0.0024691 | 0.046792 | 384.7 | µS |  |  | 522 | 403.7 | 424 | 317 | 381.5667 | 0.0026208 | 0.0026242 | 4.4827e-05 | 0.84231 | 519 | 0.83863
16:24 | 406 | uS | 203 | ppm | 22.7 | C | 471 | 403 | 0.0024814 | 0.04678 | 382.7 | µS |  |  | 523 | 402.7 | 423 | 316 | 380.5667 | 0.0026277 | 0.0026242 | 4.4827e-05 | 0.8401 | 520 | 0.8401
16:24 | 404 | uS | 202 | ppm | 22.7 | C | 472 | 406 | 0.0024631 | 0.046798 | 385.7 | µS |  |  | 524 | 402.7 | 424 | 316 | 380.9 | 0.0026254 | 0.0026265 | 4.4866e-05 | 0.84084 | 521 | 0.8401
16:24 | 405 | uS | 203 | ppm | 22.7 | C | 473 | 404 | 0.0024752 | 0.046786 | 383.7 | µS |  |  | 525 | 403.7 | 424 | 317 | 381.5667 | 0.0026208 | 0.0026231 | 4.4807e-05 | 0.84231 | 522 | 0.84231
16:24 | 406 | uS | 203 | ppm | 22.7 | C | 474 | 405 | 0.0024691 | 0.046792 | 384.7 | µS |  |  | 526 | 404.7 | 425 | 317 | 382.2333 | 0.0026162 | 0.0026185 | 4.4729e-05 | 0.84378 | 523 | 0.8401
16:24 | 407 | uS | 204 | ppm | 22.7 | C | 475 | 406 | 0.0024631 | 0.046798 | 385.7 | µS |  |  | 527 | 403.7 | 423 | 317 | 381.2333 | 0.0026231 | 0.0026196 | 4.4748e-05 | 0.84157 | 524 | 0.84084
16:24 | 408 | uS | 204 | ppm | 22.7 | C | 476 | 407 | 0.002457 | 0.046804 | 386.7 | µS |  |  | 528 | 405.7 | 426 | 317 | 382.9 | 0.0026116 | 0.0026174 | 4.4709e-05 | 0.84525 | 525 | 0.84231
16:24 | 409 | uS | 205 | ppm | 22.7 | C | 477 | 408 | 0.002451 | 0.04681 | 387.7 | µS |  |  | 529 | 405.7 | 424 | 318 | 382.5667 | 0.0026139 | 0.0026128 | 4.4631e-05 | 0.84452 | 526 | 0.84378
16:24 | 408 | uS | 204 | ppm | 22.7 | C | 478 | 409 | 0.002445 | 0.046816 | 388.7 | µS |  |  | 530 | 405.7 | 425 | 318 | 382.9 | 0.0026116 | 0.0026128 | 4.4631e-05 | 0.84525 | 527 | 0.84157
16:24 | 409 | uS | 205 | ppm | 22.7 | C | 479 | 408 | 0.002451 | 0.04681 | 387.7 | µS |  |  | 531 | 406.7 | 425 | 320 | 383.9 | 0.0026048 | 0.0026082 | 4.4554e-05 | 0.84746 | 528 | 0.84525
16:24 | 410 | uS | 205 | ppm | 22.7 | C | 480 | 409 | 0.002445 | 0.046816 | 388.7 | µS |  |  | 532 | 406.7 | 426 | 320 | 384.2333 | 0.0026026 | 0.0026037 | 4.4476e-05 | 0.8482 | 529 | 0.84452
16:24 | 409 | uS | 205 | ppm | 22.7 | C | 481 | 410 | 0.002439 | 0.046822 | 389.7 | µS |  |  | 533 | 408.7 | 425 | 320 | 384.5667 | 0.0026003 | 0.0026015 | 4.4438e-05 | 0.84893 | 530 | 0.84525
16:24 | 412 | uS | 206 | ppm | 22.7 | C | 482 | 409 | 0.002445 | 0.046816 | 388.7 | µS |  |  | 534 | 405.7 | 426 | 320 | 383.9 | 0.0026048 | 0.0026026 | 4.4457e-05 | 0.84746 | 531 | 0.84746
16:24 | 409 | uS | 205 | ppm | 22.7 | C | 483 | 412 | 0.0024272 | 0.046834 | 391.7 | µS |  |  | 535 | 407.7 | 426 | 323 | 385.5667 | 0.0025936 | 0.0025992 | 4.44e-05 | 0.85114 | 532 | 0.8482
16:24 | 411 | uS | 206 | ppm | 22.7 | C | 484 | 409 | 0.002445 | 0.046816 | 388.7 | µS |  |  | 536 | 407.7 | 425 | 323 | 385.2333 | 0.0025958 | 0.0025947 | 4.4323e-05 | 0.8504 | 533 | 0.84893
16:24 | 411 | uS | 206 | ppm | 22.7 | C | 485 | 411 | 0.0024331 | 0.046828 | 390.7 | µS |  |  | 537 | 407.7 | 426 | 320 | 384.5667 | 0.0026003 | 0.0025981 | 4.438e-05 | 0.84893 | 534 | 0.84746
16:24 | 412 | uS | 206 | ppm | 22.7 | C | 486 | 411 | 0.0024331 | 0.046828 | 390.7 | µS |  |  | 538 | 408.7 | 426 | 320 | 384.9 | 0.0025981 | 0.0025992 | 4.4399e-05 | 0.84967 | 535 | 0.85114
16:24 | 412 | uS | 206 | ppm | 22.7 | C | 487 | 412 | 0.0024272 | 0.046834 | 391.7 | µS |  |  | 539 | 408.7 | 427 | 323 | 386.2333 | 0.0025891 | 0.0025936 | 4.4304e-05 | 0.85261 | 536 | 0.8504
16:24 | 412 | uS | 206 | ppm | 22.7 | C | 488 | 412 | 0.0024272 | 0.046834 | 391.7 | µS |  |  | 540 | 408.7 | 426 | 323 | 385.9 | 0.0025913 | 0.0025902 | 4.4246e-05 | 0.85188 | 537 | 0.84893
16:24 | 412 | uS | 206 | ppm | 22.7 | C | 489 | 412 | 0.0024272 | 0.046834 | 391.7 | µS |  |  | 541 | 409.7 | 428 | 323 | 386.9 | 0.0025846 | 0.002588 | 4.4208e-05 | 0.85408 | 538 | 0.84967
16:24 | 413 | uS | 207 | ppm | 22.7 | C | 490 | 412 | 0.0024272 | 0.046834 | 391.7 | µS |  |  | 542 | 409.7 | 425 | 323 | 385.9 | 0.0025913 | 0.002588 | 4.4208e-05 | 0.85188 | 539 | 0.85261
16:24 | 413 | uS | 207 | ppm | 22.7 | C | 491 | 413 | 0.0024213 | 0.04684 | 392.7 | µS |  |  | 543 | 409.7 | 427 | 324 | 386.9 | 0.0025846 | 0.002588 | 4.4208e-05 | 0.85408 | 540 | 0.85188
16:24 | 413 | uS | 207 | ppm | 22.7 | C | 492 | 413 | 0.0024213 | 0.04684 | 392.7 | µS |  |  | 544 | 409.7 | 426 | 324 | 386.5667 | 0.0025869 | 0.0025858 | 4.417e-05 | 0.85335 | 541 | 0.85408
16:24 | 415 | uS | 208 | ppm | 22.7 | C | 493 | 413 | 0.0024213 | 0.04684 | 392.7 | µS |  |  | 545 | 409.7 | 427 | 327 | 387.9 | 0.002578 | 0.0025824 | 4.4113e-05 | 0.85629 | 542 | 0.85188
16:24 | 414 | uS | 207 | ppm | 22.7 | C | 494 | 415 | 0.0024096 | 0.046851 | 394.7 | µS |  |  | 546 | 411.7 | 427 | 327 | 388.5667 | 0.0025736 | 0.0025758 | 4.3999e-05 | 0.85776 | 543 | 0.85408
16:24 | 416 | uS | 208 | ppm | 22.7 | C | 495 | 414 | 0.0024155 | 0.046846 | 393.7 | µS |  |  | 547 | 408.7 | 428 | 329 | 388.5667 | 0.0025736 | 0.0025736 | 4.3961e-05 | 0.85776 | 544 | 0.85335
16:24 | 415 | uS | 208 | ppm | 22.7 | C | 496 | 416 | 0.0024038 | 0.046857 | 395.7 | µS |  |  | 548 | 411.7 | 429 | 329 | 389.9 | 0.0025648 | 0.0025692 | 4.3886e-05 | 0.86071 | 545 | 0.85629
16:24 | 416 | uS | 208 | ppm | 22.7 | C | 497 | 415 | 0.0024096 | 0.046851 | 394.7 | µS |  |  | 549 | 410.7 | 428 | 332 | 390.2333 | 0.0025626 | 0.0025637 | 4.3792e-05 | 0.86144 | 546 | 0.85776
16:24 | 416 | uS | 208 | ppm | 22.7 | C | 498 | 416 | 0.0024038 | 0.046857 | 395.7 | µS |  |  | 550 | 410.7 | 429 | 332 | 390.5667 | 0.0025604 | 0.0025615 | 4.3755e-05 | 0.86218 | 547 | 0.85776
16:24 | 416 | uS | 208 | ppm | 22.7 | C | 499 | 416 | 0.0024038 | 0.046857 | 395.7 | µS |  |  | 551 | 411.7 | 429 | 332 | 390.9 | 0.0025582 | 0.0025593 | 4.3718e-05 | 0.86291 | 548 | 0.86071
16:24 | 417 | uS | 209 | ppm | 22.7 | C | 500 | 416 | 0.0024038 | 0.046857 | 395.7 | µS |  |  | 552 | 411.7 | 429 | 332 | 390.9 | 0.0025582 | 0.0025582 | 4.3699e-05 | 0.86291 | 549 | 0.86144
16:24 | 418 | uS | 209 | ppm | 22.7 | C | 501 | 417 | 0.0023981 | 0.046863 | 396.7 | µS |  |  | 553 | 411.7 | 428 | 334 | 391.2333 | 0.002556 | 0.0025571 | 4.368e-05 | 0.86365 | 550 | 0.86218
16:24 | 417 | uS | 209 | ppm | 22.7 | C | 502 | 418 | 0.0023923 | 0.046869 | 397.7 | µS |  |  | 554 | 412.7 | 430 | 334 | 392.2333 | 0.0025495 | 0.0025528 | 4.3606e-05 | 0.86586 | 551 | 0.86291
16:24 | 418 | uS | 209 | ppm | 22.7 | C | 503 | 417 | 0.0023981 | 0.046863 | 396.7 | µS |  |  | 555 | 411.7 | 428 | 334 | 391.2333 | 0.002556 | 0.0025528 | 4.3606e-05 | 0.86365 | 552 | 0.86291
16:25 | 419 | uS | 210 | ppm | 22.7 | C | 504 | 418 | 0.0023923 | 0.046869 | 397.7 | µS |  |  | 556 | 412.7 | 429 | 334 | 391.9 | 0.0025517 | 0.0025538 | 4.3625e-05 | 0.86512 | 553 | 0.86365
16:25 | 419 | uS | 210 | ppm | 22.7 | C | 505 | 419 | 0.0023866 | 0.046874 | 398.7 | µS |  |  | 557 | 412.7 | 428 | 334 | 391.5667 | 0.0025538 | 0.0025528 | 4.3606e-05 | 0.86439 | 554 | 0.86586
16:25 | 421 | uS | 211 | ppm | 22.7 | C | 506 | 419 | 0.0023866 | 0.046874 | 398.7 | µS |  |  | 558 | 412.7 | 430 | 334 | 392.2333 | 0.0025495 | 0.0025517 | 4.3587e-05 | 0.86586 | 555 | 0.86365
16:25 | 418 | uS | 209 | ppm | 22.7 | C | 507 | 421 | 0.0023753 | 0.046886 | 400.7 | µS |  |  | 559 | 413.7 | 429 | 336 | 392.9 | 0.0025452 | 0.0025473 | 4.3513e-05 | 0.86733 | 556 | 0.86512
16:25 | 420 | uS | 210 | ppm | 22.7 | C | 508 | 418 | 0.0023923 | 0.046869 | 397.7 | µS |  |  | 560 | 411.7 | 430 | 336 | 392.5667 | 0.0025473 | 0.0025463 | 4.3495e-05 | 0.86659 | 557 | 0.86439
16:25 | 419 | uS | 210 | ppm | 22.7 | C | 509 | 420 | 0.002381 | 0.04688 | 399.7 | µS |  |  | 561 | 413.7 | 431 | 334 | 392.9 | 0.0025452 | 0.0025463 | 4.3495e-05 | 0.86733 | 558 | 0.86586
16:25 | 420 | uS | 210 | ppm | 22.7 | C | 510 | 419 | 0.0023866 | 0.046874 | 398.7 | µS |  |  | 562 | 411.7 | 430 | 334 | 391.9 | 0.0025517 | 0.0025484 | 4.3532e-05 | 0.86512 | 559 | 0.86733
16:25 | 421 | uS | 211 | ppm | 22.7 | C | 511 | 420 | 0.002381 | 0.04688 | 399.7 | µS |  |  | 563 | 413.7 | 430 | 338 | 393.9 | 0.0025387 | 0.0025452 | 4.3477e-05 | 0.86954 | 560 | 0.86659
16:25 | 422 | uS | 211 | ppm | 22.7 | C | 512 | 421 | 0.0023753 | 0.046886 | 400.7 | µS |  |  | 564 | 413.7 | 431 | 338 | 394.2333 | 0.0025366 | 0.0025376 | 4.3348e-05 | 0.87027 | 561 | 0.86733
16:25 | 421 | uS | 211 | ppm | 22.7 | C | 513 | 422 | 0.0023697 | 0.046891 | 401.7 | µS |  |  | 565 | 413.7 | 431 | 338 | 394.2333 | 0.0025366 | 0.0025366 | 4.3329e-05 | 0.87027 | 562 | 0.86512
16:25 | 422 | uS | 211 | ppm | 22.7 | C | 514 | 421 | 0.0023753 | 0.046886 | 400.7 | µS |  |  | 566 | 414.7 | 430 | 338 | 394.2333 | 0.0025366 | 0.0025366 | 4.3329e-05 | 0.87027 | 563 | 0.86954
16:25 | 421 | uS | 211 | ppm | 22.7 | C | 515 | 422 | 0.0023697 | 0.046891 | 401.7 | µS |  |  | 567 | 414.7 | 432 | 338 | 394.9 | 0.0025323 | 0.0025344 | 4.3293e-05 | 0.87174 | 564 | 0.87027
16:25 | 422 | uS | 211 | ppm | 22.7 | C | 516 | 421 | 0.0023753 | 0.046886 | 400.7 | µS |  |  | 568 | 413.7 | 429 | 338 | 393.5667 | 0.0025409 | 0.0025366 | 4.333e-05 | 0.8688 | 565 | 0.87027
16:25 | 423 | uS | 212 | ppm | 22.7 | C | 517 | 422 | 0.0023697 | 0.046891 | 401.7 | µS |  |  | 569 | 414.7 | 431 | 338 | 394.5667 | 0.0025344 | 0.0025376 | 4.3348e-05 | 0.87101 | 566 | 0.87027
16:25 | 422 | uS | 211 | ppm | 22.7 | C | 518 | 423 | 0.0023641 | 0.046897 | 402.7 | µS |  |  | 570 | 414.7 | 431 | 338 | 394.5667 | 0.0025344 | 0.0025344 | 4.3293e-05 | 0.87101 | 567 | 0.87174
16:25 | 424 | uS | 212 | ppm | 22.7 | C | 519 | 422 | 0.0023697 | 0.046891 | 401.7 | µS |  |  | 571 | 414.7 | 431 | 340 | 395.2333 | 0.0025302 | 0.0025323 | 4.3256e-05 | 0.87248 | 568 | 0.8688
16:25 | 422 | uS | 211 | ppm | 22.7 | C | 520 | 424 | 0.0023585 | 0.046903 | 403.7 | µS |  |  | 572 | 415.7 | 431 | 340 | 395.5667 | 0.002528 | 0.0025291 | 4.3202e-05 | 0.87322 | 569 | 0.87101
16:25 | 424 | uS | 212 | ppm | 22.7 | C | 521 | 422 | 0.0023697 | 0.046891 | 401.7 | µS |  |  | 573 | 414.7 | 432 | 343 | 396.5667 | 0.0025216 | 0.0025248 | 4.3129e-05 | 0.87542 | 570 | 0.87101
16:25 | 423 | uS | 212 | ppm | 22.7 | C | 522 | 424 | 0.0023585 | 0.046903 | 403.7 | µS |  |  | 574 | 416.7 | 433 | 343 | 397.5667 | 0.0025153 | 0.0025185 | 4.302e-05 | 0.87763 | 571 | 0.87248
16:25 | 423 | uS | 212 | ppm | 22.7 | C | 523 | 423 | 0.0023641 | 0.046897 | 402.7 | µS |  |  | 575 | 413.7 | 432 | 341 | 395.5667 | 0.002528 | 0.0025217 | 4.3075e-05 | 0.87322 | 572 | 0.87322
16:25 | 424 | uS | 212 | ppm | 22.7 | C | 524 | 423 | 0.0023641 | 0.046897 | 402.7 | µS |  |  | 576 | 415.7 | 433 | 341 | 396.5667 | 0.0025216 | 0.0025248 | 4.3129e-05 | 0.87542 | 573 | 0.87542
16:25 | 425 | uS | 213 | ppm | 22.7 | C | 525 | 424 | 0.0023585 | 0.046903 | 403.7 | µS |  |  | 577 | 414.7 | 433 | 342 | 396.5667 | 0.0025216 | 0.0025216 | 4.3074e-05 | 0.87542 | 574 | 0.87763
16:25 | 424 | uS | 212 | ppm | 22.7 | C | 526 | 425 | 0.0023529 | 0.046908 | 404.7 | µS |  |  | 578 | 415.7 | 434 | 342 | 397.2333 | 0.0025174 | 0.0025195 | 4.3038e-05 | 0.87689 | 575 | 0.87322
16:25 | 426 | uS | 213 | ppm | 22.7 | C | 527 | 424 | 0.0023585 | 0.046903 | 403.7 | µS |  |  | 579 | 415.7 | 432 | 343 | 396.9 | 0.0025195 | 0.0025185 | 4.302e-05 | 0.87616 | 576 | 0.87542
16:25 | 426 | uS | 213 | ppm | 22.7 | C | 528 | 426 | 0.0023474 | 0.046914 | 405.7 | µS |  |  | 580 | 416.7 | 434 | 343 | 397.9 | 0.0025132 | 0.0025164 | 4.2984e-05 | 0.87837 | 577 | 0.87542
16:25 | 426 | uS | 213 | ppm | 22.7 | C | 529 | 426 | 0.0023474 | 0.046914 | 405.7 | µS |  |  | 581 | 416.7 | 432 | 344 | 397.5667 | 0.0025153 | 0.0025142 | 4.2948e-05 | 0.87763 | 578 | 0.87689
16:25 | 427 | uS | 214 | ppm | 22.7 | C | 530 | 426 | 0.0023474 | 0.046914 | 405.7 | µS |  |  | 582 | 417.7 | 434 | 344 | 398.5667 | 0.002509 | 0.0025121 | 4.2912e-05 | 0.87984 | 579 | 0.87616
16:25 | 427 | uS | 214 | ppm | 22.7 | C | 531 | 427 | 0.0023419 | 0.046919 | 406.7 | µS |  |  | 583 | 416.7 | 433 | 345 | 398.2333 | 0.0025111 | 0.00251 | 4.2876e-05 | 0.8791 | 580 | 0.87837
16:25 | 429 | uS | 215 | ppm | 22.7 | C | 532 | 427 | 0.0023419 | 0.046919 | 406.7 | µS |  |  | 584 | 417.7 | 434 | 345 | 398.9 | 0.0025069 | 0.002509 | 4.2858e-05 | 0.88057 | 581 | 0.87763
16:25 | 426 | uS | 213 | ppm | 22.7 | C | 533 | 429 | 0.002331 | 0.04693 | 408.7 | µS |  |  | 585 | 417.7 | 434 | 345 | 398.9 | 0.0025069 | 0.0025069 | 4.2823e-05 | 0.88057 | 582 | 0.87984
16:25 | 428 | uS | 214 | ppm | 22.7 | C | 534 | 426 | 0.0023474 | 0.046914 | 405.7 | µS |  |  | 586 | 417.7 | 434 | 345 | 398.9 | 0.0025069 | 0.0025069 | 4.2823e-05 | 0.88057 | 583 | 0.8791
16:25 | 428 | uS | 214 | ppm | 22.7 | C | 535 | 428 | 0.0023364 | 0.046925 | 407.7 | µS |  |  | 587 | 418.7 | 435 | 349 | 400.9 | 0.0024944 | 0.0025006 | 4.2716e-05 | 0.88499 | 584 | 0.88057
16:25 | 428 | uS | 214 | ppm | 22.7 | C | 536 | 428 | 0.0023364 | 0.046925 | 407.7 | µS |  |  | 588 | 417.7 | 435 | 349 | 400.5667 | 0.0024965 | 0.0024954 | 4.2627e-05 | 0.88425 | 585 | 0.88057
16:25 | 429 | uS | 215 | ppm | 22.7 | C | 537 | 428 | 0.0023364 | 0.046925 | 407.7 | µS |  |  | 589 | 419.7 | 435 | 348 | 400.9 | 0.0024944 | 0.0024954 | 4.2627e-05 | 0.88499 | 586 | 0.88057
16:25 | 429 | uS | 215 | ppm | 22.7 | C | 538 | 429 | 0.002331 | 0.04693 | 408.7 | µS |  |  | 590 | 417.7 | 435 | 348 | 400.2333 | 0.0024985 | 0.0024965 | 4.2644e-05 | 0.88352 | 587 | 0.88499
16:25 | 429 | uS | 215 | ppm | 22.8 | C | 539 | 429 | 0.002331 | 0.04693 | 408.7 | µS |  |  | 591 | 419.7 | 436 | 349 | 401.5667 | 0.0024902 | 0.0024944 | 4.2609e-05 | 0.88646 | 588 | 0.88425
16:25 | 430 | uS | 215 | ppm | 22.8 | C | 540 | 429 | 0.002331 | 0.04693 | 408.7 | µS |  |  | 592 | 418.7 | 434 | 349 | 400.5667 | 0.0024965 | 0.0024934 | 4.2591e-05 | 0.88425 | 589 | 0.88499
16:25 | 430 | uS | 215 | ppm | 22.8 | C | 541 | 430 | 0.0023256 | 0.046936 | 409.7 | µS |  |  | 593 | 418.7 | 436 | 350 | 401.5667 | 0.0024902 | 0.0024934 | 4.2591e-05 | 0.88646 | 590 | 0.88352
16:25 | 430 | uS | 215 | ppm | 22.8 | C | 542 | 430 | 0.0023256 | 0.046936 | 409.7 | µS |  |  | 594 | 419.7 | 435 | 350 | 401.5667 | 0.0024902 | 0.0024902 | 4.2538e-05 | 0.88646 | 591 | 0.88646
16:25 | 430 | uS | 215 | ppm | 22.8 | C | 543 | 430 | 0.0023256 | 0.046936 | 409.7 | µS |  |  | 595 | 419.7 | 436 | 350 | 401.9 | 0.0024882 | 0.0024892 | 4.2521e-05 | 0.8872 | 592 | 0.88425
16:25 | 430 | uS | 215 | ppm | 22.8 | C | 544 | 430 | 0.0023256 | 0.046936 | 409.7 | µS |  |  | 596 | 418.7 | 436 | 350 | 401.5667 | 0.0024902 | 0.0024892 | 4.2521e-05 | 0.88646 | 593 | 0.88646
16:25 | 432 | uS | 216 | ppm | 22.8 | C | 545 | 430 | 0.0023256 | 0.046936 | 409.7 | µS |  |  | 597 | 419.7 | 436 | 350 | 401.9 | 0.0024882 | 0.0024892 | 4.2521e-05 | 0.8872 | 594 | 0.88646
16:25 | 429 | uS | 215 | ppm | 22.8 | C | 546 | 432 | 0.0023148 | 0.046946 | 411.7 | µS |  |  | 598 | 418.7 | 436 | 350 | 401.5667 | 0.0024902 | 0.0024892 | 4.2521e-05 | 0.88646 | 595 | 0.8872
16:25 | 432 | uS | 216 | ppm | 22.8 | C | 547 | 429 | 0.002331 | 0.04693 | 408.7 | µS |  |  | 599 | 419.7 | 435 | 349 | 401.2333 | 0.0024923 | 0.0024913 | 4.2556e-05 | 0.88572 | 596 | 0.88646
16:25 | 431 | uS | 216 | ppm | 22.8 | C | 548 | 432 | 0.0023148 | 0.046946 | 411.7 | µS |  |  | 600 | 419.7 | 436 | 349 | 401.5667 | 0.0024902 | 0.0024913 | 4.2556e-05 | 0.88646 | 597 | 0.8872
16:25 | 431 | uS | 216 | ppm | 22.8 | C | 549 | 431 | 0.0023202 | 0.046941 | 410.7 | µS |  |  | 601 | 419.7 | 436 | 352 | 402.5667 | 0.0024841 | 0.0024872 | 4.2485e-05 | 0.88867 | 598 | 0.88646
16:25 | 432 | uS | 216 | ppm | 22.8 | C | 550 | 431 | 0.0023202 | 0.046941 | 410.7 | µS |  |  | 602 | 420.7 | 437 | 352 | 403.2333 | 0.00248 | 0.002482 | 4.2397e-05 | 0.89014 | 599 | 0.88572
16:25 | 432 | uS | 216 | ppm | 22.8 | C | 551 | 432 | 0.0023148 | 0.046946 | 411.7 | µS |  |  | 603 | 419.7 | 436 | 353 | 402.9 | 0.002482 | 0.002481 | 4.238e-05 | 0.8894 | 600 | 0.88646
16:25 | 432 | uS | 216 | ppm | 22.8 | C | 552 | 432 | 0.0023148 | 0.046946 | 411.7 | µS |  |  | 604 | 421.7 | 438 | 353 | 404.2333 | 0.0024738 | 0.0024779 | 4.2327e-05 | 0.89235 | 601 | 0.88867
16:25 | 433 | uS | 217 | ppm | 22.8 | C | 553 | 432 | 0.0023148 | 0.046946 | 411.7 | µS |  |  | 605 | 418.7 | 435 | 354 | 402.5667 | 0.0024841 | 0.0024789 | 4.2345e-05 | 0.88867 | 602 | 0.89014
16:25 | 432 | uS | 216 | ppm | 22.8 | C | 554 | 433 | 0.0023095 | 0.046952 | 412.7 | µS |  |  | 606 | 420.7 | 437 | 354 | 403.9 | 0.0024759 | 0.00248 | 4.2362e-05 | 0.89161 | 603 | 0.8894
16:25 | 433 | uS | 217 | ppm | 22.8 | C | 555 | 432 | 0.0023148 | 0.046946 | 411.7 | µS |  |  | 607 | 419.7 | 437 | 355 | 403.9 | 0.0024759 | 0.0024759 | 4.2292e-05 | 0.89161 | 604 | 0.89235
16:25 | 433 | uS | 217 | ppm | 22.8 | C | 556 | 433 | 0.0023095 | 0.046952 | 412.7 | µS |  |  | 608 | 420.7 | 438 | 355 | 404.5667 | 0.0024718 | 0.0024738 | 4.2258e-05 | 0.89308 | 605 | 0.88867
16:25 | 433 | uS | 217 | ppm | 22.8 | C | 557 | 433 | 0.0023095 | 0.046952 | 412.7 | µS |  |  | 609 | 421.7 | 437 | 354 | 404.2333 | 0.0024738 | 0.0024728 | 4.224e-05 | 0.89235 | 606 | 0.89161
16:25 | 434 | uS | 217 | ppm | 22.8 | C | 558 | 433 | 0.0023095 | 0.046952 | 412.7 | µS |  |  | 610 | 421.7 | 438 | 354 | 404.5667 | 0.0024718 | 0.0024728 | 4.224e-05 | 0.89308 | 607 | 0.89161
16:25 | 432 | uS | 216 | ppm | 22.8 | C | 559 | 434 | 0.0023041 | 0.046957 | 413.7 | µS |  |  | 611 | 421.7 | 438 | 355 | 404.9 | 0.0024697 | 0.0024708 | 4.2205e-05 | 0.89382 | 608 | 0.89308
16:25 | 434 | uS | 217 | ppm | 22.8 | C | 560 | 432 | 0.0023148 | 0.046946 | 411.7 | µS |  |  | 612 | 421.7 | 438 | 355 | 404.9 | 0.0024697 | 0.0024697 | 4.2188e-05 | 0.89382 | 609 | 0.89235
16:25 | 432 | uS | 216 | ppm | 22.8 | C | 561 | 434 | 0.0023041 | 0.046957 | 413.7 | µS |  |  | 613 | 420.7 | 438 | 352 | 403.5667 | 0.0024779 | 0.0024738 | 4.2258e-05 | 0.89088 | 610 | 0.89308
16:25 | 434 | uS | 217 | ppm | 22.8 | C | 562 | 432 | 0.0023148 | 0.046946 | 411.7 | µS |  |  | 614 | 421.7 | 438 | 352 | 403.9 | 0.0024759 | 0.0024769 | 4.231e-05 | 0.89161 | 611 | 0.89382
16:25 | 434 | uS | 217 | ppm | 22.8 | C | 563 | 434 | 0.0023041 | 0.046957 | 413.7 | µS |  |  | 615 | 421.7 | 439 | 355 | 405.2333 | 0.0024677 | 0.0024718 | 4.2223e-05 | 0.89455 | 612 | 0.89382
16:26 | 434 | uS | 217 | ppm | 22.8 | C | 564 | 434 | 0.0023041 | 0.046957 | 413.7 | µS |  |  | 616 | 421.7 | 437 | 355 | 404.5667 | 0.0024718 | 0.0024697 | 4.2188e-05 | 0.89308 | 613 | 0.89088
16:26 | 435 | uS | 218 | ppm | 22.8 | C | 565 | 434 | 0.0023041 | 0.046957 | 413.7 | µS |  |  | 617 | 422.7 | 439 | 357 | 406.2333 | 0.0024616 | 0.0024667 | 4.2136e-05 | 0.89676 | 614 | 0.89161
16:26 | 435 | uS | 218 | ppm | 22.8 | C | 566 | 435 | 0.0022989 | 0.046962 | 414.7 | µS |  |  | 618 | 421.7 | 437 | 357 | 405.2333 | 0.0024677 | 0.0024647 | 4.2101e-05 | 0.89455 | 615 | 0.89455
16:26 | 434 | uS | 217 | ppm | 22.8 | C | 567 | 435 | 0.0022989 | 0.046962 | 414.7 | µS |  |  | 619 | 423.7 | 438 | 358 | 406.5667 | 0.0024596 | 0.0024637 | 4.2084e-05 | 0.8975 | 616 | 0.89308
16:26 | 435 | uS | 218 | ppm | 22.8 | C | 568 | 434 | 0.0023041 | 0.046957 | 413.7 | µS |  |  | 620 | 420.7 | 438 | 358 | 405.5667 | 0.0024657 | 0.0024627 | 4.2067e-05 | 0.89529 | 617 | 0.89676
16:26 | 435 | uS | 218 | ppm | 22.8 | C | 569 | 435 | 0.0022989 | 0.046962 | 414.7 | µS |  |  | 621 | 423.7 | 439 | 359 | 407.2333 | 0.0024556 | 0.0024606 | 4.2032e-05 | 0.89897 | 618 | 0.89455
16:26 | 435 | uS | 218 | ppm | 22.8 | C | 570 | 435 | 0.0022989 | 0.046962 | 414.7 | µS |  |  | 622 | 421.7 | 439 | 359 | 406.5667 | 0.0024596 | 0.0024576 | 4.1981e-05 | 0.8975 | 619 | 0.8975
16:26 | 436 | uS | 218 | ppm | 22.8 | C | 571 | 435 | 0.0022989 | 0.046962 | 414.7 | µS |  |  | 623 | 422.7 | 438 | 361 | 407.2333 | 0.0024556 | 0.0024576 | 4.1981e-05 | 0.89897 | 620 | 0.89529
16:26 | 435 | uS | 218 | ppm | 22.8 | C | 572 | 436 | 0.0022936 | 0.046968 | 415.7 | µS | Start |  | 624 | 422.7 | 439 | 361 | 407.5667 | 0.0024536 | 0.0024546 | 4.1929e-05 | 0.89971 | 621 | 0.89897
16:26 | 437 | uS | 219 | ppm | 22.8 | C | 573 | 435 | 0.0022989 | 0.046962 | 414.7 | µS |  |  | 625 | 422.7 | 439 | 363 | 408.2333 | 0.0024496 | 0.0024516 | 4.1878e-05 | 0.90118 | 622 | 0.8975
16:26 | 434 | uS | 217 | ppm | 22.8 | C | 574 | 437 | 0.0022883 | 0.046973 | 416.7 | µS | 575.5 |  | 626 | 423.7 | 440 | 363 | 408.9 | 0.0024456 | 0.0024476 | 4.1809e-05 | 0.90265 | 623 | 0.89897
16:26 | 436 | uS | 218 | ppm | 22.8 | C | 575 | 434 | 0.0023041 | 0.046957 | 413.7 | µS |  |  | 627 | 423.7 | 439 | 362 | 408.2333 | 0.0024496 | 0.0024476 | 4.1809e-05 | 0.90118 | 624 | 0.89971
16:26 | 435 | uS | 218 | ppm | 22.8 | C | 576 | 436 | 0.0022936 | 0.046968 | 415.7 | µS |  |  | 628 | 422.7 | 441 | 362 | 408.5667 | 0.0024476 | 0.0024486 | 4.1826e-05 | 0.90191 | 625 | 0.90118
16:26 | 436 | uS | 218 | ppm | 22.8 | C | 577 | 435 | 0.0022989 | 0.046962 | 414.7 | µS |  |  | 629 | 423.7 | 438 | 363 | 408.2333 | 0.0024496 | 0.0024486 | 4.1826e-05 | 0.90118 | 626 | 0.90265
16:26 | 436 | uS | 218 | ppm | 22.8 | C | 578 | 436 | 0.0022936 | 0.046968 | 415.7 | µS |  |  | 630 | 423.7 | 440 | 363 | 408.9 | 0.0024456 | 0.0024476 | 4.1809e-05 | 0.90265 | 627 | 0.90118
16:26 | 437 | uS | 219 | ppm | 22.8 | C | 579 | 436 | 0.0022936 | 0.046968 | 415.7 | µS | End |  | 631 | 423.7 | 439 | 363 | 408.5667 | 0.0024476 | 0.0024466 | 4.1792e-05 | 0.90191 | 628 | 0.90191
16:26 | 437 | uS | 219 | ppm | 22.8 | C | 580 | 437 | 0.0022883 | 0.046973 | 416.7 | µS |  |  | 632 | 424.7 | 440 | 363 | 409.2333 | 0.0024436 | 0.0024456 | 4.1775e-05 | 0.90338 | 629 | 0.90118
16:26 | 438 | uS | 219 | ppm | 22.8 | C | 581 | 437 | 0.0022883 | 0.046973 | 416.7 | µS |  |  | 633 | 423.7 | 440 | 363 | 408.9 | 0.0024456 | 0.0024446 | 4.1758e-05 | 0.90265 | 630 | 0.90265
16:26 | 437 | uS | 219 | ppm | 22.8 | C | 582 | 438 | 0.0022831 | 0.046978 | 417.7 | µS |  |  | 634 | 425.7 | 440 | 363 | 409.5667 | 0.0024416 | 0.0024436 | 4.1741e-05 | 0.90412 | 631 | 0.90191
16:26 | 438 | uS | 219 | ppm | 22.8 | C | 583 | 437 | 0.0022883 | 0.046973 | 416.7 | µS |  |  | 635 | 422.7 | 441 | 365 | 409.5667 | 0.0024416 | 0.0024416 | 4.1707e-05 | 0.90412 | 632 | 0.90338
16:26 | 438 | uS | 219 | ppm | 22.8 | C | 584 | 438 | 0.0022831 | 0.046978 | 417.7 | µS |  |  | 636 | 425.7 | 440 | 364 | 409.9 | 0.0024396 | 0.0024406 | 4.169e-05 | 0.90486 | 633 | 0.90265
16:26 | 438 | uS | 219 | ppm | 22.8 | C | 585 | 438 | 0.0022831 | 0.046978 | 417.7 | µS |  |  | 637 | 423.7 | 440 | 368 | 410.5667 | 0.0024357 | 0.0024376 | 4.164e-05 | 0.90633 | 634 | 0.90412
16:26 | 439 | uS | 220 | ppm | 22.8 | C | 586 | 438 | 0.0022831 | 0.046978 | 417.7 | µS |  |  | 638 | 424.7 | 441 | 368 | 411.2333 | 0.0024317 | 0.0024337 | 4.1572e-05 | 0.9078 | 635 | 0.90412
16:26 | 438 | uS | 219 | ppm | 22.8 | C | 587 | 439 | 0.0022779 | 0.046983 | 418.7 | µS |  |  | 639 | 424.7 | 442 | 368 | 411.5667 | 0.0024297 | 0.0024307 | 4.1521e-05 | 0.90854 | 636 | 0.90486
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 588 | 438 | 0.0022831 | 0.046978 | 417.7 | µS |  |  | 640 | 424.7 | 439 | 370 | 411.2333 | 0.0024317 | 0.0024307 | 4.1521e-05 | 0.9078 | 637 | 0.90633
16:26 | 438 | uS | 219 | ppm | 22.8 | C | 589 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 641 | 425.7 | 441 | 371 | 412.5667 | 0.0024239 | 0.0024278 | 4.1471e-05 | 0.91074 | 638 | 0.9078
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 590 | 438 | 0.0022831 | 0.046978 | 417.7 | µS |  |  | 642 | 425.7 | 440 | 371 | 412.2333 | 0.0024258 | 0.0024248 | 4.1421e-05 | 0.91001 | 639 | 0.90854
16:26 | 439 | uS | 220 | ppm | 22.8 | C | 591 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 643 | 424.7 | 441 | 372 | 412.5667 | 0.0024239 | 0.0024248 | 4.1421e-05 | 0.91074 | 640 | 0.9078
16:26 | 439 | uS | 220 | ppm | 22.8 | C | 592 | 439 | 0.0022779 | 0.046983 | 418.7 | µS |  |  | 644 | 425.7 | 441 | 372 | 412.9 | 0.0024219 | 0.0024229 | 4.1387e-05 | 0.91148 | 641 | 0.91074
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 593 | 439 | 0.0022779 | 0.046983 | 418.7 | µS |  |  | 645 | 424.7 | 441 | 373 | 412.9 | 0.0024219 | 0.0024219 | 4.1371e-05 | 0.91148 | 642 | 0.91001
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 594 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 646 | 425.7 | 442 | 373 | 413.5667 | 0.002418 | 0.0024199 | 4.1337e-05 | 0.91295 | 643 | 0.91074
16:26 | 439 | uS | 220 | ppm | 22.8 | C | 595 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 647 | 425.7 | 441 | 371 | 412.5667 | 0.0024239 | 0.0024209 | 4.1354e-05 | 0.91074 | 644 | 0.91148
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 596 | 439 | 0.0022779 | 0.046983 | 418.7 | µS |  |  | 648 | 425.7 | 441 | 371 | 412.5667 | 0.0024239 | 0.0024239 | 4.1404e-05 | 0.91074 | 645 | 0.91148
16:26 | 439 | uS | 220 | ppm | 22.8 | C | 597 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 649 | 426.7 | 442 | 372 | 413.5667 | 0.002418 | 0.0024209 | 4.1354e-05 | 0.91295 | 646 | 0.91295
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 598 | 439 | 0.0022779 | 0.046983 | 418.7 | µS |  |  | 650 | 425.7 | 442 | 372 | 413.2333 | 0.0024199 | 0.002419 | 4.1321e-05 | 0.91221 | 647 | 0.91074
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 599 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 651 | 426.7 | 440 | 372 | 412.9 | 0.0024219 | 0.0024209 | 4.1354e-05 | 0.91148 | 648 | 0.91074
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 600 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 652 | 424.7 | 443 | 372 | 413.2333 | 0.0024199 | 0.0024209 | 4.1354e-05 | 0.91221 | 649 | 0.91295
16:26 | 441 | uS | 221 | ppm | 22.8 | C | 601 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 653 | 426.7 | 441 | 372 | 413.2333 | 0.0024199 | 0.0024199 | 4.1337e-05 | 0.91221 | 650 | 0.91221
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 602 | 441 | 0.0022676 | 0.046994 | 420.7 | µS |  |  | 654 | 424.7 | 442 | 372 | 412.9 | 0.0024219 | 0.0024209 | 4.1354e-05 | 0.91148 | 651 | 0.91148
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 603 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 655 | 426.7 | 442 | 373 | 413.9 | 0.002416 | 0.002419 | 4.1321e-05 | 0.91369 | 652 | 0.91221
16:26 | 439 | uS | 220 | ppm | 22.8 | C | 604 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 656 | 426.7 | 442 | 373 | 413.9 | 0.002416 | 0.002416 | 4.1271e-05 | 0.91369 | 653 | 0.91221
16:26 | 441 | uS | 221 | ppm | 22.8 | C | 605 | 439 | 0.0022779 | 0.046983 | 418.7 | µS |  |  | 657 | 426.7 | 442 | 374 | 414.2333 | 0.0024141 | 0.0024151 | 4.1254e-05 | 0.91442 | 654 | 0.91148
16:26 | 440 | uS | 220 | ppm | 22.8 | C | 606 | 441 | 0.0022676 | 0.046994 | 420.7 | µS |  |  | 658 | 427.7 | 442 | 374 | 414.5667 | 0.0024122 | 0.0024131 | 4.1221e-05 | 0.91516 | 655 | 0.91369
16:26 | 441 | uS | 221 | ppm | 22.8 | C | 607 | 440 | 0.0022727 | 0.046988 | 419.7 | µS |  |  | 659 | 427.7 | 442 | 376 | 415.2333 | 0.0024083 | 0.0024102 | 4.1171e-05 | 0.91663 | 656 | 0.91369
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 608 | 441 | 0.0022676 | 0.046994 | 420.7 | µS |  |  | 660 | 426.7 | 442 | 376 | 414.9 | 0.0024102 | 0.0024093 | 4.1155e-05 | 0.91589 | 657 | 0.91442
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 609 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 661 | 427.7 | 443 | 376 | 415.5667 | 0.0024064 | 0.0024083 | 4.1138e-05 | 0.91737 | 658 | 0.91516
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 610 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 662 | 426.7 | 441 | 376 | 414.5667 | 0.0024122 | 0.0024093 | 4.1155e-05 | 0.91516 | 659 | 0.91663
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 611 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 663 | 427.7 | 443 | 378 | 416.2333 | 0.0024025 | 0.0024073 | 4.1122e-05 | 0.91884 | 660 | 0.91589
16:26 | 441 | uS | 221 | ppm | 22.8 | C | 612 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 664 | 427.7 | 441 | 378 | 415.5667 | 0.0024064 | 0.0024044 | 4.1072e-05 | 0.91737 | 661 | 0.91737
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 613 | 441 | 0.0022676 | 0.046994 | 420.7 | µS |  |  | 665 | 427.7 | 443 | 379 | 416.5667 | 0.0024006 | 0.0024035 | 4.1056e-05 | 0.91957 | 662 | 0.91516
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 614 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 666 | 428.7 | 442 | 379 | 416.5667 | 0.0024006 | 0.0024006 | 4.1006e-05 | 0.91957 | 663 | 0.91884
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 615 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 667 | 427.7 | 443 | 380 | 416.9 | 0.0023987 | 0.0023996 | 4.099e-05 | 0.92031 | 664 | 0.91737
16:26 | 443 | uS | 222 | ppm | 22.8 | C | 616 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 668 | 428.7 | 443 | 380 | 417.2333 | 0.0023967 | 0.0023977 | 4.0957e-05 | 0.92104 | 665 | 0.91957
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 617 | 443 | 0.0022573 | 0.047004 | 422.7 | µS |  |  | 669 | 426.7 | 443 | 382 | 417.2333 | 0.0023967 | 0.0023967 | 4.0941e-05 | 0.92104 | 666 | 0.91957
16:26 | 444 | uS | 222 | ppm | 22.8 | C | 618 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 670 | 428.7 | 443 | 382 | 417.9 | 0.0023929 | 0.0023948 | 4.0908e-05 | 0.92252 | 667 | 0.92031
16:26 | 441 | uS | 221 | ppm | 22.8 | C | 619 | 444 | 0.0022523 | 0.047009 | 423.7 | µS |  |  | 671 | 426.7 | 443 | 381 | 416.9 | 0.0023987 | 0.0023958 | 4.0925e-05 | 0.92031 | 668 | 0.92104
16:26 | 444 | uS | 222 | ppm | 22.8 | C | 620 | 441 | 0.0022676 | 0.046994 | 420.7 | µS |  |  | 672 | 427.7 | 444 | 381 | 417.5667 | 0.0023948 | 0.0023967 | 4.0941e-05 | 0.92178 | 669 | 0.92104
16:26 | 442 | uS | 221 | ppm | 22.8 | C | 621 | 444 | 0.0022523 | 0.047009 | 423.7 | µS |  |  | 673 | 427.7 | 442 | 384 | 417.9 | 0.0023929 | 0.0023939 | 4.0892e-05 | 0.92252 | 670 | 0.92252
16:26 | 443 | uS | 222 | ppm | 22.8 | C | 622 | 442 | 0.0022624 | 0.046999 | 421.7 | µS |  |  | 674 | 428.7 | 445 | 384 | 419.2333 | 0.0023853 | 0.0023891 | 4.0811e-05 | 0.92546 | 671 | 0.92031
16:26 | 443 | uS | 222 | ppm | 22.8 | C | 623 | 443 | 0.0022573 | 0.047004 | 422.7 | µS |  |  | 675 | 428.7 | 442 | 384 | 418.2333 | 0.002391 | 0.0023882 | 4.0794e-05 | 0.92325 | 672 | 0.92178
16:26 | 443 | uS | 222 | ppm | 22.8 | C | 624 | 443 | 0.0022573 | 0.047004 | 422.7 | µS |  |  | 676 | 428.7 | 444 | 384 | 418.9 | 0.0023872 | 0.0023891 | 4.081e-05 | 0.92472 | 673 | 0.92252
16:27 | 444 | uS | 222 | ppm | 22.8 | C | 625 | 443 | 0.0022573 | 0.047004 | 422.7 | µS |  |  | 677 | 428.7 | 443 | 385 | 418.9 | 0.0023872 | 0.0023872 | 4.0778e-05 | 0.92472 | 674 | 0.92546
16:27 | 444 | uS | 222 | ppm | 22.8 | C | 626 | 444 | 0.0022523 | 0.047009 | 423.7 | µS |  |  | 678 | 428.7 | 444 | 385 | 419.2333 | 0.0023853 | 0.0023863 | 4.0762e-05 | 0.92546 | 675 | 0.92325
16:27 | 443 | uS | 222 | ppm | 22.8 | C | 627 | 444 | 0.0022523 | 0.047009 | 423.7 | µS |  |  | 679 | 427.7 | 444 | 386 | 419.2333 | 0.0023853 | 0.0023853 | 4.0746e-05 | 0.92546 | 676 | 0.92472
16:27 | 444 | uS | 222 | ppm | 22.8 | C | 628 | 443 | 0.0022573 | 0.047004 | 422.7 | µS |  |  | 680 | 428.7 | 444 | 386 | 419.5667 | 0.0023834 | 0.0023844 | 4.0729e-05 | 0.9262 | 677 | 0.92472
16:27 | 444 | uS | 222 | ppm | 22.8 | C | 629 | 444 | 0.0022523 | 0.047009 | 423.7 | µS |  |  | 681 | 428.7 | 444 | 386 | 419.5667 | 0.0023834 | 0.0023834 | 4.0713e-05 | 0.9262 | 678 | 0.92546
16:27 | 444 | uS | 222 | ppm | 22.8 | C | 630 | 444 | 0.0022523 | 0.047009 | 423.7 | µS |  |  | 682 | 428.7 | 444 | 386 | 419.5667 | 0.0023834 | 0.0023834 | 4.0713e-05 | 0.9262 | 679 | 0.92546
16:27 | 445 | uS | 223 | ppm | 22.8 | C | 631 | 444 | 0.0022523 | 0.047009 | 423.7 | µS |  |  | 683 | 429.7 | 444 | 387 | 420.2333 | 0.0023796 | 0.0023815 | 4.0681e-05 | 0.92767 | 680 | 0.9262
16:27 | 444 | uS | 222 | ppm | 22.8 | C | 632 | 445 | 0.0022472 | 0.047014 | 424.7 | µS |  |  | 684 | 428.7 | 443 | 387 | 419.5667 | 0.0023834 | 0.0023815 | 4.0681e-05 | 0.9262 | 681 | 0.9262
16:27 | 446 | uS | 223 | ppm | 22.8 | C | 633 | 444 | 0.0022523 | 0.047009 | 423.7 | µS |  |  | 685 | 429.7 | 445 | 386 | 420.2333 | 0.0023796 | 0.0023815 | 4.0681e-05 | 0.92767 | 682 | 0.9262
16:27 | 443 | uS | 222 | ppm | 22.8 | C | 634 | 446 | 0.0022422 | 0.047019 | 425.7 | µS |  |  | 686 | 428.7 | 443 | 386 | 419.2333 | 0.0023853 | 0.0023825 | 4.0697e-05 | 0.92546 | 683 | 0.92767
16:27 | 446 | uS | 223 | ppm | 22.8 | C | 635 | 443 | 0.0022573 | 0.047004 | 422.7 | µS |  |  | 687 | 430.7 | 444 | 386 | 420.2333 | 0.0023796 | 0.0023825 | 4.0697e-05 | 0.92767 | 684 | 0.9262
16:27 | 444 | uS | 222 | ppm | 22.8 | C | 636 | 446 | 0.0022422 | 0.047019 | 425.7 | µS |  |  | 688 | 427.7 | 444 | 386 | 419.2333 | 0.0023853 | 0.0023825 | 4.0697e-05 | 0.92546 | 685 | 0.92767
16:27 | 445 | uS | 223 | ppm | 22.8 | C | 637 | 444 | 0.0022523 | 0.047009 | 423.7 | µS |  |  | 689 | 429.7 | 445 | 387 | 420.5667 | 0.0023777 | 0.0023815 | 4.0681e-05 | 0.9284 | 686 | 0.92546
16:27 | 445 | uS | 223 | ppm | 22.8 | C | 638 | 445 | 0.0022472 | 0.047014 | 424.7 | µS |  |  | 690 | 428.7 | 445 | 387 | 420.2333 | 0.0023796 | 0.0023787 | 4.0633e-05 | 0.92767 | 687 | 0.92767
16:27 | 445 | uS | 223 | ppm | 22.8 | C | 639 | 445 | 0.0022472 | 0.047014 | 424.7 | µS |  |  | 691 | 429.7 | 444 | 387 | 420.2333 | 0.0023796 | 0.0023796 | 4.0649e-05 | 0.92767 | 688 | 0.92546
16:27 | 446 | uS | 223 | ppm | 22.8 | C | 640 | 445 | 0.0022472 | 0.047014 | 424.7 | µS |  |  | 692 | 429.7 | 445 | 387 | 420.5667 | 0.0023777 | 0.0023787 | 4.0633e-05 | 0.9284 | 689 | 0.9284
16:27 | 446 | uS | 223 | ppm | 22.8 | C | 641 | 446 | 0.0022422 | 0.047019 | 425.7 | µS |  |  | 693 | 429.7 | 445 | 388 | 420.9 | 0.0023759 | 0.0023768 | 4.06e-05 | 0.92914 | 690 | 0.92767
16:27 | 445 | uS | 223 | ppm | 22.8 | C | 642 | 446 | 0.0022422 | 0.047019 | 425.7 | µS |  |  | 694 | 430.7 | 445 | 388 | 421.2333 | 0.002374 | 0.0023749 | 4.0568e-05 | 0.92987 | 691 | 0.92767
16:27 | 446 | uS | 223 | ppm | 22.8 | C | 643 | 445 | 0.0022472 | 0.047014 | 424.7 | µS |  |  | 695 | 430.7 | 444 | 388 | 420.9 | 0.0023759 | 0.0023749 | 4.0568e-05 | 0.92914 | 692 | 0.9284
16:27 | 445 | uS | 223 | ppm | 22.8 | C | 644 | 446 | 0.0022422 | 0.047019 | 425.7 | µS |  |  | 696 | 429.7 | 446 | 388 | 421.2333 | 0.002374 | 0.0023749 | 4.0568e-05 | 0.92987 | 693 | 0.92914
16:27 | 446 | uS | 223 | ppm | 22.8 | C | 645 | 445 | 0.0022472 | 0.047014 | 424.7 | µS |  |  | 697 | 430.7 | 444 | 390 | 421.5667 | 0.0023721 | 0.002373 | 4.0536e-05 | 0.93061 | 694 | 0.92987
16:27 | 446 | uS | 223 | ppm | 22.8 | C | 646 | 446 | 0.0022422 | 0.047019 | 425.7 | µS |  |  | 698 | 429.7 | 445 | 390 | 421.5667 | 0.0023721 | 0.0023721 | 4.052e-05 | 0.93061 | 695 | 0.92914
16:27 | 446 | uS | 223 | ppm | 22.8 | C | 647 | 446 | 0.0022422 | 0.047019 | 425.7 | µS |  |  | 699 | 430.7 | 445 | 390 | 421.9 | 0.0023702 | 0.0023712 | 4.0504e-05 | 0.93135 | 696 | 0.92987
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 648 | 446 | 0.0022422 | 0.047019 | 425.7 | µS |  |  | 700 | 430.7 | 446 | 390 | 422.2333 | 0.0023684 | 0.0023693 | 4.0472e-05 | 0.93208 | 697 | 0.93061
16:27 | 446 | uS | 223 | ppm | 22.8 | C | 649 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 701 | 430.7 | 446 | 391 | 422.5667 | 0.0023665 | 0.0023674 | 4.044e-05 | 0.93282 | 698 | 0.93061
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 650 | 446 | 0.0022422 | 0.047019 | 425.7 | µS |  |  | 702 | 430.7 | 445 | 391 | 422.2333 | 0.0023684 | 0.0023674 | 4.044e-05 | 0.93208 | 699 | 0.93135
16:27 | 445 | uS | 223 | ppm | 22.8 | C | 651 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 703 | 430.7 | 446 | 392 | 422.9 | 0.0023646 | 0.0023665 | 4.0424e-05 | 0.93355 | 700 | 0.93208
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 652 | 445 | 0.0022472 | 0.047014 | 424.7 | µS |  |  | 704 | 431.7 | 446 | 392 | 423.2333 | 0.0023628 | 0.0023637 | 4.0376e-05 | 0.93429 | 701 | 0.93282
16:27 | 445 | uS | 223 | ppm | 22.8 | C | 653 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 705 | 429.7 | 446 | 392 | 422.5667 | 0.0023665 | 0.0023646 | 4.0392e-05 | 0.93282 | 702 | 0.93208
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 654 | 445 | 0.0022472 | 0.047014 | 424.7 | µS |  |  | 706 | 431.7 | 445 | 392 | 422.9 | 0.0023646 | 0.0023656 | 4.0408e-05 | 0.93355 | 703 | 0.93355
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 655 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 707 | 429.7 | 447 | 393 | 423.2333 | 0.0023628 | 0.0023637 | 4.0376e-05 | 0.93429 | 704 | 0.93429
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 656 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 708 | 431.7 | 445 | 393 | 423.2333 | 0.0023628 | 0.0023628 | 4.036e-05 | 0.93429 | 705 | 0.93282
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 657 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 709 | 429.7 | 446 | 390 | 421.9 | 0.0023702 | 0.0023665 | 4.0424e-05 | 0.93135 | 706 | 0.93355
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 658 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 710 | 430.7 | 445 | 390 | 421.9 | 0.0023702 | 0.0023702 | 4.0488e-05 | 0.93135 | 707 | 0.93429
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 659 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 711 | 430.7 | 446 | 392 | 422.9 | 0.0023646 | 0.0023674 | 4.044e-05 | 0.93355 | 708 | 0.93429
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 660 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 712 | 430.7 | 446 | 392 | 422.9 | 0.0023646 | 0.0023646 | 4.0392e-05 | 0.93355 | 709 | 0.93135
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 661 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 713 | 431.7 | 445 | 393 | 423.2333 | 0.0023628 | 0.0023637 | 4.0376e-05 | 0.93429 | 710 | 0.93135
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 662 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 714 | 431.7 | 446 | 393 | 423.5667 | 0.0023609 | 0.0023618 | 4.0345e-05 | 0.93503 | 711 | 0.93355
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 663 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 715 | 431.7 | 446 | 393 | 423.5667 | 0.0023609 | 0.0023609 | 4.0329e-05 | 0.93503 | 712 | 0.93355
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 664 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 716 | 431.7 | 447 | 393 | 423.9 | 0.002359 | 0.00236 | 4.0313e-05 | 0.93576 | 713 | 0.93429
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 665 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 717 | 430.7 | 445 | 395 | 423.5667 | 0.0023609 | 0.00236 | 4.0313e-05 | 0.93503 | 714 | 0.93503
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 666 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 718 | 431.7 | 447 | 395 | 424.5667 | 0.0023553 | 0.0023581 | 4.0281e-05 | 0.93723 | 715 | 0.93503
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 667 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 719 | 430.7 | 445 | 395 | 423.5667 | 0.0023609 | 0.0023581 | 4.0281e-05 | 0.93503 | 716 | 0.93576
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 668 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 720 | 430.7 | 447 | 395 | 424.2333 | 0.0023572 | 0.002359 | 4.0297e-05 | 0.9365 | 717 | 0.93503
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 669 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 721 | 431.7 | 446 | 398 | 425.2333 | 0.0023517 | 0.0023544 | 4.0218e-05 | 0.9387 | 718 | 0.93723
16:27 | 447 | uS | 224 | ppm | 22.8 | C | 670 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 722 | 431.7 | 447 | 398 | 425.5667 | 0.0023498 | 0.0023507 | 4.0155e-05 | 0.93944 | 719 | 0.93503
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 671 | 447 | 0.0022371 | 0.047024 | 426.7 | µS |  |  | 723 | 431.7 | 447 | 398 | 425.5667 | 0.0023498 | 0.0023498 | 4.0139e-05 | 0.93944 | 720 | 0.9365
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 672 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 724 | 431.7 | 446 | 398 | 425.2333 | 0.0023517 | 0.0023507 | 4.0155e-05 | 0.9387 | 721 | 0.9387
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 673 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 725 | 432.7 | 447 | 399 | 426.2333 | 0.0023461 | 0.0023489 | 4.0124e-05 | 0.94091 | 722 | 0.93944
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 674 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 726 | 430.7 | 447 | 399 | 425.5667 | 0.0023498 | 0.002348 | 4.0108e-05 | 0.93944 | 723 | 0.93944
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 675 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 727 | 432.7 | 448 | 400 | 426.9 | 0.0023425 | 0.0023461 | 4.0077e-05 | 0.94238 | 724 | 0.9387
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 676 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 728 | 430.7 | 445 | 400 | 425.2333 | 0.0023517 | 0.0023471 | 4.0092e-05 | 0.9387 | 725 | 0.94091
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 677 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 729 | 431.7 | 447 | 400 | 426.2333 | 0.0023461 | 0.0023489 | 4.0124e-05 | 0.94091 | 726 | 0.93944
16:27 | 448 | uS | 224 | ppm | 22.8 | C | 678 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 730 | 430.7 | 446 | 400 | 425.5667 | 0.0023498 | 0.002348 | 4.0108e-05 | 0.93944 | 727 | 0.94238
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 679 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 731 | 431.7 | 447 | 401 | 426.5667 | 0.0023443 | 0.0023471 | 4.0092e-05 | 0.94165 | 728 | 0.9387
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 680 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 732 | 431.7 | 447 | 401 | 426.5667 | 0.0023443 | 0.0023443 | 4.0045e-05 | 0.94165 | 729 | 0.94091
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 681 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 733 | 431.7 | 447 | 398 | 425.5667 | 0.0023498 | 0.0023471 | 4.0092e-05 | 0.93944 | 730 | 0.93944
16:27 | 450 | uS | 225 | ppm | 22.8 | C | 682 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 734 | 432.7 | 448 | 398 | 426.2333 | 0.0023461 | 0.002348 | 4.0108e-05 | 0.94091 | 731 | 0.94165
16:27 | 449 | uS | 225 | ppm | 22.8 | C | 683 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 735 | 432.7 | 447 | 400 | 426.5667 | 0.0023443 | 0.0023452 | 4.0061e-05 | 0.94165 | 732 | 0.94165
16:27 | 450 | uS | 225 | ppm | 22.8 | C | 684 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 736 | 432.7 | 447 | 400 | 426.5667 | 0.0023443 | 0.0023443 | 4.0045e-05 | 0.94165 | 733 | 0.93944
16:28 | 449 | uS | 225 | ppm | 22.8 | C | 685 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 737 | 432.7 | 447 | 401 | 426.9 | 0.0023425 | 0.0023434 | 4.0029e-05 | 0.94238 | 734 | 0.94091
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 686 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 738 | 431.7 | 448 | 401 | 426.9 | 0.0023425 | 0.0023425 | 4.0014e-05 | 0.94238 | 735 | 0.94165
16:28 | 448 | uS | 224 | ppm | 22.8 | C | 687 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 739 | 432.7 | 446 | 402 | 426.9 | 0.0023425 | 0.0023425 | 4.0014e-05 | 0.94238 | 736 | 0.94165
16:28 | 450 | uS | 225 | ppm | 22.8 | C | 688 | 448 | 0.0022321 | 0.047029 | 427.7 | µS |  |  | 740 | 431.7 | 447 | 402 | 426.9 | 0.0023425 | 0.0023425 | 4.0014e-05 | 0.94238 | 737 | 0.94238
16:28 | 449 | uS | 225 | ppm | 22.8 | C | 689 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 741 | 432.7 | 447 | 403 | 427.5667 | 0.0023388 | 0.0023406 | 3.9983e-05 | 0.94386 | 738 | 0.94238
16:28 | 450 | uS | 225 | ppm | 22.8 | C | 690 | 449 | 0.0022272 | 0.047034 | 428.7 | µS |  |  | 742 | 432.7 | 448 | 403 | 427.9 | 0.002337 | 0.0023379 | 3.9936e-05 | 0.94459 | 739 | 0.94238
16:28 | 450 | uS | 225 | ppm | 22.8 | C | 691 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 743 | 432.7 | 448 | 404 | 428.2333 | 0.0023352 | 0.0023361 | 3.9905e-05 | 0.94533 | 740 | 0.94238
16:28 | 450 | uS | 225 | ppm | 22.8 | C | 692 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 744 | 432.7 | 447 | 404 | 427.9 | 0.002337 | 0.0023361 | 3.9905e-05 | 0.94459 | 741 | 0.94386
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 693 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 745 | 432.7 | 448 | 407 | 429.2333 | 0.0023297 | 0.0023334 | 3.9858e-05 | 0.94753 | 742 | 0.94459
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 694 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 746 | 433.7 | 448 | 407 | 429.5667 | 0.0023279 | 0.0023288 | 3.9781e-05 | 0.94827 | 743 | 0.94533
16:28 | 450 | uS | 225 | ppm | 22.8 | C | 695 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 747 | 431.7 | 448 | 406 | 428.5667 | 0.0023334 | 0.0023306 | 3.9812e-05 | 0.94606 | 744 | 0.94459
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 696 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 748 | 433.7 | 447 | 406 | 428.9 | 0.0023315 | 0.0023325 | 3.9843e-05 | 0.9468 | 745 | 0.94753
16:28 | 450 | uS | 225 | ppm | 22.8 | C | 697 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 749 | 430.7 | 449 | 407 | 428.9 | 0.0023315 | 0.0023315 | 3.9827e-05 | 0.9468 | 746 | 0.94827
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 698 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 750 | 432.7 | 447 | 407 | 428.9 | 0.0023315 | 0.0023315 | 3.9827e-05 | 0.9468 | 747 | 0.94606
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 699 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 751 | 431.7 | 448 | 407 | 428.9 | 0.0023315 | 0.0023315 | 3.9827e-05 | 0.9468 | 748 | 0.9468
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 700 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 752 | 432.7 | 448 | 407 | 429.2333 | 0.0023297 | 0.0023306 | 3.9812e-05 | 0.94753 | 749 | 0.9468
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 701 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 753 | 432.7 | 448 | 407 | 429.2333 | 0.0023297 | 0.0023297 | 3.9796e-05 | 0.94753 | 750 | 0.9468
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 702 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 754 | 432.7 | 449 | 407 | 429.5667 | 0.0023279 | 0.0023288 | 3.9781e-05 | 0.94827 | 751 | 0.9468
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 703 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 755 | 433.7 | 448 | 408 | 429.9 | 0.0023261 | 0.002327 | 3.975e-05 | 0.94901 | 752 | 0.94753
16:28 | 450 | uS | 225 | ppm | 22.8 | C | 704 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 756 | 433.7 | 448 | 408 | 429.9 | 0.0023261 | 0.0023261 | 3.9735e-05 | 0.94901 | 753 | 0.94753
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 705 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 757 | 433.7 | 448 | 405 | 428.9 | 0.0023315 | 0.0023288 | 3.9781e-05 | 0.9468 | 754 | 0.94827
16:28 | 450 | uS | 225 | ppm | 22.8 | C | 706 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 758 | 433.7 | 450 | 405 | 429.5667 | 0.0023279 | 0.0023297 | 3.9796e-05 | 0.94827 | 755 | 0.94901
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 707 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 759 | 432.7 | 447 | 407 | 428.9 | 0.0023315 | 0.0023297 | 3.9796e-05 | 0.9468 | 756 | 0.94901
16:28 | 450 | uS | 225 | ppm | 22.8 | C | 708 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 760 | 433.7 | 449 | 407 | 429.9 | 0.0023261 | 0.0023288 | 3.9781e-05 | 0.94901 | 757 | 0.9468
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 709 | 450 | 0.0022222 | 0.047039 | 429.7 | µS |  |  | 761 | 432.7 | 448 | 407 | 429.2333 | 0.0023297 | 0.0023279 | 3.9765e-05 | 0.94753 | 758 | 0.94827
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 710 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 762 | 433.7 | 449 | 407 | 429.9 | 0.0023261 | 0.0023279 | 3.9765e-05 | 0.94901 | 759 | 0.9468
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 711 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 763 | 433.7 | 449 | 408 | 430.2333 | 0.0023243 | 0.0023252 | 3.9719e-05 | 0.94974 | 760 | 0.94901
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 712 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 764 | 433.7 | 448 | 410 | 430.5667 | 0.0023225 | 0.0023234 | 3.9688e-05 | 0.95048 | 761 | 0.94753
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 713 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 765 | 433.7 | 449 | 410 | 430.9 | 0.0023207 | 0.0023216 | 3.9658e-05 | 0.95121 | 762 | 0.94901
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 714 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 766 | 433.7 | 449 | 410 | 430.9 | 0.0023207 | 0.0023207 | 3.9642e-05 | 0.95121 | 763 | 0.94974
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 715 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 767 | 434.7 | 449 | 412 | 431.9 | 0.0023154 | 0.002318 | 3.9596e-05 | 0.95342 | 764 | 0.95048
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 716 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 768 | 433.7 | 448 | 412 | 431.2333 | 0.0023189 | 0.0023171 | 3.9581e-05 | 0.95195 | 765 | 0.95121
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 717 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 769 | 434.7 | 450 | 410 | 431.5667 | 0.0023171 | 0.002318 | 3.9596e-05 | 0.95269 | 766 | 0.95121
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 718 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 770 | 432.7 | 448 | 410 | 430.2333 | 0.0023243 | 0.0023207 | 3.9642e-05 | 0.94974 | 767 | 0.95342
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 719 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 771 | 434.7 | 449 | 412 | 431.9 | 0.0023154 | 0.0023198 | 3.9627e-05 | 0.95342 | 768 | 0.95195
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 720 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 772 | 432.7 | 449 | 412 | 431.2333 | 0.0023189 | 0.0023171 | 3.9581e-05 | 0.95195 | 769 | 0.95269
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 721 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 773 | 434.7 | 450 | 412 | 432.2333 | 0.0023136 | 0.0023162 | 3.9566e-05 | 0.95416 | 770 | 0.94974
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 722 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 774 | 432.7 | 450 | 412 | 431.5667 | 0.0023171 | 0.0023154 | 3.9551e-05 | 0.95269 | 771 | 0.95342
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 723 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 775 | 433.7 | 449 | 411 | 431.2333 | 0.0023189 | 0.002318 | 3.9596e-05 | 0.95195 | 772 | 0.95195
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 724 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 776 | 433.7 | 449 | 411 | 431.2333 | 0.0023189 | 0.0023189 | 3.9612e-05 | 0.95195 | 773 | 0.95416
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 725 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 777 | 433.7 | 449 | 413 | 431.9 | 0.0023154 | 0.0023171 | 3.9581e-05 | 0.95342 | 774 | 0.95269
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 726 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 778 | 434.7 | 451 | 413 | 432.9 | 0.00231 | 0.0023127 | 3.9505e-05 | 0.95563 | 775 | 0.95195
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 727 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 779 | 434.7 | 448 | 411 | 431.2333 | 0.0023189 | 0.0023145 | 3.9535e-05 | 0.95195 | 776 | 0.95195
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 728 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 780 | 434.7 | 449 | 411 | 431.5667 | 0.0023171 | 0.002318 | 3.9596e-05 | 0.95269 | 777 | 0.95342
16:28 | 451 | uS | 226 | ppm | 22.8 | C | 729 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 781 | 434.7 | 449 | 413 | 432.2333 | 0.0023136 | 0.0023154 | 3.9551e-05 | 0.95416 | 778 | 0.95563
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 730 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 782 | 433.7 | 450 | 413 | 432.2333 | 0.0023136 | 0.0023136 | 3.952e-05 | 0.95416 | 779 | 0.95195
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 731 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 783 | 434.7 | 449 | 414 | 432.5667 | 0.0023118 | 0.0023127 | 3.9505e-05 | 0.95489 | 780 | 0.95269
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 732 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 784 | 433.7 | 449 | 414 | 432.2333 | 0.0023136 | 0.0023127 | 3.9505e-05 | 0.95416 | 781 | 0.95416
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 733 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 785 | 434.7 | 450 | 415 | 433.2333 | 0.0023082 | 0.0023109 | 3.9474e-05 | 0.95636 | 782 | 0.95416
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 734 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 786 | 433.7 | 450 | 415 | 432.9 | 0.00231 | 0.0023091 | 3.9444e-05 | 0.95563 | 783 | 0.95489
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 735 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 787 | 434.7 | 450 | 415 | 433.2333 | 0.0023082 | 0.0023091 | 3.9444e-05 | 0.95636 | 784 | 0.95416
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 736 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 788 | 434.7 | 448 | 415 | 432.5667 | 0.0023118 | 0.00231 | 3.9459e-05 | 0.95489 | 785 | 0.95636
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 737 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 789 | 434.7 | 451 | 416 | 433.9 | 0.0023047 | 0.0023082 | 3.9429e-05 | 0.95784 | 786 | 0.95563
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 738 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 790 | 434.7 | 449 | 416 | 433.2333 | 0.0023082 | 0.0023065 | 3.9399e-05 | 0.95636 | 787 | 0.95636
16:28 | 452 | uS | 226 | ppm | 22.8 | C | 739 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 791 | 434.7 | 450 | 419 | 434.5667 | 0.0023011 | 0.0023047 | 3.9368e-05 | 0.95931 | 788 | 0.95489
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 740 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 792 | 435.7 | 449 | 419 | 434.5667 | 0.0023011 | 0.0023011 | 3.9308e-05 | 0.95931 | 789 | 0.95784
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 741 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 793 | 433.7 | 450 | 418 | 433.9 | 0.0023047 | 0.0023029 | 3.9338e-05 | 0.95784 | 790 | 0.95636
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 742 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 794 | 435.7 | 450 | 418 | 434.5667 | 0.0023011 | 0.0023029 | 3.9338e-05 | 0.95931 | 791 | 0.95931
16:28 | 453 | uS | 227 | ppm | 22.8 | C | 743 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 795 | 433.7 | 450 | 420 | 434.5667 | 0.0023011 | 0.0023011 | 3.9308e-05 | 0.95931 | 792 | 0.95931
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 744 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 796 | 435.7 | 450 | 420 | 435.2333 | 0.0022976 | 0.0022994 | 3.9278e-05 | 0.96078 | 793 | 0.95784
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 745 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 797 | 432.7 | 450 | 419 | 433.9 | 0.0023047 | 0.0023011 | 3.9308e-05 | 0.95784 | 794 | 0.95931
16:29 | 452 | uS | 226 | ppm | 22.8 | C | 746 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 798 | 435.7 | 451 | 419 | 435.2333 | 0.0022976 | 0.0023011 | 3.9308e-05 | 0.96078 | 795 | 0.95931
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 747 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 799 | 433.7 | 448 | 419 | 433.5667 | 0.0023065 | 0.002302 | 3.9323e-05 | 0.9571 | 796 | 0.96078
16:29 | 451 | uS | 226 | ppm | 22.8 | C | 748 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 800 | 434.7 | 450 | 419 | 434.5667 | 0.0023011 | 0.0023038 | 3.9353e-05 | 0.95931 | 797 | 0.95784
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 749 | 451 | 0.0022173 | 0.047044 | 430.7 | µS |  |  | 801 | 434.7 | 450 | 420 | 434.9 | 0.0022994 | 0.0023003 | 3.9293e-05 | 0.96004 | 798 | 0.96078
16:29 | 452 | uS | 226 | ppm | 22.8 | C | 750 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 802 | 435.7 | 450 | 420 | 435.2333 | 0.0022976 | 0.0022985 | 3.9263e-05 | 0.96078 | 799 | 0.9571
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 751 | 452 | 0.0022124 | 0.047049 | 431.7 | µS |  |  | 803 | 435.7 | 450 | 418 | 434.5667 | 0.0023011 | 0.0022994 | 3.9278e-05 | 0.95931 | 800 | 0.95931
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 752 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 804 | 435.7 | 449 | 418 | 434.2333 | 0.0023029 | 0.002302 | 3.9323e-05 | 0.95857 | 801 | 0.96004
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 753 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 805 | 435.7 | 450 | 420 | 435.2333 | 0.0022976 | 0.0023003 | 3.9293e-05 | 0.96078 | 802 | 0.96078
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 754 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 806 | 435.7 | 450 | 420 | 435.2333 | 0.0022976 | 0.0022976 | 3.9248e-05 | 0.96078 | 803 | 0.95931
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 755 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 807 | 435.7 | 451 | 420 | 435.5667 | 0.0022959 | 0.0022967 | 3.9233e-05 | 0.96152 | 804 | 0.95857
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 756 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 808 | 435.7 | 448 | 420 | 434.5667 | 0.0023011 | 0.0022985 | 3.9263e-05 | 0.95931 | 805 | 0.96078
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 757 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 809 | 435.7 | 451 | 422 | 436.2333 | 0.0022924 | 0.0022967 | 3.9233e-05 | 0.96299 | 806 | 0.96078
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 758 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 810 | 435.7 | 450 | 422 | 435.9 | 0.0022941 | 0.0022932 | 3.9173e-05 | 0.96225 | 807 | 0.96152
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 759 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 811 | 435.7 | 450 | 423 | 436.2333 | 0.0022924 | 0.0022932 | 3.9173e-05 | 0.96299 | 808 | 0.95931
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 760 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 812 | 435.7 | 450 | 423 | 436.2333 | 0.0022924 | 0.0022924 | 3.9158e-05 | 0.96299 | 809 | 0.96299
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 761 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 813 | 435.7 | 450 | 423 | 436.2333 | 0.0022924 | 0.0022924 | 3.9158e-05 | 0.96299 | 810 | 0.96225
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 762 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 814 | 435.7 | 451 | 423 | 436.5667 | 0.0022906 | 0.0022915 | 3.9143e-05 | 0.96372 | 811 | 0.96299
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 763 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 815 | 436.7 | 451 | 426 | 437.9 | 0.0022836 | 0.0022871 | 3.9068e-05 | 0.96667 | 812 | 0.96299
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 764 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 816 | 435.7 | 451 | 426 | 437.5667 | 0.0022854 | 0.0022845 | 3.9024e-05 | 0.96593 | 813 | 0.96299
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 765 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 817 | 436.7 | 450 | 425 | 437.2333 | 0.0022871 | 0.0022862 | 3.9053e-05 | 0.96519 | 814 | 0.96372
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 766 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 818 | 435.7 | 452 | 425 | 437.5667 | 0.0022854 | 0.0022862 | 3.9053e-05 | 0.96593 | 815 | 0.96667
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 767 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 819 | 436.7 | 450 | 425 | 437.2333 | 0.0022871 | 0.0022862 | 3.9053e-05 | 0.96519 | 816 | 0.96593
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 768 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 820 | 435.7 | 451 | 425 | 437.2333 | 0.0022871 | 0.0022871 | 3.9068e-05 | 0.96519 | 817 | 0.96519
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 769 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 821 | 436.7 | 450 | 426 | 437.5667 | 0.0022854 | 0.0022862 | 3.9053e-05 | 0.96593 | 818 | 0.96593
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 770 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 822 | 434.7 | 451 | 426 | 437.2333 | 0.0022871 | 0.0022862 | 3.9053e-05 | 0.96519 | 819 | 0.96519
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 771 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 823 | 436.7 | 451 | 425 | 437.5667 | 0.0022854 | 0.0022862 | 3.9053e-05 | 0.96593 | 820 | 0.96519
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 772 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 824 | 434.7 | 450 | 425 | 436.5667 | 0.0022906 | 0.002288 | 3.9083e-05 | 0.96372 | 821 | 0.96593
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 773 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 825 | 436.7 | 451 | 426 | 437.9 | 0.0022836 | 0.0022871 | 3.9068e-05 | 0.96667 | 822 | 0.96519
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 774 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 826 | 434.7 | 450 | 426 | 436.9 | 0.0022889 | 0.0022862 | 3.9053e-05 | 0.96446 | 823 | 0.96593
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 775 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 827 | 435.7 | 452 | 424 | 437.2333 | 0.0022871 | 0.002288 | 3.9083e-05 | 0.96519 | 824 | 0.96372
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 776 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 828 | 435.7 | 449 | 424 | 436.2333 | 0.0022924 | 0.0022897 | 3.9113e-05 | 0.96299 | 825 | 0.96667
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 777 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 829 | 436.7 | 451 | 426 | 437.9 | 0.0022836 | 0.002288 | 3.9083e-05 | 0.96667 | 826 | 0.96446
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 778 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 830 | 436.7 | 450 | 426 | 437.5667 | 0.0022854 | 0.0022845 | 3.9024e-05 | 0.96593 | 827 | 0.96519
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 779 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 831 | 436.7 | 451 | 427 | 438.2333 | 0.0022819 | 0.0022836 | 3.9009e-05 | 0.9674 | 828 | 0.96299
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 780 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 832 | 436.7 | 451 | 427 | 438.2333 | 0.0022819 | 0.0022819 | 3.8979e-05 | 0.9674 | 829 | 0.96667
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 781 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 833 | 436.7 | 450 | 427 | 437.9 | 0.0022836 | 0.0022828 | 3.8994e-05 | 0.96667 | 830 | 0.96593
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 782 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 834 | 436.7 | 451 | 427 | 438.2333 | 0.0022819 | 0.0022828 | 3.8994e-05 | 0.9674 | 831 | 0.9674
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 783 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 835 | 436.7 | 451 | 429 | 438.9 | 0.0022784 | 0.0022802 | 3.8949e-05 | 0.96887 | 832 | 0.9674
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 784 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 836 | 436.7 | 452 | 429 | 439.2333 | 0.0022767 | 0.0022776 | 3.8905e-05 | 0.96961 | 833 | 0.96667
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 785 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 837 | 436.7 | 449 | 429 | 438.2333 | 0.0022819 | 0.0022793 | 3.8935e-05 | 0.9674 | 834 | 0.9674
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 786 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 838 | 435.7 | 451 | 429 | 438.5667 | 0.0022802 | 0.002281 | 3.8964e-05 | 0.96814 | 835 | 0.96887
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 787 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 839 | 436.7 | 450 | 431 | 439.2333 | 0.0022767 | 0.0022784 | 3.892e-05 | 0.96961 | 836 | 0.96961
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 788 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 840 | 436.7 | 451 | 431 | 439.5667 | 0.002275 | 0.0022758 | 3.8876e-05 | 0.97035 | 837 | 0.9674
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 789 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 841 | 436.7 | 451 | 430 | 439.2333 | 0.0022767 | 0.0022758 | 3.8876e-05 | 0.96961 | 838 | 0.96814
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 790 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 842 | 436.7 | 451 | 430 | 439.2333 | 0.0022767 | 0.0022767 | 3.889e-05 | 0.96961 | 839 | 0.96961
16:29 | 456 | uS | 228 | ppm | 22.8 | C | 791 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 843 | 436.7 | 451 | 431 | 439.5667 | 0.002275 | 0.0022758 | 3.8876e-05 | 0.97035 | 840 | 0.97035
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 792 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 844 | 437.7 | 451 | 431 | 439.9 | 0.0022732 | 0.0022741 | 3.8846e-05 | 0.97108 | 841 | 0.96961
16:29 | 456 | uS | 228 | ppm | 22.8 | C | 793 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 845 | 436.7 | 452 | 431 | 439.9 | 0.0022732 | 0.0022732 | 3.8831e-05 | 0.97108 | 842 | 0.96961
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 794 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 846 | 437.7 | 449 | 431 | 439.2333 | 0.0022767 | 0.002275 | 3.8861e-05 | 0.96961 | 843 | 0.97035
16:29 | 456 | uS | 228 | ppm | 22.8 | C | 795 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 847 | 436.7 | 451 | 431 | 439.5667 | 0.002275 | 0.0022758 | 3.8876e-05 | 0.97035 | 844 | 0.97108
16:29 | 453 | uS | 227 | ppm | 22.8 | C | 796 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 848 | 437.7 | 451 | 431 | 439.9 | 0.0022732 | 0.0022741 | 3.8846e-05 | 0.97108 | 845 | 0.97108
16:29 | 456 | uS | 228 | ppm | 22.8 | C | 797 | 453 | 0.0022075 | 0.047054 | 432.7 | µS |  |  | 849 | 436.7 | 452 | 432 | 440.2333 | 0.0022715 | 0.0022724 | 3.8817e-05 | 0.97182 | 846 | 0.96961
16:29 | 454 | uS | 227 | ppm | 22.8 | C | 798 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 850 | 437.7 | 451 | 432 | 440.2333 | 0.0022715 | 0.0022715 | 3.8802e-05 | 0.97182 | 847 | 0.97035
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 799 | 454 | 0.0022026 | 0.047058 | 433.7 | µS |  |  | 851 | 435.7 | 451 | 430 | 438.9 | 0.0022784 | 0.002275 | 3.8861e-05 | 0.96887 | 848 | 0.97108
16:29 | 455 | uS | 228 | ppm | 22.8 | C | 800 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 852 | 437.7 | 452 | 430 | 439.9 | 0.0022732 | 0.0022758 | 3.8876e-05 | 0.97108 | 849 | 0.97182
16:29 | 456 | uS | 228 | ppm | 22.8 | C | 801 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 853 | 435.7 | 452 | 431 | 439.5667 | 0.002275 | 0.0022741 | 3.8846e-05 | 0.97035 | 850 | 0.97182
16:29 | 456 | uS | 228 | ppm | 22.8 | C | 802 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 854 | 437.7 | 452 | 431 | 440.2333 | 0.0022715 | 0.0022732 | 3.8831e-05 | 0.97182 | 851 | 0.96887
16:29 | 456 | uS | 228 | ppm | 22.8 | C | 803 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 855 | 435.7 | 450 | 432 | 439.2333 | 0.0022767 | 0.0022741 | 3.8846e-05 | 0.96961 | 852 | 0.97108
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 804 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 856 | 437.7 | 452 | 432 | 440.5667 | 0.0022698 | 0.0022732 | 3.8831e-05 | 0.97255 | 853 | 0.97035
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 805 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 857 |  | 451 | 433 | 442 | 0.0022624 | 0.0022661 | 3.871e-05 | 0.97572 | 854 | 0.97182
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 806 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 858 |  | 452 | 433 | 442.5 | 0.0022599 | 0.0022612 | 3.8625e-05 | 0.97682 | 855 | 0.96961
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 807 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 859 |  | 451 | 434 | 442.5 | 0.0022599 | 0.0022599 | 3.8603e-05 | 0.97682 | 856 | 0.97255
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 808 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 860 |  | 451 | 434 | 442.5 | 0.0022599 | 0.0022599 | 3.8603e-05 | 0.97682 | 857 | 0.97572
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 809 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 861 |  | 452 | 434 | 443 | 0.0022573 | 0.0022586 | 3.8581e-05 | 0.97792 | 858 | 0.97682
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 810 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 862 |  | 452 | 434 | 443 | 0.0022573 | 0.0022573 | 3.856e-05 | 0.97792 | 859 | 0.97682
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 811 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 863 |  | 452 | 437 | 444.5 | 0.0022497 | 0.0022535 | 3.8495e-05 | 0.98124 | 860 | 0.97682
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 812 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 864 |  | 450 | 437 | 443.5 | 0.0022548 | 0.0022523 | 3.8473e-05 | 0.97903 | 861 | 0.97792
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 813 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 865 |  | 453 | 436 | 444.5 | 0.0022497 | 0.0022523 | 3.8473e-05 | 0.98124 | 862 | 0.97792
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 814 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 866 |  | 451 | 436 | 443.5 | 0.0022548 | 0.0022523 | 3.8473e-05 | 0.97903 | 863 | 0.98124
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 815 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 867 |  | 452 | 437 | 444.5 | 0.0022497 | 0.0022523 | 3.8473e-05 | 0.98124 | 864 | 0.97903
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 816 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 868 |  | 452 | 437 | 444.5 | 0.0022497 | 0.0022497 | 3.8429e-05 | 0.98124 | 865 | 0.98124
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 817 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 869 |  | 452 | 438 | 445 | 0.0022472 | 0.0022485 | 3.8408e-05 | 0.98234 | 866 | 0.97903
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 818 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 870 |  | 452 | 438 | 445 | 0.0022472 | 0.0022472 | 3.8386e-05 | 0.98234 | 867 | 0.98124
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 819 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 871 |  | 452 | 438 | 445 | 0.0022472 | 0.0022472 | 3.8386e-05 | 0.98234 | 868 | 0.98124
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 820 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 872 |  | 452 | 438 | 445 | 0.0022472 | 0.0022472 | 3.8386e-05 | 0.98234 | 869 | 0.98234
16:30 | 455 | uS | 228 | ppm | 22.8 | C | 821 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 873 |  | 451 | 439 | 445 | 0.0022472 | 0.0022472 | 3.8386e-05 | 0.98234 | 870 | 0.98234
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 822 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 874 |  | 453 | 439 | 446 | 0.0022422 | 0.0022447 | 3.8343e-05 | 0.98455 | 871 | 0.98234
16:30 | 455 | uS | 228 | ppm | 22.8 | C | 823 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 875 |  | 451 | 437 | 444 | 0.0022523 | 0.0022472 | 3.8386e-05 | 0.98013 | 872 | 0.98234
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 824 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 876 |  | 452 | 437 | 444.5 | 0.0022497 | 0.002251 | 3.8451e-05 | 0.98124 | 873 | 0.98234
16:30 | 455 | uS | 228 | ppm | 22.8 | C | 825 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 877 |  | 452 | 440 | 446 | 0.0022422 | 0.0022459 | 3.8365e-05 | 0.98455 | 874 | 0.98455
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 826 | 455 | 0.0021978 | 0.047063 | 434.7 | µS |  |  | 878 |  | 452 | 440 | 446 | 0.0022422 | 0.0022422 | 3.83e-05 | 0.98455 | 875 | 0.98013
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 827 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 879 |  | 453 | 440 | 446.5 | 0.0022396 | 0.0022409 | 3.8279e-05 | 0.98565 | 876 | 0.98124
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 828 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 880 |  | 452 | 440 | 446 | 0.0022422 | 0.0022409 | 3.8279e-05 | 0.98455 | 877 | 0.98455
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 829 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 881 |  | 452 | 441 | 446.5 | 0.0022396 | 0.0022409 | 3.8279e-05 | 0.98565 | 878 | 0.98455
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 830 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 882 |  | 451 | 441 | 446 | 0.0022422 | 0.0022409 | 3.8279e-05 | 0.98455 | 879 | 0.98565
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 831 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 883 |  | 453 | 442 | 447.5 | 0.0022346 | 0.0022384 | 3.8236e-05 | 0.98786 | 880 | 0.98455
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 832 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 884 |  | 451 | 442 | 446.5 | 0.0022396 | 0.0022371 | 3.8215e-05 | 0.98565 | 881 | 0.98565
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 833 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 885 |  | 452 | 441 | 446.5 | 0.0022396 | 0.0022396 | 3.8257e-05 | 0.98565 | 882 | 0.98455
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 834 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 886 |  | 452 | 441 | 446.5 | 0.0022396 | 0.0022396 | 3.8257e-05 | 0.98565 | 883 | 0.98786
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 835 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 887 |  | 452 | 443 | 447.5 | 0.0022346 | 0.0022371 | 3.8215e-05 | 0.98786 | 884 | 0.98565
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 836 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 888 |  | 453 | 443 | 448 | 0.0022321 | 0.0022334 | 3.8151e-05 | 0.98896 | 885 | 0.98565
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 837 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 889 |  | 452 | 443 | 447.5 | 0.0022346 | 0.0022334 | 3.8151e-05 | 0.98786 | 886 | 0.98565
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 838 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 890 |  | 453 | 443 | 448 | 0.0022321 | 0.0022334 | 3.8151e-05 | 0.98896 | 887 | 0.98786
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 839 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 891 |  | 451 | 444 | 447.5 | 0.0022346 | 0.0022334 | 3.8151e-05 | 0.98786 | 888 | 0.98896
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 840 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 892 |  | 453 | 444 | 448.5 | 0.0022297 | 0.0022321 | 3.8129e-05 | 0.99007 | 889 | 0.98786
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 841 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 893 |  | 452 | 444 | 448 | 0.0022321 | 0.0022309 | 3.8108e-05 | 0.98896 | 890 | 0.98896
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 842 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 894 |  | 453 | 445 | 449 | 0.0022272 | 0.0022297 | 3.8087e-05 | 0.99117 | 891 | 0.98786
16:30 | 458 | uS | 229 | ppm | 22.8 | C | 843 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 895 |  | 452 | 445 | 448.5 | 0.0022297 | 0.0022284 | 3.8066e-05 | 0.99007 | 892 | 0.99007
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 844 | 458 | 0.0021834 | 0.047078 | 437.7 | µS |  |  | 896 |  | 452 | 443 | 447.5 | 0.0022346 | 0.0022321 | 3.8129e-05 | 0.98786 | 893 | 0.98896
16:30 | 458 | uS | 229 | ppm | 22.8 | C | 845 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 897 |  | 453 | 443 | 448 | 0.0022321 | 0.0022334 | 3.8151e-05 | 0.98896 | 894 | 0.99117
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 846 | 458 | 0.0021834 | 0.047078 | 437.7 | µS |  |  | 898 |  | 453 | 444 | 448.5 | 0.0022297 | 0.0022309 | 3.8108e-05 | 0.99007 | 895 | 0.99007
16:30 | 458 | uS | 229 | ppm | 22.8 | C | 847 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 899 |  | 453 | 444 | 448.5 | 0.0022297 | 0.0022297 | 3.8087e-05 | 0.99007 | 896 | 0.98786
16:30 | 457 | uS | 229 | ppm | 22.8 | C | 848 | 458 | 0.0021834 | 0.047078 | 437.7 | µS |  |  | 900 |  | 451 | 444 | 447.5 | 0.0022346 | 0.0022321 | 3.8129e-05 | 0.98786 | 897 | 0.98896
16:30 | 458 | uS | 229 | ppm | 22.8 | C | 849 | 457 | 0.0021882 | 0.047073 | 436.7 | µS |  |  | 901 |  | 453 | 444 | 448.5 | 0.0022297 | 0.0022321 | 3.8129e-05 | 0.99007 | 898 | 0.99007
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 850 | 458 | 0.0021834 | 0.047078 | 437.7 | µS |  |  | 902 |  | 452 | 444 | 448 | 0.0022321 | 0.0022309 | 3.8108e-05 | 0.98896 | 899 | 0.99007
16:30 | 458 | uS | 229 | ppm | 22.8 | C | 851 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 903 |  | 453 | 444 | 448.5 | 0.0022297 | 0.0022309 | 3.8108e-05 | 0.99007 | 900 | 0.98786
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 852 | 458 | 0.0021834 | 0.047078 | 437.7 | µS |  |  | 904 |  |  | 445 | 445 | 0.0022472 | 0.0022384 | 3.8237e-05 | 0.98234 | 901 | 0.99007
16:30 | 458 | uS | 229 | ppm | 22.8 | C | 853 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 905 |  |  | 445 | 445 | 0.0022472 | 0.0022472 | 3.8386e-05 | 0.98234 | 902 | 0.98896
16:30 | 456 | uS | 228 | ppm | 22.8 | C | 854 | 458 | 0.0021834 | 0.047078 | 437.7 | µS |  |  | 906 |  |  | 443 | 443 | 0.0022573 | 0.0022523 | 3.8473e-05 | 0.97792 | 903 | 0.99007
16:30 | 458 | uS | 229 | ppm | 22.8 | C | 855 | 456 | 0.002193 | 0.047068 | 435.7 | µS |  |  | 907 |  |  | 443 | 443 | 0.0022573 | 0.0022573 | 3.856e-05 | 0.97792 | 904 | 0.98234
[7 empty] | 856 | 458 | 0.0021834 | 0.047078 | 437.7 | µS |  |  | 908 |  |  | 446 | 446 | 0.0022422 | 0.0022497 | 3.843e-05 | 0.98455 | 905 | 0.98234
[7 empty] | 95% | 435.1 | 0.0022983 | 0.046963 |  |  |  |  | 909 |  |  | 446 | 446 | 0.0022422 | 0.0022422 | 3.83e-05 | 0.98455 | 906 | 0.97792
[15 empty] | 910 |  |  | 446 | 446 | 0.0022422 | 0.0022422 | 3.83e-05 | 0.98455 | 907 | 0.97792
[15 empty] | 911 |  |  | 446 | 446 | 0.0022422 | 0.0022422 | 3.83e-05 | 0.98455 | 908 | 0.98455
[15 empty] | 912 |  |  | 447 | 447 | 0.0022371 | 0.0022396 | 3.8257e-05 | 0.98675 | 909 | 0.98455
[15 empty] | 913 |  |  | 447 | 447 | 0.0022371 | 0.0022371 | 3.8215e-05 | 0.98675 | 910 | 0.98455
[15 empty] | 914 |  |  | 448 | 448 | 0.0022321 | 0.0022346 | 3.8172e-05 | 0.98896 | 911 | 0.98455
[15 empty] | 915 |  |  | 448 | 448 | 0.0022321 | 0.0022321 | 3.8129e-05 | 0.98896 | 912 | 0.98675
[15 empty] | 916 |  |  | 448 | 448 | 0.0022321 | 0.0022321 | 3.8129e-05 | 0.98896 | 913 | 0.98675
[15 empty] | 917 |  |  | 448 | 448 | 0.0022321 | 0.0022321 | 3.8129e-05 | 0.98896 | 914 | 0.98896
[15 empty] | 918 |  |  | 450 | 450 | 0.0022222 | 0.0022272 | 3.8045e-05 | 0.99338 | 915 | 0.98896
[15 empty] | 919 |  |  | 450 | 450 | 0.0022222 | 0.0022222 | 3.796e-05 | 0.99338 | 916 | 0.98896
[15 empty] | 920 |  |  | 448 | 448 | 0.0022321 | 0.0022272 | 3.8045e-05 | 0.98896 | 917 | 0.98896
[15 empty] | 921 |  |  | 448 | 448 | 0.0022321 | 0.0022321 | 3.8129e-05 | 0.98896 | 918 | 0.99338
[15 empty] | 922 |  |  | 449 | 449 | 0.0022272 | 0.0022297 | 3.8087e-05 | 0.99117 | 919 | 0.99338
[15 empty] | 923 |  |  | 449 | 449 | 0.0022272 | 0.0022272 | 3.8044e-05 | 0.99117 | 920 | 0.98896
[15 empty] | 924 |  |  | 449 | 449 | 0.0022272 | 0.0022272 | 3.8044e-05 | 0.99117 | 921 | 0.98896
[15 empty] | 925 |  |  | 449 | 449 | 0.0022272 | 0.0022272 | 3.8044e-05 | 0.99117 | 922 | 0.99117
[15 empty] | 926 |  |  | 448 | 448 | 0.0022321 | 0.0022297 | 3.8087e-05 | 0.98896 | 923 | 0.99117
[15 empty] | 927 |  |  | 448 | 448 | 0.0022321 | 0.0022321 | 3.8129e-05 | 0.98896 | 924 | 0.99117
[15 empty] | 928 |  |  | 449 | 449 | 0.0022272 | 0.0022297 | 3.8087e-05 | 0.99117 | 925 | 0.99117
[15 empty] | 929 |  |  | 449 | 449 | 0.0022272 | 0.0022272 | 3.8044e-05 | 0.99117 | 926 | 0.98896
[15 empty] | 930 |  |  | 446 | 446 | 0.0022422 | 0.0022347 | 3.8172e-05 | 0.98455 | 927 | 0.98896
[15 empty] | 931 |  |  | 446 | 446 | 0.0022422 | 0.0022422 | 3.83e-05 | 0.98455 | 928 | 0.99117
[15 empty] | 932 |  |  | 448 | 448 | 0.0022321 | 0.0022371 | 3.8215e-05 | 0.98896 | 929 | 0.99117
[15 empty] | 933 |  |  | 448 | 448 | 0.0022321 | 0.0022321 | 3.8129e-05 | 0.98896 | 930 | 0.98455
[15 empty] | 934 |  |  | 448 | 448 | 0.0022321 | 0.0022321 | 3.8129e-05 | 0.98896 | 931 | 0.98455
[15 empty] | 935 |  |  | 448 | 448 | 0.0022321 | 0.0022321 | 3.8129e-05 | 0.98896 | 932 | 0.98896
[15 empty] | 936 |  |  | 449 | 449 | 0.0022272 | 0.0022297 | 3.8087e-05 | 0.99117 | 933 | 0.98896
[15 empty] | 937 |  |  | 449 | 449 | 0.0022272 | 0.0022272 | 3.8044e-05 | 0.99117 | 934 | 0.98896
[15 empty] | 938 |  |  | 450 | 450 | 0.0022222 | 0.0022247 | 3.8002e-05 | 0.99338 | 935 | 0.98896
[15 empty] | 939 |  |  | 450 | 450 | 0.0022222 | 0.0022222 | 3.796e-05 | 0.99338 | 936 | 0.99117
[15 empty] | 940 |  |  | 450 | 450 | 0.0022222 | 0.0022222 | 3.796e-05 | 0.99338 | 937 | 0.99117
[15 empty] | 941 |  |  | 450 | 450 | 0.0022222 | 0.0022222 | 3.796e-05 | 0.99338 | 938 | 0.99338
[15 empty] | 942 |  |  | 452 | 452 | 0.0022124 | 0.0022173 | 3.7876e-05 | 0.99779 | 939 | 0.99338
[15 empty] | 943 |  |  | 452 | 452 | 0.0022124 | 0.0022124 | 3.7792e-05 | 0.99779 | 940 | 0.99338
[15 empty] | 944 |  |  | 451 | 451 | 0.0022173 | 0.0022148 | 3.7834e-05 | 0.99558 | 941 | 0.99338
[15 empty] | 945 |  |  | 451 | 451 | 0.0022173 | 0.0022173 | 3.7876e-05 | 0.99558 | 942 | 0.99779
[15 empty] | 946 |  |  | 452 | 452 | 0.0022124 | 0.0022148 | 3.7834e-05 | 0.99779 | 943 | 0.99779
[15 empty] | 947 |  |  | 452 | 452 | 0.0022124 | 0.0022124 | 3.7792e-05 | 0.99779 | 944 | 0.99558
[15 empty] | 948 |  |  | 453 | 453 | 0.0022075 | 0.0022099 | 3.775e-05 | 1 | 945 | 0.99558
[15 empty] | 949 |  |  | 453 | 453 | 0.0022075 | 0.0022075 | 3.7708e-05 | 1 | 946 | 0.99779
[15 empty] | 950 |  |  | 453 | 453 | 0.0022075 | 0.0022075 | 3.7708e-05 | 1 | 947 | 0.99779
[15 empty] | 951 |  |  | 453 | 453 | 0.0022075 | 0.0022075 | 3.7708e-05 | 1 | 948 | 1
[15 empty] | 952 |  |  | 454 | 454 | 0.0022026 | 0.0022051 | 3.7667e-05 | 1.0022 | 949 | 1
[15 empty] | 953 |  |  | 454 | 454 | 0.0022026 | 0.0022026 | 3.7625e-05 | 1.0022 | 950 | 1
[15 empty] | 954 |  |  | 452 | 452 | 0.0022124 | 0.0022075 | 3.7709e-05 | 0.99779 | 951 | 1
[15 empty] | 955 |  |  | 452 | 452 | 0.0022124 | 0.0022124 | 3.7792e-05 | 0.99779 | 952 | 1.0022
[15 empty] | 956 |  |  | 453 | 453 | 0.0022075 | 0.0022099 | 3.775e-05 | 1 | 953 | 1.0022
[15 empty] | 957 |  |  | 453 | 453 | 0.0022075 | 0.0022075 | 3.7708e-05 | 1 | 954 | 0.99779
[15 empty] | 958 |  |  | 453 | 453 | 0.0022075 | 0.0022075 | 3.7708e-05 | 1 | 955 | 0.99779
[15 empty] | 959 |  |  | 453 | 453 | 0.0022075 | 0.0022075 | 3.7708e-05 | 1 | 956 | 1
[15 empty] | 960 |  |  | 452 | 452 | 0.0022124 | 0.0022099 | 3.775e-05 | 0.99779 | 957 | 1
[15 empty] | 961 |  |  | 452 | 452 | 0.0022124 | 0.0022124 | 3.7792e-05 | 0.99779 | 958 | 1
[15 empty] | 962 |  |  | 453 | 453 | 0.0022075 | 0.0022099 | 3.775e-05 | 1 | 959 | 1
[15 empty] | 963 |  |  | 453 | 453 | 0.0022075 | 0.0022075 | 3.7708e-05 | 1 | 960 | 0.99779
[19 empty] | 430.35 | 0.0023237 | 58.5415 |  |  | 961 | 0.99779
[24 empty] | 962 | 1
[24 empty] | 963 | 1

## 4
Mixing time through conductivity method and modelling | [9 empty] | 1000 | [5 empty] | stirring speed | immersion | Mixing time | Error | stirring speed | immersion | Mixing time | Error |  |  |  | Two-factor ANOVA analysis
| 4 % Impeller immersion |  |  |  | 42 % Impeller immersion |  |  |  | 79 % Impeller immersion |  |  |  | Impeller fully immersed |  |  | 100 rpm | 4% | 659.3333 | 181.9924 | 400 rpm | 4% | 84.6667 | 4.7428 |  |  |  | Experimental results of mixing time in triplicate
1 L | RPM | Time (s) | Error | 2 L | RPM | Time (s) | Error |  | RPM | Time (s) | Error |  | RPM | Time (s) | Error |  | 42% | 72 | 6.1524 |  | 42% | 20.6667 | 1.7759 |  |  |  |  | Impeller immersion | Volume | Speed
| 100 | 659.3333 | 181.9924 |  | 100 | 72 | 6.1524 | 3 L | 100 | 70.6667 | 1.9001 | 4 L | 100 | 135.3333 | 3.7784 |  | 79% | 70.6667 | 1.9001 |  | 79% | 17.6667 | 0.68968 | [6 empty] | 100 | 200 | 300 | 400 | 500
| 200 | 162 | 15.3609 |  | 200 | 41.3333 | 2.9352 |  | 200 | 24.6667 | 1.781 |  | 200 | 68.6667 | 2.986 |  | 121% | 135.3333 | 3.7784 |  | 121% | 35.3333 | 4.8204 |  |  |  |  | 4% | 1 | 576 | 157 | 128 | 89 | 85
| 300 | 136.6667 | 25.7872 |  | 300 | 27.6667 | 0.7218 |  | 300 | 19.6667 | 1.3478 |  | 300 | 46 | 1.243 | 200 rpm | 4% | 162 | 15.3609 | 500 rpm | 4% | 81.6667 | 4.1366 | [6 empty] | 561 | 177 | 162 | 81 | 82
| 400 | 84.6667 | 4.7428 |  | 400 | 20.6667 | 1.7759 |  | 400 | 17.6667 | 0.68968 |  | 400 | 35.3333 | 4.8204 |  | 42% | 41.3333 | 2.9352 |  | 42% | 20.3333 | 2.6744 | [6 empty] | 841 | 152 | 120 | 84 | 78
| 500 | 81.6667 | 4.1366 |  | 500 | 20.3333 | 2.6744 |  | 500 | 16.3333 | 2.6717 |  | 500 | 33.3333 | 0.74536 |  | 79% | 24.6667 | 1.781 |  | 79% | 16.3333 | 2.6717 |  |  |  |  | 42% | 2 | 74 | 44 | 27 | 21 | 23
|  | Impeller speed | [14 empty] | 121% | 68.6667 | 2.986 |  | 121% | 33.3333 | 0.74536
Experimental result matrix |  | 100 | 200 | 300 | 400 | 500 |  |  |  | Legends | [5 empty] | 300 rpm | 4% | 136.6667 | 25.7872 | [10 empty] | 76 | 41 | 28 | 19 | 19
Volume (L) | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  |  | 100 rpm | 100 rpm model |  | 4 % model |  |  |  | 42% | 27.6667 | 0.7218 | [10 empty] | 66 | 39 | 28 | 22 | 19
1 | 4% | 659.3333 | 162 | 136.6667 | 84.6667 | 81.6667 | 224.8667 |  |  | 200 rpm | 200 rpm model |  | 42 % model |  |  |  | 79% | 19.6667 | 1.3478 | [8 empty] | 79% | 3 | 69 | 23 | 17 | 19 | 23
2 | 42% | 72 | 41.3333 | 27.6667 | 20.6667 | 20.3333 | 36.4 |  |  | 300 rpm | 300 rpm model |  | 79 % model |  |  |  | 121% | 46 | 1.243 | [10 empty] | 71 | 26 | 18 | 15 | 19
3 | 79% | 70.6667 | 24.6667 | 19.6667 | 17.6667 | 16.3333 | 29.8 |  |  | 400 rpm | 400 rpm model |  | 121 % model | [16 empty] | 72 | 25 | 18 | 15 | 19
4 | 1.21 | 135.3333 | 68.6667 | 46 | 35.3333 | 33.3333 | 63.7333 |  |  | 500 rpm | 500 rpm model | [16 empty] | Fully immersed | 4 | 138 | 71 | 46 | 40 | 33
|  | 234.3333 | 74.1667 | 57.5 | 39.5833 | 37.9167
[30 empty] | 136 | 66 | 47 | 32 | 34
Model 1 |  |  | Nq | 0.19 |  |  |  |  | Model 2 | [5 empty] | Parameters |  |  | Model 3 |  |  |  | Parameters | [7 empty] | 132 | 69 | 45 | 34 | 33
|  |  | L | 0.5 | [10 empty] | a | 1.0002 | [5 empty] | K | 160944.5535
|  |  | D | 0.11 | [10 empty] | b | 0.9788 | [5 empty] | a | 0.5051 |  |  |  |  | Anova: Two-Factor With Replication
|  |  | d | 0.05 | [10 empty] | c | 0.45246 | [5 empty] | b | 0.05984
Parameters |  |  | L/D | 4.5455 | [10 empty] | d | 0.18574 | [5 empty] | c | -0.88952 |  |  |  |  | SUMMARY | 100 | 200 | 300 | 400 | 500 | Total
a | 1.1478 |  | D/d | 2.2 | [13 empty] | ρ | 993.33 | Kg m-3 |  | d | -1.1922 |  |  | Figure 4-13 |  | 1
b | 0.45705 |  | K | 152.0531 | [13 empty] | µ | 0.0006915 | [8 empty] | Count | 3 | 3 | 3 | 3 | 3 | 15
Nq | [8 empty] | Nq | [9 empty] | 600 | as per Tsui and Hu (2008) | [7 empty] | Sum | 1978 | 486 | 410 | 254 | 245 | 3373
|  | Speed | [8 empty] | Speed | [16 empty] | Average | 659.3333 | 162 | 136.6667 | 84.6667 | 81.6667 | 224.8667
Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  | Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  | Reynolds number (Nre) | [9 empty] | Variance | 24808.3333 | 175 | 497.3333 | 16.3333 | 12.3333 | 55213.1238
1 | 0.04 | 0.02912 | 0.059259 | 0.046829 | 0.056693 | 0.04702 |  |  | 1 | 0.04 | 0.02912 | 0.059259 | 0.046829 | 0.056693 | 0.04702 |  |  | Speed (N)
2 | 0.42 | 0.53333 | 0.46452 | 0.46265 | 0.46452 | 0.3777 |  |  | 2 | 0.42 | 0.53333 | 0.46452 | 0.46265 | 0.46452 | 0.3777 |  |  | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 | [5 empty] | 2
3 | 0.79 | 0.81509 | 1.1676 | 0.97627 | 0.81509 | 0.70531 |  |  | 3 | 0.79 | 0.81509 | 1.1676 | 0.97627 | 0.81509 | 0.70531 |  |  | 5985.3579 | 11970.7158 | 17956.0738 | 23941.4317 | 29926.7896 | [5 empty] | Count | 3 | 3 | 3 | 3 | 3 | 15
4 | 1.21 | 0.56749 | 0.55922 | 0.55652 | 0.5434 | 0.4608 |  |  | 4 | 1.21 | 0.56749 | 0.55922 | 0.55652 | 0.5434 | 0.4608 | [12 empty] | Sum | 216 | 124 | 83 | 62 | 61 | 546
[18 empty] | Np | [9 empty] | Average | 72 | 41.3333 | 27.6667 | 20.6667 | 20.3333 | 36.4
Predicted values for model 1 | [8 empty] | Predicted values for model 2 | [10 empty] | Speed | [7 empty] | Variance | 28 | 6.3333 | 0.33333 | 2.3333 | 5.3333 | 407.5429
|  | Speed | [8 empty] | Speed | [6 empty] | Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333
Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  | Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  | 1 | 0.04 | 0.0029192 | 0.0029702 | 0.0015648 | 0.0014208 | 0.00094271 |  |  |  | 3
1 | 0.04 | 667.1702 | 147.9654 | 117.5668 | 69.8025 | 65.1451 |  |  | 1 | 0.04 | 659.3373 | 161.9841 | 136.645 | 84.6496 | 81.6474 |  |  | 2 | 0.42 | 0.053464 | 0.023283 | 0.015459 | 0.011641 | 0.0075726 |  |  |  | Count | 3 | 3 | 3 | 3 | 3 | 15
2 | 0.42 | 106.7011 | 55.2905 | 34.8565 | 24.9536 | 23.7548 |  |  | 2 | 0.42 | 72.4394 | 41.5812 | 27.8309 | 20.7885 | 20.4525 |  |  | 3 | 0.79 | 0.081709 | 0.058521 | 0.032622 | 0.020427 | 0.014141 |  |  |  | Sum | 212 | 74 | 53 | 49 | 61 | 449
3 | 0.79 | 93.1889 | 29.3612 | 22.0481 | 18.9815 | 16.9797 |  |  | 3 | 0.79 | 69.6427 | 24.3067 | 19.3785 | 17.4071 | 16.0928 |  |  | 4 | 1.21 | 0.056888 | 0.02803 | 0.018596 | 0.013618 | 0.0092385 |  |  |  | Average | 70.6667 | 24.6667 | 17.6667 | 16.3333 | 20.3333 | 29.9333
4 | 1.21 | 162.6455 | 74.4898 | 46.9987 | 34.5978 | 31.5807 |  |  | 4 | 1.21 | 135.576 | 68.7826 | 46.0749 | 35.3893 | 33.385 | [12 empty] | Variance | 2.3333 | 2.3333 | 0.33333 | 5.3333 | 5.3333 | 455.3524
Predicted values for model 3
Sum of squares error | [8 empty] | Sum of squares error | [10 empty] | Speed | [7 empty] | 4
|  | Speed | [8 empty] | Speed | [6 empty] | Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  |  | Count | 3 | 3 | 3 | 3 | 3 | 15
Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  | Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  | 1 | 0.04 | 660.8191 | 153.0711 | 135.9848 | 83.3449 | 83.3436 |  |  |  | Sum | 406 | 206 | 138 | 106 | 100 | 956
1 | 0.04 | 61.4164 | 196.9701 | 364.8034 | 220.9426 | 272.9605 | 1117.093 |  | 1 | 0.04 | 1.61e-05 | 0.00025192 | 0.00047125 | 0.00029257 | 0.00036995 | 0.0014018 |  | 2 | 0.42 | 80.5253 | 48.7624 | 33.3342 | 25.2522 | 25.8248 |  |  |  | Average | 135.3333 | 68.6667 | 46 | 35.3333 | 33.3333 | 63.7333
2 | 0.42 | 1204.1648 | 194.8017 | 51.6937 | 18.3781 | 11.7061 | 1480.7445 |  | 2 | 0.42 | 0.19304 | 0.061456 | 0.026975 | 0.014832 | 0.014194 | 0.3105 |  | 3 | 0.79 | 68.5376 | 23.627 | 19.6894 | 18.3804 | 17.5186 |  |  |  | Variance | 9.3333 | 6.3333 | 1 | 17.3333 | 0.33333 | 1546.9238
3 | 0.79 | 507.2499 | 22.0384 | 5.6711 | 1.7289 | 0.41779 | 537.1061 |  | 3 | 0.79 | 1.0485 | 0.12957 | 0.083042 | 0.067401 | 0.05787 | 1.3864 |  | 4 | 1.21 | 128.0911 | 67.4445 | 46.1481 | 36.081 | 35.1851
4 | 1.21 | 745.9559 | 33.909 | 0.99739 | 0.54094 | 3.0717 | 784.4749 |  | 4 | 1.21 | 0.058875 | 0.013446 | 0.0056061 | 0.0031341 | 0.0026725 | 0.083733 | [11 empty] | Total
|  | 2518.787 | 447.7192 | 423.1657 | 241.5905 | 288.1562 | 3919.4185 |  |  |  | 1.3005 | 0.20472 | 0.11609 | 0.08566 | 0.075106 | 1.782 |  | Sum of squares error | [9 empty] | Count | 12 | 12 | 12 | 12 | 12
Errors | [8 empty] | Errors | [10 empty] | Speed | [7 empty] | Sum | 2812 | 890 | 684 | 471 | 467
|  | Speed | [8 empty] | Speed | [6 empty] | Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  |  | Average | 234.3333 | 74.1667 | 57 | 39.25 | 38.9167
Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  | Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  | 1 | 0.04 | 2.2075 | 79.7257 | 0.46491 | 1.7471 | 2.8121 | 86.9573 |  |  | Variance | 70944.6061 | 3109.0606 | 2511.2727 | 811.6591 | 699.5379
1 | 0.04 | 1.189% | 8.663% | 13.98% | 17.56% | 20.23% | 12.32% |  | 1 | 0.04 | 0.0006086% | 0.009797% | 0.01588% | 0.0202% | 0.02355% | 0.01401% |  | 2 | 0.42 | 72.6811 | 55.1916 | 32.1209 | 21.0269 | 30.1562 | 211.1767
2 | 0.42 | 48.2% | 33.77% | 25.99% | 20.74% | 16.83% | 29.1% |  | 2 | 0.42 | 0.6102% | 0.5998% | 0.5936% | 0.5893% | 0.5859% | 0.5958% |  | 3 | 0.79 | 4.5329 | 1.0808 | 0.00051732 | 0.50946 | 1.4049 | 7.5286
3 | 0.79 | 31.87% | 19.03% | 12.11% | 7.443% | 3.957% | 14.88% |  | 3 | 0.79 | 1.449% | 1.459% | 1.465% | 1.47% | 1.473% | 1.463% |  | 4 | 1.21 | 52.4505 | 1.4936 | 0.02193 | 0.55907 | 3.4289 | 57.954 |  |  | ANOVA
4 | 1.21 | 20.18% | 8.48% | 2.171% | 2.082% | 5.258% | 7.634% |  | 4 | 1.21 | 0.1793% | 0.1689% | 0.1628% | 0.1584% | 0.1551% | 0.1649% |  |  |  | 131.872 | 137.4917 | 32.6082 | 23.8425 | 37.8022 | 363.6166 |  |  | Source of Variation | SS | df | MS | F | P-value | F crit
|  | 25.36% | 17.49% | 13.56% | 11.96% | 11.57% | 15.99% |  |  |  | 0.5598% | 0.5594% | 0.5594% | 0.5594% | 0.5593% | 0.5595% |  | Errors | [9 empty] | Sample | 380302.5333 | 3 | 126767.5111 | 99.0307 | 1.4831e-18 | 2.8387
[20 empty] | Speed | [7 empty] | Columns | 328186.2333 | 4 | 82046.5583 | 64.0947 | 7.3574e-17 | 2.606
Residual plot wrt coverage | [8 empty] | Residual plot wrt coverage | [8 empty] | Volume | Impeller coverage | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  |  | Interaction | 427331.6333 | 12 | 35610.9694 | 27.8193 | 1.2016e-15 | 2.0035
| Residuals | [8 empty] | Residuals | [7 empty] | 1 | 0.04 | 0.2253% | 5.512% | 0.4989% | 1.561% | 2.053% | 1.97% |  |  | Within | 51203.3333 | 40 | 1280.0833
Impeller coverage | 100 | 200 | 300 | 400 | 500 |  |  |  | Impeller coverage | 100 | 200 | 300 | 400 | 500 |  |  |  | 2 | 0.42 | 11.84% | 17.97% | 20.49% | 22.19% | 27.01% | 19.9%
| 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  |  |  | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333 |  |  |  | 3 | 0.79 | 3.013% | 4.215% | 0.1157% | 4.04% | 7.257% | 3.728% |  |  | Total | 1187023.7333 | 59
4% | -7.8369 | 14.0346 | 19.0998 | 14.8641 | 16.5215 |  |  |  | 4% | -0.0040125 | 0.015872 | 0.021708 | 0.017105 | 0.019234 |  |  |  | 4 | 1.21 | 5.351% | 1.78% | 0.3219% | 2.116% | 5.555% | 3.025%
42% | -34.7011 | -13.9571 | -7.1898 | -4.287 | -3.4214 |  |  |  | 42% | -0.43937 | -0.2479 | -0.16424 | -0.12179 | -0.11914 | [5 empty] | 5.108% | 7.37% | 5.355% | 7.476% | 10.47% | 7.155%
79% | -22.5222 | -4.6945 | -2.3814 | -1.3149 | -0.64636 |  |  |  | 79% | 1.024 | 0.35995 | 0.28817 | 0.25962 | 0.24056
121% | -27.3122 | -5.8231 | -0.9987 | 0.73549 | 1.7526 |  |  |  | 121% | -0.24264 | -0.11596 | -0.074874 | -0.055983 | -0.051696 |  |  |  | Residual plot wrt coverage
Residuals
[18 empty] | Impeller coverage | 100 | 200 | 300 | 400 | 500
[19 empty] | 1.6667 | 3.3333 | 5 | 6.6667 | 8.3333
[18 empty] | 4% | -1.4858 | 8.9289 | 0.68185 | 1.3218 | -1.6769
[18 empty] | 42% | -8.5253 | -7.4291 | -5.6675 | -4.5855 | -5.4915
[18 empty] | 79% | 2.1291 | 1.0396 | -0.022745 | -0.71376 | -1.1853
[18 empty] | 121% | 7.2423 | 1.2221 | -0.14809 | -0.74771 | -1.8517
R-squared calculation with respect to 'impeller immersion' | [8 empty] | R-squared calculation with respect to 'impeller immersion'
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
100 rpm | hL,d | tmix | [6 empty] | 100 rpm | hL,d | tmix
mean y | 0.04 | 659.3333 | 667.1702 | 61.4164 | 180625 |  |  |  | mean y | 0.04 | 659.3333 | 659.3373 | 1.61e-05 | 180625
234.3333 | 0.42 | 72 | 106.7011 | 1204.1648 | 26352.1111 |  |  |  | 234.3333 | 0.42 | 72 | 72.4394 | 0.19304 | 26352.1111
| 0.79 | 70.6667 | 93.1889 | 507.2499 | 26786.7778 |  |  |  |  | 0.79 | 70.6667 | 69.6427 | 1.0485 | 26786.7778
| 1.21 | 135.3333 | 162.6455 | 745.9559 | 9801 |  |  |  |  | 1.21 | 135.3333 | 135.576 | 0.058875 | 9801
|  |  | SUM | 2518.787 | 243564.8889 | [6 empty] | SUM | 1.3005 | 243564.8889 |  |  |  | R-squared calculation with respect to 'impeller immersion'
|  |  |  | variation | 0.010341 | [7 empty] | variation | 5.3393e-06 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  |  |  | R-square | 0.98966 | [7 empty] | R-square | 0.99999 |  |  |  | 100 rpm | hL,d | tmix
[18 empty] | mean y | 0.04 | 659.3333 | 660.8191 | 2.2075 | 180625
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  | 234.3333 | 0.42 | 72 | 80.5253 | 72.6811 | 26352.1111
200 rpm | hL,d | tmix | [6 empty] | 200 rpm | hL,d | tmix | [7 empty] | 0.79 | 70.6667 | 68.5376 | 4.5329 | 26786.7778
mean y | 0.04 | 162 | 147.9654 | 196.9701 | 7714.6944 |  |  |  | mean y | 0.04 | 162 | 161.9841 | 0.00025192 | 7714.6944 |  |  |  |  | 1.21 | 135.3333 | 128.0911 | 52.4505 | 9801
74.1667 | 0.42 | 41.3333 | 55.2905 | 194.8017 | 1078.0278 |  |  |  | 74.1667 | 0.42 | 41.3333 | 41.5812 | 0.061456 | 1078.0278 | [6 empty] | SUM | 131.872 | 243564.8889
| 0.79 | 24.6667 | 29.3612 | 22.0384 | 2450.25 |  |  |  |  | 0.79 | 24.6667 | 24.3067 | 0.12957 | 2450.25 | [7 empty] | variation | 0.00054142
| 1.21 | 68.6667 | 74.4898 | 33.909 | 30.25 |  |  |  |  | 1.21 | 68.6667 | 68.7826 | 0.013446 | 30.25 | [7 empty] | R-square | 0.99946
|  |  |  | 447.7192 | 11273.2222 | [7 empty] | 0.20472 | 11273.2222
|  |  |  | variation | 0.039715 | [7 empty] | variation | 1.816e-05 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  |  |  | R-square | 0.96028 | [7 empty] | R-square | 0.99998 |  |  |  | 200 rpm | hL,d | tmix
[18 empty] | mean y | 0.04 | 162 | 153.0711 | 79.7257 | 7714.6944
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  | 74.1667 | 0.42 | 41.3333 | 48.7624 | 55.1916 | 1078.0278
300 rpm | hL,d | tmix | [6 empty] | 300 rpm | hL,d | tmix | [7 empty] | 0.79 | 24.6667 | 23.627 | 1.0808 | 2450.25
mean y | 0.04 | 136.6667 | 117.5668 | 364.8034 | 6267.3611 |  |  |  | mean y | 0.04 | 136.6667 | 136.645 | 0.00047125 | 6267.3611 |  |  |  |  | 1.21 | 68.6667 | 67.4445 | 1.4936 | 30.25
57.5 | 0.42 | 27.6667 | 34.8565 | 51.6937 | 890.0278 |  |  |  | 57.5 | 0.42 | 27.6667 | 27.8309 | 0.026975 | 890.0278 | [7 empty] | 137.4917 | 11273.2222
| 0.79 | 19.6667 | 22.0481 | 5.6711 | 1431.3611 |  |  |  |  | 0.79 | 19.6667 | 19.3785 | 0.083042 | 1431.3611 | [7 empty] | variation | 0.012196
| 1.21 | 46 | 46.9987 | 0.99739 | 132.25 |  |  |  |  | 1.21 | 46 | 46.0749 | 0.0056061 | 132.25 | [7 empty] | R-square | 0.9878
|  |  |  | 423.1657 | 8721 | [7 empty] | 0.11609 | 8721
|  |  |  | variation | 0.048523 | [7 empty] | variation | 1.3312e-05 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  |  |  | R-square | 0.95148 | [7 empty] | R-square | 0.99999 |  |  |  | 300 rpm | hL,d | tmix
[18 empty] | mean y | 0.04 | 136.6667 | 135.9848 | 0.46491 | 6267.3611
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  | 57.5 | 0.42 | 27.6667 | 33.3342 | 32.1209 | 890.0278
400 rpm | hL,d | tmix | [6 empty] | 400 rpm | hL,d | tmix | [7 empty] | 0.79 | 19.6667 | 19.6894 | 0.00051732 | 1431.3611
mean y | 0.04 | 84.6667 | 69.8025 | 220.9426 | 2032.5069 |  |  |  | mean y | 0.04 | 84.6667 | 84.6496 | 0.00029257 | 2032.5069 |  |  |  |  | 1.21 | 46 | 46.1481 | 0.02193 | 132.25
39.5833 | 0.42 | 20.6667 | 24.9536 | 18.3781 | 357.8403 |  |  |  | 39.5833 | 0.42 | 20.6667 | 20.7885 | 0.014832 | 357.8403 | [7 empty] | 32.6082 | 8721
| 0.79 | 17.6667 | 18.9815 | 1.7289 | 480.3403 |  |  |  |  | 0.79 | 17.6667 | 17.4071 | 0.067401 | 480.3403 | [7 empty] | variation | 0.003739
| 1.21 | 35.3333 | 34.5978 | 0.54094 | 18.0625 |  |  |  |  | 1.21 | 35.3333 | 35.3893 | 0.0031341 | 18.0625 | [7 empty] | R-square | 0.99626
|  |  |  | 241.5905 | 2888.75 | [7 empty] | 0.08566 | 2888.75
|  |  |  | variation | 0.083632 | [7 empty] | variation | 2.9653e-05 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  |  |  | R-square | 0.91637 | [7 empty] | R-square | 0.99997 |  |  |  | 400 rpm | hL,d | tmix
[18 empty] | mean y | 0.04 | 84.6667 | 83.3449 | 1.7471 | 2032.5069
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  | 39.5833 | 0.42 | 20.6667 | 25.2522 | 21.0269 | 357.8403
500 rpm | hL,d | tmix | [6 empty] | 500 rpm | hL,d | tmix | [7 empty] | 0.79 | 17.6667 | 18.3804 | 0.50946 | 480.3403
mean y | 0.04 | 81.6667 | 65.1451 | 272.9605 | 1914.0625 |  |  |  | mean y | 0.04 | 81.6667 | 81.6474 | 0.00036995 | 1914.0625 |  |  |  |  | 1.21 | 35.3333 | 36.081 | 0.55907 | 18.0625
37.9167 | 0.42 | 20.3333 | 23.7548 | 11.7061 | 309.1736 |  |  |  | 37.9167 | 0.42 | 20.3333 | 20.4525 | 0.014194 | 309.1736 | [7 empty] | 23.8425 | 2888.75
| 0.79 | 16.3333 | 16.9797 | 0.41779 | 465.8403 |  |  |  |  | 0.79 | 16.3333 | 16.0928 | 0.05787 | 465.8403 | [7 empty] | variation | 0.0082536
| 1.21 | 33.3333 | 31.5807 | 3.0717 | 21.0069 |  |  |  |  | 1.21 | 33.3333 | 33.385 | 0.0026725 | 21.0069 | [7 empty] | R-square | 0.99175
|  |  |  | 288.1562 | 2710.0833 | [7 empty] | 0.075106 | 2710.0833
|  |  |  | variation | 0.10633 | [7 empty] | variation | 2.7714e-05 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  |  |  | R-square | 0.89367 | [7 empty] | R-square | 0.99997 |  |  |  | 500 rpm | hL,d | tmix
[18 empty] | mean y | 0.04 | 81.6667 | 83.3436 | 2.8121 | 1914.0625
R-squared calculation with respect to 'impeller speed' | [8 empty] | R-squared calculation with respect to 'impeller speed' | [8 empty] | 37.9167 | 0.42 | 20.3333 | 25.8248 | 30.1562 | 309.1736
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | 0.79 | 16.3333 | 17.5186 | 1.4049 | 465.8403
1 L | RPM | tmix | [6 empty] | 1 L | RPM | tmix | [7 empty] | 1.21 | 33.3333 | 35.1851 | 3.4289 | 21.0069
mean y | 100 | 659.3333 | 667.1702 | 61.4164 | 188761.2844 |  |  |  | mean y | 100 | 659.3333 | 659.3373 | 1.61e-05 | 188761.2844 | [7 empty] | 37.8022 | 2710.0833
224.8667 | 200 | 162 | 147.9654 | 196.9701 | 3952.2178 |  |  |  | 224.8667 | 200 | 162 | 161.9841 | 0.00025192 | 3952.2178 | [7 empty] | variation | 0.013949
| 300 | 136.6667 | 117.5668 | 364.8034 | 7779.24 |  |  |  |  | 300 | 136.6667 | 136.645 | 0.00047125 | 7779.24 | [7 empty] | R-square | 0.98605
| 400 | 84.6667 | 69.8025 | 220.9426 | 19656.04 |  |  |  |  | 400 | 84.6667 | 84.6496 | 0.00029257 | 19656.04
| 500 | 81.6667 | 65.1451 | 272.9605 | 20506.24 |  |  |  |  | 500 | 81.6667 | 81.6474 | 0.00036995 | 20506.24 |  |  |  | R-squared calculation with respect to 'impeller speed'
|  |  | SUM | 1117.093 | 240655.0222 | [6 empty] | SUM | 0.0014018 | 240655.0222 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  |  |  | variation | 0.0046419 | [7 empty] | variation | 5.8249e-09 |  |  |  | 1 L | RPM | tmix
|  |  |  | R-square | 0.99536 | [7 empty] | R-square | 1 |  |  |  | mean y | 100 | 659.3333 | 660.8191 | 2.2075 | 188761.2844
[18 empty] | 224.8667 | 200 | 162 | 153.0711 | 79.7257 | 3952.2178
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | 300 | 136.6667 | 135.9848 | 0.46491 | 7779.24
2 L | RPM | tmix | [6 empty] | 2 L | RPM | tmix | [7 empty] | 400 | 84.6667 | 83.3449 | 1.7471 | 19656.04
mean y | 100 | 72 | 106.7011 | 1204.1648 | 1267.36 |  |  |  | mean y | 100 | 72 | 72.4394 | 0.19304 | 1267.36 |  |  |  |  | 500 | 81.6667 | 83.3436 | 2.8121 | 20506.24
36.4 | 200 | 41.3333 | 55.2905 | 194.8017 | 24.3378 |  |  |  | 36.4 | 200 | 41.3333 | 41.5812 | 0.061456 | 24.3378 | [6 empty] | SUM | 86.9573 | 240655.0222
| 300 | 27.6667 | 34.8565 | 51.6937 | 76.2711 |  |  |  |  | 300 | 27.6667 | 27.8309 | 0.026975 | 76.2711 | [7 empty] | variation | 0.00036134
| 400 | 20.6667 | 24.9536 | 18.3781 | 247.5378 |  |  |  |  | 400 | 20.6667 | 20.7885 | 0.014832 | 247.5378 | [7 empty] | R-square | 0.99964
| 500 | 20.3333 | 23.7548 | 11.7061 | 258.1378 |  |  |  |  | 500 | 20.3333 | 20.4525 | 0.014194 | 258.1378
|  |  | SUM | 1480.7445 | 1873.6444 | [6 empty] | SUM | 0.3105 | 1873.6444 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  |  |  | variation | 0.7903 | [7 empty] | variation | 0.00016572 |  |  |  | 2 L | RPM | tmix
|  |  |  | R-square | 0.2097 | [7 empty] | R-square | 0.99983 |  |  |  | mean y | 100 | 72 | 80.5253 | 72.6811 | 1267.36
[18 empty] | 36.4 | 200 | 41.3333 | 48.7624 | 55.1916 | 24.3378
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | 300 | 27.6667 | 33.3342 | 32.1209 | 76.2711
3 L | RPM | tmix | [6 empty] | 3 L | RPM | tmix | [7 empty] | 400 | 20.6667 | 25.2522 | 21.0269 | 247.5378
mean y | 100 | 70.6667 | 93.1889 | 507.2499 | 1670.0844 |  |  |  | mean y | 100 | 70.6667 | 69.6427 | 1.0485 | 1670.0844 |  |  |  |  | 500 | 20.3333 | 25.8248 | 30.1562 | 258.1378
29.8 | 200 | 24.6667 | 29.3612 | 22.0384 | 26.3511 |  |  |  | 29.8 | 200 | 24.6667 | 24.3067 | 0.12957 | 26.3511 | [6 empty] | SUM | 211.1767 | 1873.6444
| 300 | 19.6667 | 22.0481 | 5.6711 | 102.6844 |  |  |  |  | 300 | 19.6667 | 19.3785 | 0.083042 | 102.6844 | [7 empty] | variation | 0.11271
| 400 | 17.6667 | 18.9815 | 1.7289 | 147.2178 |  |  |  |  | 400 | 17.6667 | 17.4071 | 0.067401 | 147.2178 | [7 empty] | R-square | 0.88729
| 500 | 16.3333 | 16.9797 | 0.41779 | 181.3511 |  |  |  |  | 500 | 16.3333 | 16.0928 | 0.05787 | 181.3511
|  |  | SUM | 537.1061 | 2127.6889 | [6 empty] | SUM | 1.3864 | 2127.6889 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  |  |  | variation | 0.25244 | [7 empty] | variation | 0.0006516 |  |  |  | 3 L | RPM | tmix
|  |  |  | R-square | 0.74756 | [7 empty] | R-square | 0.99935 |  |  |  | mean y | 100 | 70.6667 | 68.5376 | 4.5329 | 1670.0844
[18 empty] | 29.8 | 200 | 24.6667 | 23.627 | 1.0808 | 26.3511
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | 300 | 19.6667 | 19.6894 | 0.00051732 | 102.6844
4 L | RPM | tmix | [6 empty] | 4 L | RPM | tmix | [7 empty] | 400 | 17.6667 | 18.3804 | 0.50946 | 147.2178
mean y | 100 | 135.3333 | 162.6455 | 745.9559 | 5126.56 |  |  |  | mean y | 100 | 135.3333 | 135.576 | 0.058875 | 5126.56 |  |  |  |  | 500 | 16.3333 | 17.5186 | 1.4049 | 181.3511
63.7333 | 200 | 68.6667 | 74.4898 | 33.909 | 24.3378 |  |  |  | 63.7333 | 200 | 68.6667 | 68.7826 | 0.013446 | 24.3378 | [6 empty] | SUM | 7.5286 | 2127.6889
| 300 | 46 | 46.9987 | 0.99739 | 314.4711 |  |  |  |  | 300 | 46 | 46.0749 | 0.0056061 | 314.4711 | [7 empty] | variation | 0.0035384
| 400 | 35.3333 | 34.5978 | 0.54094 | 806.56 |  |  |  |  | 400 | 35.3333 | 35.3893 | 0.0031341 | 806.56 | [7 empty] | R-square | 0.99646
| 500 | 33.3333 | 31.5807 | 3.0717 | 924.16 |  |  |  |  | 500 | 33.3333 | 33.385 | 0.0026725 | 924.16
|  |  | SUM | 784.4749 | 7196.0889 | [6 empty] | SUM | 0.083733 | 7196.0889 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  |  |  | variation | 0.10901 | [7 empty] | variation | 1.1636e-05 |  |  |  | 4 L | RPM | tmix
|  |  |  | R-square | 0.89099 | [7 empty] | R-square | 0.99999 |  |  |  | mean y | 100 | 135.3333 | 128.0911 | 52.4505 | 5126.56
[18 empty] | 63.7333 | 200 | 68.6667 | 67.4445 | 1.4936 | 24.3378
Overall R-squared calculation | [8 empty] | Overall R-squared calculation | [9 empty] | 300 | 46 | 46.1481 | 0.02193 | 314.4711
| x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2 |  |  |  |  | 400 | 35.3333 | 36.081 | 0.55907 | 806.56
mean y |  | tmix | [6 empty] | mean y |  | tmix | [7 empty] | 500 | 33.3333 | 35.1851 | 3.4289 | 924.16
88.7 |  | 659.3333 | 667.1702 | 61.4164 | 434720.4444 |  |  |  | 88.7 |  | 659.3333 | 659.3373 | 1.61e-05 | 325622.4011 | [6 empty] | SUM | 57.954 | 7196.0889
|  | 72 | 106.7011 | 1204.1648 | 5184 | [5 empty] | 72 | 72.4394 | 0.19304 | 278.89 | [7 empty] | variation | 0.0080535
|  | 70.6667 | 93.1889 | 507.2499 | 4993.7778 | [5 empty] | 70.6667 | 69.6427 | 1.0485 | 325.2011 | [7 empty] | R-square | 0.99195
|  | 135.3333 | 162.6455 | 745.9559 | 18315.1111 | [5 empty] | 135.3333 | 135.576 | 0.058875 | 2174.6678
|  | 162 | 147.9654 | 196.9701 | 26244 | [5 empty] | 162 | 161.9841 | 0.00025192 | 5372.89 |  |  |  | Overall R-squared calculation
|  | 41.3333 | 55.2905 | 194.8017 | 1708.4444 | [5 empty] | 41.3333 | 41.5812 | 0.061456 | 2243.6011 |  |  |  |  | x | y | y_est | (y-y_est)^2 | (y-y_mean)^2
|  | 24.6667 | 29.3612 | 22.0384 | 608.4444 | [5 empty] | 24.6667 | 24.3067 | 0.12957 | 4100.2678 |  |  |  | mean y |  | tmix
|  | 68.6667 | 74.4898 | 33.909 | 4715.1111 | [5 empty] | 68.6667 | 68.7826 | 0.013446 | 401.3344 |  |  |  | 88.7 |  | 659.3333 | 660.8191 | 2.2075 | 325622.4011
|  | 136.6667 | 117.5668 | 364.8034 | 18677.7778 | [5 empty] | 136.6667 | 136.645 | 0.00047125 | 2300.8011 | [5 empty] | 72 | 80.5253 | 72.6811 | 278.89
|  | 27.6667 | 34.8565 | 51.6937 | 765.4444 | [5 empty] | 27.6667 | 27.8309 | 0.026975 | 3725.0678 | [5 empty] | 70.6667 | 68.5376 | 4.5329 | 325.2011
|  | 19.6667 | 22.0481 | 5.6711 | 386.7778 | [5 empty] | 19.6667 | 19.3785 | 0.083042 | 4765.6011 | [5 empty] | 135.3333 | 128.0911 | 52.4505 | 2174.6678
|  | 46 | 46.9987 | 0.99739 | 2116 | [5 empty] | 46 | 46.0749 | 0.0056061 | 1823.29 | [5 empty] | 162 | 153.0711 | 79.7257 | 5372.89
|  | 84.6667 | 69.8025 | 220.9426 | 7168.4444 | [5 empty] | 84.6667 | 84.6496 | 0.00029257 | 16.2678 | [5 empty] | 41.3333 | 48.7624 | 55.1916 | 2243.6011
|  | 20.6667 | 24.9536 | 18.3781 | 427.1111 | [5 empty] | 20.6667 | 20.7885 | 0.014832 | 4628.5344 | [5 empty] | 24.6667 | 23.627 | 1.0808 | 4100.2678
|  | 17.6667 | 18.9815 | 1.7289 | 312.1111 | [5 empty] | 17.6667 | 17.4071 | 0.067401 | 5045.7344 | [5 empty] | 68.6667 | 67.4445 | 1.4936 | 401.3344
|  | 35.3333 | 34.5978 | 0.54094 | 1248.4444 | [5 empty] | 35.3333 | 35.3893 | 0.0031341 | 2848.0011 | [5 empty] | 136.6667 | 135.9848 | 0.46491 | 2300.8011
|  | 81.6667 | 65.1451 | 272.9605 | 6669.4444 | [5 empty] | 81.6667 | 81.6474 | 0.00036995 | 49.4678 | [5 empty] | 27.6667 | 33.3342 | 32.1209 | 3725.0678
|  | 20.3333 | 23.7548 | 11.7061 | 413.4444 | [5 empty] | 20.3333 | 20.4525 | 0.014194 | 4674.0011 | [5 empty] | 19.6667 | 19.6894 | 0.00051732 | 4765.6011
|  | 16.3333 | 16.9797 | 0.41779 | 266.7778 | [5 empty] | 16.3333 | 16.0928 | 0.05787 | 5236.9344 | [5 empty] | 46 | 46.1481 | 0.02193 | 1823.29
|  | 33.3333 | 31.5807 | 3.0717 | 1111.1111 | [5 empty] | 33.3333 | 33.385 | 0.0026725 | 3065.4678 | [5 empty] | 84.6667 | 83.3449 | 1.7471 | 16.2678
|  |  | SUM | 3919.4185 | 536052.2222 | [6 empty] | SUM | 1.782 | 378698.4222 | [5 empty] | 20.6667 | 25.2522 | 21.0269 | 4628.5344
|  |  |  | variation | 0.0073116 | [7 empty] | variation | 4.7057e-06 | [5 empty] | 17.6667 | 18.3804 | 0.50946 | 5045.7344
|  |  |  | R-square | 0.99269 | [7 empty] | R-square | 1 | [5 empty] | 35.3333 | 36.081 | 0.55907 | 2848.0011
[20 empty] | 81.6667 | 83.3436 | 2.8121 | 49.4678
[20 empty] | 20.3333 | 25.8248 | 30.1562 | 4674.0011
[20 empty] | 16.3333 | 17.5186 | 1.4049 | 5236.9344
[20 empty] | 33.3333 | 35.1851 | 3.4289 | 3065.4678
[21 empty] | SUM | 363.6166 | 378698.4222
[22 empty] | variation | 0.00096017
[22 empty] | R-square | 0.99904
𝜽_𝒎=𝝅/(𝒏^𝒂 𝑵_𝑸 ) (𝑳/𝑫) (𝑫/𝒅)^𝟑 (𝒉_(𝑳,𝒅)/𝒅)^𝒃
𝜃_𝑚=𝜋/(𝑛^𝑎 𝑁_𝑄 ) (𝐿/𝐷) (𝐷/𝑑)^3 [(ℎ_(𝐿,𝑑)/𝑑)^𝑏−𝑐(ℎ_(𝐿,𝑑)/𝑑)+𝑑]
𝜽_𝒎=𝑲(𝒉_(𝑳,𝒅)/𝒅)^𝒂 (𝑵_𝑹𝒆 )^𝒃 (𝑵_𝑷 )^𝒄 (𝑵_𝑸 )^𝒅
〖(𝑵〗_𝑷 𝑵_𝑹𝒆/𝑵_𝑸)
𝝆𝑵𝒅^𝟐/µ

## 5
Determination of minimum agitation speed at different impeller immersion ratios from STR model
Parameters | Symbol | Values | Units | Impeller coverage ratio |  | From Literature
Kinematic viscosity | ν | 6.915e-07 | m2 s-1 | Volume (L) | h,i/d
Diameter of particle | dp | 0.0002 | m | 1 | 0.04037
Gravity | g | 9.81 | m s-2 | 2 | 0.42156
Liquid density | ρL | 993.33 | kg m-3 | 3 | 0.79207 | Equation 1 |  |  | Equation 2
Solid density | ρS | 1024 | kg m-3 | 4 | 1.2085
Diameter of impeller | d | 0.05 | m
detrmination at different impeller immersion ratio in rps |  |  |  |  | Minimum speed in rpm (experimental)
h,i/d | χ (solid loading percentage) |  |  |  | h,i/d | χ (solid loading percentage)
| 0.25 | 0.5 | 0.75 | 1 |  | 0.25 | 0.5 | 0.75 | 1
4.037% | Not determined due to low impeller coverage |  |  |  | 4.037% | Not determined due to low impeller coverage
42.16% | 7.5 | 8.8333 | 11.1667 | 12.5 | 42.16% | 450 | 529.9998 | 670.002 | 750
79.21% | 7.75 | 9.3333 | 9.8333 | 9.6667 | 79.21% | 465 | 559.9998 | 589.998 | 580.002
120.9% | 3.6667 | 4 | 5.3333 | 7.6667 | 120.9% | 220.002 | 240 | 319.998 | 460.002
The value of S at each stage (Eq 2) |  | Number of runs at each impeller immersion (N) |  |  |  | 4
h,i/d | χ (solid loading percentage) |  |  |  | Mean at each impeller immersion | Std Deviation | Standard error | U (95 %) | % error
| 0.25 | 0.5 | 0.75 | 1
4.037% | Not determined due to low impeller coverage
42.16% | 27.3331 | 29.4183 | 35.2798 | 38.0425 | 32.5184 | 4.988 | 2.494 | 4.988 | 15.34%
79.21% | 28.2442 | 31.0835 | 31.0671 | 29.4196 | 29.9536 | 1.3813 | 0.69064 | 1.3813 | 4.611%
120.9% | 13.3629 | 13.3215 | 16.8499 | 23.3328 | 16.7168 | 4.7105 | 2.3552 | 4.7105 | 28.18%
Error Propagation
Predicted Value of from equation 1 in rps |  |  |  |  | Predicted Value of in rpm | [5 empty] | Error in predicted Value of in rps |  |  |  |  | Error in predicted Value of in rpm
h,i/d | χ (solid loading percentage) |  |  |  | h,i/d | χ (solid loading percentage) |  |  |  |  | h,i/d | χ (solid loading percentage) |  |  |  | h,i/d | χ (solid loading percentage)
| 0.25 | 0.5 | 0.75 | 1 |  | 0.25 | 0.5 | 0.75 | 1 |  |  | 0.25 | 0.5 | 0.75 | 1 |  | 0.25 | 0.5 | 0.75 | 1
4.037% | Not determined due to low impeller coverage |  |  |  | 0% | Not determined due to low impeller coverage |  |  |  |  | 3.733% | Not determined due to low impeller coverage |  |  |  | 4% | Not determined due to low impeller coverage
42.16% | 8.9228 | 9.7642 | 10.2927 | 10.6849 | 42% | 535.3695 | 585.8515 | 617.5604 | 641.0936 |  | 41.85% | 1.3687 | 1.4977 | 1.5788 | 1.6389 | 42% | 82.1198 | 89.8632 | 94.727 | 98.3368
79.21% | 8.2191 | 8.9941 | 9.4809 | 9.8421 | 79% | 493.1434 | 539.6437 | 568.8516 | 590.5287 |  | 78.9% | 0.37901 | 0.41475 | 0.4372 | 0.45386 | 79% | 22.7408 | 24.8851 | 26.232 | 27.2316
120.9% | 4.587 | 5.0195 | 5.2912 | 5.4928 | 121% | 275.2183 | 301.1697 | 317.4703 | 329.5681 |  | 120.5% | 1.2925 | 1.4144 | 1.491 | 1.5478 | 121% | 77.5515 | 84.8641 | 89.4573 | 92.8663
Error in % from observed and predictied values of | [10 empty] | Upper range of error in predicting value |  |  |  |  | Lower range of error in predicting value
h,i/d | χ (solid loading percentage) | [9 empty] | h,i/d | χ (solid loading percentage) |  |  |  | h,i/d | χ (solid loading percentage)
| 0.25 | 0.5 | 0.75 | 1 | [7 empty] | 0.25 | 0.5 | 0.75 | 1 |  | 0.25 | 0.5 | 0.75 | 1
4.037% | Not determined due to low impeller coverage | [9 empty] | 3.733% | Not determined due to low impeller coverage |  |  |  | 4% | Not determined due to low impeller coverage
42.16% | 18.97% | 10.54% | 7.827% | 14.52% | [6 empty] | 41.85% | 10.2915 | 11.2619 | 11.8715 | 12.3238 | 42% | 7.5542 | 8.2665 | 8.7139 | 9.0459
79.21% | 6.052% | 3.635% | 3.584% | 1.815% | [6 empty] | 78.9% | 8.5981 | 9.4088 | 9.9181 | 10.296 | 79% | 7.84 | 8.5793 | 9.0437 | 9.3883
120.9% | 25.1% | 25.49% | 0.7899% | 28.36% | [6 empty] | 120.5% | 5.8795 | 6.4339 | 6.7821 | 7.0406 | 121% | 3.2944 | 3.6051 | 3.8002 | 3.945
R-squared calculations
Overall R2 calculations for impeller speed in rpm | [5 empty] | For solid loading (χ) | [5 empty] | For solid loading (χ) for each impeller immersion
Impeller speed (Njs) | Modelled Njs | (y-y')^2 |  | (y-ymean)^2 |  | Solid loading (χ) | Impeller speed (Njs) | Modelled Njs |  |  |  | Impeller immersion | Impeller speed (Njs) | Modelled Njs
y | y' | [5 empty] | y | y' | (y-y')^2 | (y-ymean)^2 |  |  | y | y' | (y-y')^2 | (y-ymean)^2
450 | 535.3695 | 7287.9583 |  | 1314.0843 |  | 0.25 | 0 | 0 | 0 |  |  | 42.16% | 450 | 535.3695 | 7287.9583 | 22500.135
465 | 493.1434 | 792.0499 |  | 451.5753 |  | Mean | 7.5 | 8.9228 | 2.0244 | 1.4267 |  | Mean | 529.9998 | 585.8515 | 3119.4141 | 4900.091
220.002 | 275.2183 | 3048.8417 |  | 70888.1573 |  | 6.3056 | 7.75 | 8.2191 | 0.22001 | 2.0864 |  | 600.0005 | 670.002 | 617.5604 | 2750.1242 | 4900.217
529.9998 | 585.8515 | 3119.4141 |  | 1914.0188 |  |  | 3.6667 | 4.587 | 0.8469 | 6.9636 |  |  | 750 | 641.0936 | 11860.599 | 22499.865
559.9998 | 539.6437 | 414.371 |  | 5438.9888 |  |  |  |  | 3.0913 | 10.4767 |  |  |  |  | 25018.0956 | 54800.308
240 | 301.1697 | 3741.7287 |  | 60639.2103 |  |  |  |  | Variations | 0.29507 |  |  |  |  | Variations | 0.45653
670.002 | 617.5604 | 2750.1242 |  | 33764.6873 |  |  |  |  | R-Squared | 0.70493 |  |  |  |  | R-Squared | 0.54347
589.998 | 568.8516 | 447.1709 |  | 10763.5853
319.998 | 317.4703 | 6.3893 |  | 27639.8273 |  | Solid loading (χ) | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2 |  | Impeller immersion | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2
750 | 641.0936 | 11860.599 |  | 69563.9043 |  |  | y | y' |  |  |  |  | y | y'
580.002 | 590.5287 | 110.8113 |  | 8789.3813 |  | 0.5 | 0 | 0 | 0 |  |  | 79.21% | 465 | 493.1434 | 792.0499 | 7014.0541
460.002 | 329.5681 | 17013.0089 |  | 688.9733 |  | Mean | 8.8333 | 9.7642 | 0.8665 | 2.0864 |  | Mean | 559.9998 | 539.6437 | 414.371 | 126.5591
486.2503 |  | 50592.4674 |  | 291856.393 |  | 7.3889 | 9.3333 | 8.9941 | 0.1151 | 3.7809 |  | 548.75 | 589.998 | 568.8516 | 447.1709 | 1701.4016
| Variations |  | 0.17335 |  |  |  | 4 | 5.0195 | 1.0394 | 11.4846 |  |  | 580.002 | 590.5287 | 110.8113 | 976.6906
| R-Squared |  | 0.82665 | [5 empty] | 2.021 | 17.3518 |  |  |  |  | 1764.4032 | 9818.7055
[9 empty] | Variations | 0.11647 |  |  |  |  | Variations | 0.1797
Overall R2 calculations for impeller speed in rps | [8 empty] | R-Squared | 0.88353 |  |  |  |  | R-Squared | 0.8203
Impeller speed (Njs) | Modelled Njs | (y-y')^2 |  | (y-ymean)^2
y | y' |  |  |  |  | Solid loading (χ) | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2 |  | Impeller immersion | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2
7.5 | 8.9228 | 2.0244 |  | 0.36502 |  |  | y | y' |  |  |  |  | y | y'
7.75 | 8.2191 | 0.22001 |  | 0.12544 |  | 0.75 | 0 | 0 | 0 |  |  | 1.2085 | 220.002 | 275.2183 | 3048.8417 | 8099.73
3.6667 | 4.587 | 0.8469 |  | 19.6912 |  | Mean | 11.1667 | 10.2927 | 0.76392 | 5.707 |  | Mean | 240 | 301.1697 | 3741.7287 | 4900.07
8.8333 | 9.7642 | 0.8665 |  | 0.53167 |  | 8.7778 | 9.8333 | 9.4809 | 0.12421 | 1.1142 |  | 310.0005 | 319.998 | 317.4703 | 6.3893 | 99.95
9.3333 | 8.9941 | 0.1151 |  | 1.5108 |  |  | 5.3333 | 5.2912 | 0.0017748 | 11.8644 |  |  | 460.002 | 329.5681 | 17013.0089 | 22500.45
4 | 5.0195 | 1.0394 |  | 16.8442 |  |  |  |  | 0.88991 | 18.6855 |  |  |  |  | 23809.9686 | 35600.2
11.1667 | 10.2927 | 0.76392 |  | 9.3791 |  |  |  |  | Variations | 0.047626 |  |  |  |  | Variations | 0.66882
9.8333 | 9.4809 | 0.12421 |  | 2.9899 |  |  |  |  | R-Squared | 0.95237 |  |  |  |  | R-Squared | 0.33118
5.3333 | 5.2912 | 0.0017748 |  | 7.6777
12.5 | 10.6849 | 3.2946 |  | 19.3233 |  | Solid loading (χ) | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2
9.6667 | 9.8421 | 0.030781 |  | 2.4415 |  |  | y | y'
7.6667 | 5.4928 | 4.7258 |  | 0.19138 |  | 1 | 0 | 0 | 0
8.1042 |  | 14.0535 |  | 81.0712 |  | Mean | 12.5 | 10.6849 | 3.2946 | 6.5308
| Variations |  | 0.17335 |  |  | 9.9445 | 9.6667 | 9.8421 | 0.030781 | 0.077154
| R-Squared |  | 0.82665 |  |  |  | 7.6667 | 5.4928 | 4.7258 | 5.1882
[9 empty] | 8.0512 | 11.7961
[9 empty] | Variations | 0.68253
[9 empty] | R-Squared | 0.31747
𝑵_𝒋𝒔
𝑺=(𝑵_𝒋𝒔∙𝒅^(𝟎.𝟖𝟓))/(𝝂^(𝟎.𝟏)∙𝒅_𝒑^(𝟎.𝟐)∙((𝒈∙∆𝝆)/𝝆_𝑳 )^(𝟎.𝟒𝟓) 𝝌^(𝟎.𝟏𝟑) )
𝑵_𝒋𝒔=(〖𝑺.𝝂〗^(𝟎.𝟏)∙𝒅_𝒑^(𝟎.𝟐)∙((𝒈∙∆𝝆)/𝝆_𝑳 )^(𝟎.𝟒𝟓) 𝝌^(𝟎.𝟏𝟑))/𝒅^(𝟎.𝟖𝟓)
𝑵_𝒋𝒔
𝑵_𝒋𝒔
𝑵_𝒋𝒔
𝑵_𝒋𝒔
𝑵_𝒋𝒔
𝑵_𝒋𝒔
𝑵_𝒋𝒔

## 6
Determination of minimum agitation speed at different impeller immersion ratios from HTB model
For HTB
Parameters | Symbol | Values | Units | Impeller coverage ratio |  | From Literature
Kinematic viscosity | ν | 6.915e-07 | m2 s-1 | Volume (L) | h,i/d | [9 empty] | Equation 3
Diameter of particle | dp | 0.0002 | m | 1 | 0.04037
Gravity | g | 9.81 | m s-2 | 2 | 0.42156
Liquid density | ρL | 993.33 | kg m-3 | 3 | 0.79207 | Equation 1 |  |  | Equation 2
Solid density | ρS | 1024 | kg m-3 | 4 | 1.2085
diameter of impeller | d | 0.05 | m
After solving exponents of the equation
detrmination at different impeller immersion ratio in rps | [10 empty] | K | 0.96507
h,i/d | χ (solid loading percentage) | [9 empty] | a | 0.047963
| 0.25 | 0.5 | 0.75 | 1 | [6 empty] | b | 0.046251
4.037% | Not determined due to low impeller coverage | [9 empty] | c | 0.091613
42.16% | 7.5 | 8.8333 | 11.1667 | 12.5 | [6 empty] | d | 0.31998
79.21% | 7.75 | 9.3333 | 9.8333 | 9.6667 | [6 empty] | e | 0.098616
120.9% | 3.6667 | 4 | 5.3333 | 7.6667
S
The value of S at each stage (Eq 2) |  | Number of runs at each impeller immersion (N) |  |  |  | 4
h,i/d | χ (solid loading percentage) |  |  |  | Mean at each impeller immersion | STD Deviation | Standard error | U (95 %) | % error
| 0.25 | 0.5 | 0.75 | 1
4.037% | Not determined due to low impeller coverage
42.16% | 27.3331 | 29.4183 | 35.2798 | 38.0425 | 32.5184 | 4.988 | 2.494 | 4.988 | 15.34%
79.21% | 28.2442 | 31.0835 | 31.0671 | 29.4196 | 29.9536 | 1.3813 | 0.69064 | 1.3813 | 4.611%
120.9% | 13.3629 | 13.3215 | 16.8499 | 23.3328 | 16.7168 | 4.7105 | 2.3552 | 4.7105 | 28.18%
Predicted Value of from equation 1 | [5 empty] | Predicted Value of from equation 3 for HTB | [5 empty] | Sum of square errors
h,i/d | χ (solid loading percentage) |  |  |  |  | h,i/d | χ (solid loading percentage) |  |  |  |  | h,i/d | χ (solid loading percentage)
| 0.25 | 0.5 | 0.75 | 1 |  |  | 0.25 | 0.5 | 0.75 | 1 |  |  | 0.25 | 0.5 | 0.75 | 1
4.037% | Not determined due to low impeller coverage |  |  |  |  | 4.037% | Not determined due to low impeller coverage |  |  |  |  | 4.037% | Not determined due to low impeller coverage
42.16% | 8.9228 | 9.7642 | 10.2927 | 10.6849 |  | 42.16% | 7.6015 | 9.4891 | 10.8036 | 11.8453 |  | 42.16% | 0.010308 | 0.43001 | 0.13182 | 0.42859
79.21% | 8.2191 | 8.9941 | 9.4809 | 9.8421 |  | 79.21% | 7.0221 | 8.7658 | 9.9801 | 10.9424 |  | 79.21% | 0.52985 | 0.32213 | 0.021554 | 1.6274
120.9% | 4.587 | 5.0195 | 5.2912 | 5.4928 |  | 120.9% | 3.9996 | 4.9928 | 5.6844 | 6.2325 |  | 120.9% | 0.11083 | 0.98559 | 0.12329 | 2.0568
Error Propagation
Error in % from observed and predictied values of | [5 empty] | Error in % from observed and predictied values of | [5 empty] | Error in predicted Value of in rps |  |  |  |  | Error in predicted Value of in RPM
h,i/d | χ (solid loading percentage) |  |  |  |  | h,i/d | χ (solid loading percentage) |  |  |  |  | h,i/d | χ (solid loading percentage) |  |  |  | h,i/d | χ (solid loading percentage)
| 0.25 | 0.5 | 0.75 | 1 |  |  | 0.25 | 0.5 | 0.75 | 1 |  |  | 0.25 | 0.5 | 0.75 | 1 |  | 0.25 | 0.5 | 0.75 | 1
4.037% | Not determined due to low impeller coverage |  |  |  |  | 4.037% | Not determined due to low impeller coverage |  |  |  |  | 3.733% | Not determined due to low impeller coverage |  |  |  | 4% | Not determined due to low impeller coverage
42.16% | 18.97% | 10.54% | 7.827% | 14.52% |  | 42.16% | 1.354% | 7.424% | 3.251% | 5.237% |  | 41.85% | 1.166 | 1.4555 | 1.6572 | 1.8169 | 42% | 69.9595 | 87.3313 | 99.4295 | 109.0167
79.21% | 6.052% | 3.635% | 3.584% | 1.815% |  | 79.21% | 9.392% | 6.081% | 1.493% | 13.2% |  | 78.9% | 0.32382 | 0.40422 | 0.46022 | 0.5046 | 79% | 19.429 | 24.2534 | 27.6133 | 30.2758
120.9% | 25.1% | 25.49% | 0.7899% | 28.36% |  | 120.9% | 9.079% | 24.82% | 6.584% | 18.71% |  | 120.5% | 1.127 | 1.4069 | 1.6018 | 1.7562 | 121% | 67.6211 | 84.4123 | 96.1061 | 105.3729
Minimum speed in RPM (experimental) | [5 empty] | Minimum speed in RPM (model) | [5 empty] | Upper range of error in predicting value |  |  |  |  | Lower range of error in predicting value
h,i/d | χ (solid loading percentage) |  |  |  |  | h,i/d | χ (solid loading percentage) |  |  |  |  | h,i/d | χ (solid loading percentage) |  |  |  | h,i/d | χ (solid loading percentage)
| 0.25 | 0.5 | 0.75 | 1 |  |  | 0.25 | 0.5 | 0.75 | 1 |  |  | 0.25 | 0.5 | 0.75 | 1 |  | 0.25 | 0.5 | 0.75 | 1
4.037% | Not determined due to low impeller coverage |  |  |  |  | 4.037% | Not determined due to low impeller coverage |  |  |  |  | 3.733% | Not determined due to low impeller coverage |  |  |  | 4% | Not determined due to low impeller coverage
42.16% | 450 | 529.9998 | 670.002 | 750 |  | 42.16% | 456.0916 | 569.3447 | 648.2176 | 710.7201 |  | 41.85% | 8.7675 | 10.9446 | 12.4608 | 13.6623 | 42% | 6.4355 | 8.0336 | 9.1465 | 10.0284
79.21% | 465 | 559.9998 | 589.998 | 580.002 |  | 79.21% | 421.3256 | 525.946 | 598.8067 | 656.5449 |  | 78.9% | 7.3459 | 9.17 | 10.4403 | 11.447 | 79% | 6.6983 | 8.3615 | 9.5199 | 10.4378
120.9% | 220.002 | 240 | 319.998 | 460.002 |  | 120.9% | 239.9769 | 299.5661 | 341.0658 | 373.952 |  | 120.5% | 5.1266 | 6.3996 | 7.2862 | 7.9887 | 121% | 2.8726 | 3.5859 | 4.0827 | 4.4763
R-squared calculations
For Impeller speed in RPM | [5 empty] | For solid loading (χ) | [5 empty] | For solid loading (χ) for each impeller immersion
Impeller speed (Njs) | Modelled Njs | (y-y')^2 |  | (y-ymean)^2 |  | Solid loading (χ) | Impeller speed (Njs) | Modelled Njs |  |  |  | Impeller immersion | Impeller speed (Njs) | Modelled Njs
y | y' | [5 empty] | y | y' | (y-y')^2 | (y-ymean)^2 |  |  | y | y' | (y-y')^2 | (y-ymean)^2
450 | 456.0916 | 37.1072 |  | 1314.0843 |  | 0.25 | 0 | 0 | 0 | 22.3651 |  | 42.16% | 7.5 | 7.6015 | 0.010308 | 6.25
465 | 421.3256 | 1907.4508 |  | 451.5753 |  | Mean | 7.5 | 7.6015 | 0.010308 | 7.6775 |  | Mean | 8.8333 | 9.4891 | 0.43001 | 1.3611
220.002 | 239.9769 | 398.9949 |  | 70888.1573 |  | 4.7292 | 7.75 | 7.0221 | 0.52985 | 9.1254 |  | 10 | 11.1667 | 10.8036 | 0.13182 | 1.3612
529.9998 | 569.3447 | 1548.0247 |  | 1914.0188 |  |  | 3.6667 | 3.9996 | 0.11083 | 1.1289 |  |  | 12.5 | 11.8453 | 0.42859 | 6.25
559.9998 | 525.946 | 1159.662 |  | 5438.9888 |  |  |  |  | 0.65099 | 17.9317 |  |  |  |  | 1.0007 | 15.2223
240 | 299.5661 | 3548.1175 |  | 60639.2103 |  |  |  |  | Variations | 0.036304 |  |  |  |  | Variations | 0.065741
670.002 | 648.2176 | 474.5601 |  | 33764.6873 |  |  |  |  | R-Squared | 0.9637 |  |  |  |  | R-Squared | 0.93426
589.998 | 598.8067 | 77.5932 |  | 10763.5853
319.998 | 341.0658 | 443.8509 |  | 27639.8273 |  | Solid loading (χ) | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2 |  | Impeller immersion | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2
750 | 710.7201 | 1542.9118 |  | 69563.9043 |  |  | y | y' |  |  |  |  | y | y'
580.002 | 656.5449 | 5858.8125 |  | 8789.3813 |  | 0.5 | 0 | 0 | 0 | 30.7101 |  | 79.21% | 7.75 | 7.0221 | 0.52985 | 1.9483
460.002 | 373.952 | 7404.5961 |  | 688.9733 |  | Mean | 8.8333 | 9.4891 | 0.43001 | 10.8351 |  | Mean | 9.3333 | 8.7658 | 0.32213 | 0.035155
486.2503 |  | 24401.6818 |  | 291856.393 |  | 5.5417 | 9.3333 | 8.7658 | 0.32213 | 14.3767 |  | 9.1458 | 9.8333 | 9.9801 | 0.021554 | 0.47261
| Variations |  | 0.083609 |  |  |  | 4 | 4.9928 | 0.98559 | 2.3767 |  |  | 9.6667 | 10.9424 | 1.6274 | 0.2713
| R-Squared |  | 0.91639 | [5 empty] | 1.7377 | 27.5885 |  |  |  |  | 2.501 | 2.7274
[9 empty] | Variations | 0.062987 |  |  |  |  | Variations | 0.91698
[9 empty] | R-Squared | 0.93701 |  |  |  |  | R-Squared | 0.083024
[6 empty] | Solid loading (χ) | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2 |  | Impeller immersion | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2
[7 empty] | y | y' |  |  |  |  | y | y'
[6 empty] | 0.75 | 0 | 0 | 0 | 43.3402 |  | 1.2085 | 3.6667 | 3.9996 | 0.11083 | 2.2499
[6 empty] | Mean | 11.1667 | 10.8036 | 0.13182 | 21.0073 |  | Mean | 4 | 4.9928 | 0.98559 | 1.3611
[6 empty] | 6.5833 | 9.8333 | 9.9801 | 0.021554 | 10.5623 |  | 5.1667 | 5.3333 | 5.6844 | 0.12329 | 0.027764
[7 empty] | 5.3333 | 5.6844 | 0.12329 | 1.5626 |  |  | 7.6667 | 6.2325 | 2.0568 | 6.2501
[9 empty] | 0.27667 | 33.1322 |  |  |  |  | 3.2765 | 9.8889
[9 empty] | Variations | 0.0083504 |  |  |  |  | Variations | 0.33133
[9 empty] | R-Squared | 0.99165 |  |  |  |  | R-Squared | 0.66867
[6 empty] | Solid loading (χ) | Impeller speed (Njs) | Modelled Njs | (y-y')^2 | (y-ymean)^2
[7 empty] | y | y'
[6 empty] | 1 | 0 | 0 | 0 | 43.3402
[6 empty] | Mean | 12.5 | 11.8453 | 0.42859 | 25.4182
[6 empty] | 7.4583 | 9.6667 | 10.9424 | 1.6274 | 4.8768
[7 empty] | 7.6667 | 6.2325 | 2.0568 | 0.04341
[9 empty] | 4.1129 | 30.3385
[9 empty] | Variations | 0.13557
[9 empty] | R-Squared | 0.86443
𝑁_𝑗𝑠
𝑺=(𝑵_𝒋𝒔∙𝒅^(𝟎.𝟖𝟓))/(𝝂^(𝟎.𝟏)∙𝒅_𝒑^(𝟎.𝟐)∙((𝒈∙∆𝝆)/𝝆_𝑳 )^(𝟎.𝟒𝟓) 𝝌^(𝟎.𝟏𝟑) )
𝑵_𝒋𝒔=(〖𝑺.𝝂〗^(𝟎.𝟏)∙𝒅_𝒑^(𝟎.𝟐)∙((𝒈∙∆𝝆)/𝝆_𝑳 )^(𝟎.𝟒𝟓) 𝝌^(𝟎.𝟏𝟑))/𝒅^(𝟎.𝟖𝟓)
𝑵_𝒋𝒔
𝑁_𝑗𝑠
𝑵_𝒋𝒔=(𝑺^(𝟎.𝟗𝟕) 〖.𝝂〗^(𝟎.𝟎𝟒𝟖)∙𝒅_𝒑^(𝟎.𝟎𝟒𝟔)∙((𝒈∙∆𝝆)/𝝆_𝑳 )^(𝟎.𝟎𝟗𝟐) 𝝌^(𝟎.𝟑𝟐))/𝒅^(𝟎.𝟏𝟎)
𝑵_𝒋𝒔=(〖𝟑𝟑〗^(𝟎.𝟗𝟕) 〖.𝟔.𝟗𝟏𝟓𝒆−𝟎𝟕〗^(𝟎.𝟎𝟒𝟖)∙〖𝟎.𝟎𝟎𝟎𝟐〗_^(𝟎.𝟎𝟒𝟔)∙((𝟗.𝟖𝟏∙𝟑𝟎.𝟔𝟕)/(𝟗𝟗𝟑.𝟑𝟑))^(𝟎.𝟎𝟗𝟐) 〖𝟎.𝟐𝟓〗^(𝟎.𝟑𝟐))/〖𝟎.𝟎𝟓〗^(𝟎.𝟏𝟎)
𝑵_𝒋𝒔
𝑁_𝑗𝑠
𝑵_𝒋𝒔
𝑵_𝒋𝒔
𝑵_𝒋𝒔
𝑵_𝒋𝒔

## 7
Modelling of oxygen mass transfer (Buckingham Pi method) |  |  | //NL-regresssion for kLa
clear;clc;format(12);
Oxygen mass transfer is a function of |  |  | //Dimensions, physical data
Drx = 0.109; // reactor internal diameter in m
Symbol | Parameter | Fundamental dimensions | L = 0.500; //reactor length in m
| Liquid height covering impeller | L | d = 0.050; //impeller diameter in m
| Impeller diameter | L | rho = 993.33; //density of water in g/cm3 at 37 deg C, 1 atm, CRC Handbook 2015
| Impeller rotational speed | T-1 | mu = 0.0006915; //viscosity of water in Pa.s at 37 deg C, 1 atm, CRC Handbook 2015
| Power dissipated | ML2T-3 | Nq=0.19; //Tsui & Hu 2008
| Fluid density | ML-3 | NpRe=114; //Tsui & Hu 2008
| Fluid viscosity | ML-1T-1 | k1=NpRe/Nq; //<-- this value is a constant for higher Re numbers**
| Superficial gas velocity | L2T-1 | di = 0.0032; //internal diameter of air tube in m
| Volumetric bulk liquid pumping rate | L3T-1 | d_he = 0.022; //<-- outer diameter of heating element
Additional parameters: |  |  | L_he = 0.292; //<-- length of heating element
| Vessel internal diameter | L
| Vessel length | L | //The function and loop below calculate the impeller coverage
| Impeller pitch | L | function VL=area(hL, D)
δ | Diffusivity of oxygen in liquid | L2T-1 | r = D/2;
if hL < r then
It is assumed that the parameters δ, D,L and S1 are fixed, and so are included in a pre-exponential constant coefficient K1 |  |  | t = 2.*acos((r-hL)./r);
Dimensional matrix w.r.t parameters |  |  | AL = ((r.^2).*(t-sin(t)))/2;
else
t = 2.*acos((r-(D-hL))./r);
AL = (%pi.*(r.^2))-(((r.^2).*(t-sin(t)))/2);
end
VL = AL.*L;
endfunction
Applying Buckingham theorem for number if independent dimensionless groups:
Vhe=(%pi.*(d_he.^2)/4)*L_he; //<-- heating element volume displaced
hL=linspace(0,0.109,500); //<-- all possible liquid heights
therefore 5 core variables were selected |  |  | VL=area(hL,Drx); //<-- all posssible liquid volumes\
//The loop below extracts liquid heights at integer volumes 1-4 L
Ve=[0.001 0.002 0.003 0.004]; Vf=Ve+Vhe; Vx=[]; H=[]; tol=5e-6;
for i=1:4
for n=1:length(VL)
if abs(VL(n)-Vf(i))<tol then
Vx=[Vx VL(n)];
H=[H hL(n)];
end
end
M | g + 1=0 |  | end
| g = -1 |  | //The vector H now stores the liquid height in m, at Volumes of 1- 4 L
L | e - 3g + 2=0 |  | hLd=H-((Drx-d)/2); // <-- height of water covering the impeller
| e= -5 |  | hLd_D=hLd./d; // <-- Impeller immersion ratio
T
| f = -3 |  | //Extract kLa data
//Data = readxls('kLa_analysis.xls'); D=Data(2);//<--kLa data
D=csvRead('C:\Users\SHRRAJ002\Desktop\Wesley_Andrew_data\Conductivity experiment\Data files\kLa_analysis.csv');
hLdD=D(2:41,2); Vs=D(2:41,3); N=D(2:41,4); kLa=D(2:41,5);
//Generate dimensionless groups
NRe=((rho.*N).*(d.^2))./mu;
Nq=D(2:41,6); //<-- circulation flow number, estimated from mixing time
Np=k1./(NRe./Nq); //<-- predicts power number
NVs=Vs./(N.*d);
M=[hLdD NRe Nq Np NVs];
//Call leastsq function
//Model form based on dimensional analysis
function k=kla(M, I)
k=I(1).*((M(:,1)).^(I(2))).*(M(:,2).^(I(3))).*(M(:,3).^(I(4))).*(M(:,4).^(I(5))).*(M(:,5).^(I(6)));
endfunction
//Define functions to minimize
function y=func(I, M, kLa)
y=kla(M, I)- kLa;
endfunction
I0=[0.5 0.1 0.1 0.1 0.1 0.1]'; M2=M(1:20,1:$); kLa2=kLa(1:20); M3=M(21:40,1:$); kLa3=kLa(21:40);
[K1 K2 K3] = leastsq(list(func,M,kLa),I0);
[K2_1 K2_2 K2_3] = leastsq(list(func,M2,kLa2),I0);
[K3_1 K3_2 K3_3] = leastsq(list(func,M3,kLa3),I0);
//Calculate the single variables exponents e.g. d, n, Q etc.
dd=(-1*K2(2))+(-5*K2(5))+(2*K2(3))+(-1*K2(6))+(-3*K2(4));
nn=(-3*K2(5))+(1*K2(3))+(-1*K2(6))+(-1*K2(4));
rr=(-1*K2(5))+(1*K2(3));
KK1=K2(1)*(mu^-K2(3))*(rho^rr);
Tm2=[72 41.333333 27.666667 20.666667 20.333333]'; //<-- experimentally recorded mixing times at 2L
Hypothetical model based on Buckingham Pi |  |  | Tm3=[70.666667 24.666667 19.666667 17.666667 16.333333]'; //<-- experimentally recorded mixing times at 3L
Tm=[Tm2;Tm2;Tm2;Tm2;Tm3;Tm3;Tm3;Tm3];
//Models to use for predicting kLa
//Single model 2L
function kLa=model2L(n, Vs);
NRe=((rho.*n).*(d.^2))./mu;
The values of the exponents were determined by Scilab least square regression. The Scilab code has given separately with the excel file with the relevent data |  |  | //Tm=109.6.*(n.^(-0.834)); //<-- model for mixing time
Nq=(%pi./(n.*Tm)).*(L./Drx).*((Drx./d).^3);
Np=600.*(Nq./NRe);
NVs=Vs./(n.*d);
kLa = K2_2(1).*((hLd_D(2)).^K2_2(2)).*(NRe.^K2_2(3)).*(Nq.^K2_2(4)).*(Np.^K2_2(5)).*(NVs.^K2_2(6));
To simplify the relation, exponent d, n and k was calculated as follows |  |  | endfunction
Exponents of d' |  |  | //Different models 2L
function kLa=model2L_comb(n, Vs);
NRe=((rho.*n).*(d.^2))./mu;
//Tm=109.6.*(n.^(-0.834)); //<-- model for mixing time
Exponents of n |  |  | Nq=(%pi./(n.*Tm)).*(L./Drx).*((Drx./d).^3);
Np=600.*(Nq./NRe);
NVs=Vs./(n.*d);
kLa = K2(1).*((hLd_D(2)).^K2(2)).*(NRe.^K2(3)).*(Nq.^K2(4)).*(Np.^K2(5)).*(NVs.^K2(6));
Pre-exponential constant K1 |  |  | endfunction
//Single model 3L
function kla=model3L(n, Vs);
NRe=((rho.*n).*(d.^2))./mu;
//Tm=93.77.*(n.^(-0.899)); //<-- model for mixing time
Nq=(%pi./(n.*Tm)).*(L./Drx).*((Drx./d).^3);
Np=600.*(Nq./NRe);
NVs=Vs./(n.*d);
kla = K3_2(1).*((hLd_D(2)).^K3_2(2)).*(NRe.^K3_2(3)).*(Nq.^K3_2(4)).*(Np.^K3_2(5)).*(NVs.^K3_2(6));
endfunction
//Different models 3L
function kla=model3L_comb(n, Vs);
NRe=((rho.*n).*(d.^2))./mu;
//Tm=93.77.*(n.^(-0.899)); //<-- model for mixing time
Nq=(%pi./(n.*Tm)).*(L./Drx).*((Drx./d).^3);
Np=600.*(Nq./NRe);
NVs=Vs./(n.*d);
kla = K2(1).*((hLd_D(2)).^K2(2)).*(NRe.^K2(3)).*(Nq.^K2(4)).*(Np.^K2(5)).*(NVs.^K2(6));
endfunction
//Calculate coefficient of determination
//single model
y_bar=(sum(kLa)/length(kLa)); SS_total=0; //<-- mean kLa
for c=1:length(kLa)
SS_total=SS_total+((kLa(c)-y_bar).^2)
end
R2=1-(K1/SS_total);
//Different models
y_bar2=(sum(kLa2)/length(kLa2)); SS_2total=0; //<-- mean kLa
for c=1:length(kLa2)
SS_2total=SS_2total+((kLa2(c)-y_bar2).^2)
end
R2_2=1-(K2_1/SS_2total);
y_bar3=(sum(kLa3)/length(kLa3)); SS_3total=0; //<-- mean kLa
for c=1:length(kLa3)
SS_3total=SS_3total+((kLa3(c)-y_bar3).^2)
end
R2_3=1-(K3_1/SS_3total);
//plot 3d surfaces for kLa and experimental points
//For 2L:
scf(0);clf;
Nm=N(1:5); Vm=[Vs(1) Vs(6) Vs(11) Vs(16)];
//plot experimetnal results
kLa2m=[kLa2(1:5) kLa2(6:10) kLa2(11:15) kLa2(16:20)];
plot3d(Nm,Vm,kLa2m); //plot x=N , y=Vs , z= kla
title('2L kLa model vs. experimental','fontsize', 4);
xlabel('n [1/s]','fontsize', 3);
ylabel('Vs [m/s]','fontsize', 3);
zlabel('kLa [1/h]', 'fontsize', 3);
xgrid(0);
a=gca(); b=a.children(1);
b.surface_mode="off"; b.mark_mode="on"; b.mark_size=1.0; b.mark_style=0;
//plot model predictions
//kmodel2=model2L(N, Vs); //<-- enable to use separate models
kmodel2=model2L_comb(N, Vs); //<-- enable to use single model
Km2=[kmodel2(1:5) kmodel2(6:10) kmodel2(11:15) kmodel2(16:20)];
plot3d(Nm, Vm, Km2);
f=gcf();
f.color_map = graycolormap(128);
e=gca(); g=e.children(1);
g.color_flag=1;
//plot lines linking data points to surface
scf(0);
for m=1:length(Nm)
for p=1:length(Vm)
param3d([Nm(m) Nm(m)], [Vm(p) Vm(p)], [Km2(m,p) kLa2m(m,p)]);
end
end
//For 3L:
scf(1);clf;
//plot experimetnal results
kLa3m=[kLa3(1:5) kLa3(6:10) kLa3(11:15) kLa3(16:20)];
plot3d(Nm,Vm,kLa3m); //plot x=N , y=Vs , z= kla
title('3L kLa model vs. experimental','fontsize', 4);
xlabel('n [1/s]','fontsize', 3);
ylabel('Vs [m/s]','fontsize', 3);
zlabel('kLa [1/h]', 'fontsize', 3);
xgrid(0);
a=gca(); b=a.children(1);
b.surface_mode="off"; b.mark_mode="on"; b.mark_size=1.0; b.mark_style=0;
//plot model predictions
//kmodel3=model3L(N, Vs); //<-- enable to use separate models
kmodel3=model3L_comb(N, Vs); //<-- enable to use single model
Km3=[kmodel3(21:25) kmodel2(26:30) kmodel2(31:35) kmodel2(36:40)];
plot3d(Nm, Vm, Km3);
f=gcf();
f.color_map = graycolormap(128);
e=gca(); g=e.children(1);
g.color_flag=1;
scf(1);
for m=1:length(Nm)
for p=1:length(Vm)
param3d([Nm(m) Nm(m)], [Vm(p) Vm(p)], [Km3(m,p) kLa3m(m,p)]);
end
end
//Report results of least squares regression
//Single model
disp('Single model for kLa at all volumes is');
format(5);
disp('kLa = '+string(K2(1))+'.(hLd/d^'+string(K2(2))+').(NRe^'+string(K2(3))+').(Nq^'+string(K2(4))+').(Np^'+string(K2(5))+').(NVs^'+string(K2(6))+')');
disp('R^2 coefficient of determination for combined model is '+string(R2));
disp('');
//Different models
disp('Model for kLa at 2L volume is');
format(5);
disp('kLa = '+string(K2_2(1))+'.(hLd/d^'+string(K2_2(2))+').(NRe^'+string(K2_2(3))+').(Nq^'+string(K2_2(4))+').(Np^'+string(K2_2(5))+').(NVs^'+string(K2_2(6))+')');
disp('R^2 coefficient of determination for 2L model is '+string(R2_2));
disp('');
disp('Model for kLa at 3L volume is');
format(5);
disp('kLa = '+string(K3_2(1))+'.(hLd/d^'+string(K3_2(2))+').(NRe^'+string(K3_2(3))+').(Nq^'+string(K3_2(4))+').(Np^'+string(K3_2(5))+').(NVs^'+string(K3_2(6))+')');
disp('R^2 coefficient of determination for 3L model is '+string(R2_3));
format(12);
𝒌_𝑳 𝒂=𝒇(𝒉_(𝑳,𝒅), 𝒅, 𝒏,𝑷,𝝆,𝝁,𝑽𝒔,𝑸,𝑫,𝑳,𝑺,δ)
■(&𝒉_(𝑳,𝒅)&𝒅&𝒏&𝑷&𝝆&𝝁&𝑽_𝑺&𝑸@𝐌&0&0&0&1&1&1&0&0@𝐋&1&1&0&2&−3&−1&2&3@𝐭&0&0&−1&−3&0&−1&−1&−1)
𝑖=𝑛−𝑟=8−3
5
𝑖=5
𝝅_𝟏, 𝝅_𝟐, 𝝅_𝟑, 𝝅_𝟒, 𝝅_𝟓
𝒉_(𝑳,𝒅),𝑷,𝝁,𝑽_𝑺,𝑸
𝝅_𝟏=𝒅^𝒂 𝒏^𝒃 𝝆^𝒄 𝒉_(𝑳,𝒅)=𝑴^𝟎 𝑳^𝟎 𝒕^𝟎 " "
𝑴 ": " 𝒄=𝟎
"L : " 𝒂−𝟑𝒄+𝟏=𝟎
𝒂=−𝟏
𝝅_𝟏=𝒉_(𝑳,𝒅)/𝒅=𝒊𝒎𝒑𝒆𝒍𝒍𝒆𝒓 𝒊𝒎𝒎𝒆𝒓𝒔𝒊𝒐𝒏 𝒓𝒂𝒕𝒊𝒐
𝝅_𝟐=𝒅^𝒆 𝒏^𝒇 𝝆^𝒈 𝑷=𝑴^𝟎 𝑳^𝟎 𝒕^𝟎 " "
-𝒇-𝟑 = 𝟎
𝝅_𝟐=𝑷/(𝝆𝒏^𝟑 𝒅^𝟓 )=𝑵_𝑷
𝝅_𝟑=𝒅^𝒉 𝒏^𝒊 𝝆^𝒋 𝝁=𝑴^𝟎 𝑳^𝟎 𝒕^𝟎 " "
𝑴 " : " 𝒋+𝟏=𝟎
" " 𝒋=−𝟏
𝑳 ": " 𝒉−𝟑𝒋−𝟏=𝟎
" " 𝒉=−𝟐
"T : "−𝒊−𝟏=𝟎
" " 𝒊=−𝟏
𝝅_𝟑=𝝁/(𝝆𝒏𝒅^𝟐 )=𝟏/𝑵_𝑹𝒆
𝝅_𝟒=𝒅^𝒌 𝒏^𝒍 𝝆^𝒎 𝑽_𝑺=𝑴^𝟎 𝑳^𝟎 𝒕^𝟎 " "
𝑴 ": " 𝒎=𝟎
𝑳 ": " 𝒌−𝟑𝒎+𝟏=𝟎
" " 𝒌=−𝟏
"T : "−𝒍−𝟏=𝟎
" " 𝒍=−𝟏
𝝅_𝟒=𝑽_𝑺/𝒏𝒅=𝑵_𝑽𝑺=𝒎𝒐𝒅𝒊𝒇𝒊𝒆𝒅 𝒂𝒊𝒓 𝒇𝒍𝒐𝒘 𝒏𝒖𝒎𝒃𝒆𝒓
𝝅_𝟓=𝒅^𝒒 𝒏^𝒓 𝝆^𝒔 𝑸=𝑴^𝟎 𝑳^𝟎 𝒕^𝟎 " "
𝑴 ": " 𝒔=𝟎
𝑳 ": " 𝒒−𝟑𝒔+𝟑=𝟎
" " 𝒒=−𝟑
𝑻": "−𝒓−𝟏=𝟎
" " 𝒓=−𝟏
𝝅_𝟓=𝑸/(𝒏𝒅^𝟑 )=𝑵_𝑸
𝒌_𝑳 𝒂=𝑲_𝟏 (𝒉_(𝑳,𝒅)/𝒅)^𝜶 (𝑷/(𝝆𝒏^𝟑 𝒅^𝟓 ))^𝜷 ((𝝆𝒏𝒅^𝟐)/𝝁)^𝜸 (𝑽_𝑺/𝒏𝒅)^𝜹 (𝑸/(𝒏𝒅^𝟑 ))^𝜺
=
𝒌_𝑳 𝒂 =𝝅_𝟏, 𝝅_𝟐, 𝝅_𝟑, 𝝅_𝟒, 𝝅_𝟓
𝒌_𝑳 𝒂=𝑲_𝟏 (𝒉_(𝑳,𝒅)/𝒅)^𝜶 (𝑵_𝑷 )^𝜷 (𝑵_𝑹𝒆 )^𝜸 (𝑵_𝑽𝑺 )^𝜹 (𝑵_𝑸 )^𝜺
𝒌_𝑳 𝒂 (𝒉^(−𝟏))=𝟎.𝟑𝟔(𝒉_(𝑳,𝒅)/𝒅)^(𝟎.𝟎𝟖) (𝑷/(𝝆𝒏^𝟑 𝒅^𝟓 ))^(−𝟎.𝟗𝟎) ((𝝆𝒏𝒅^𝟐)/𝝁)^(−𝟎.𝟎𝟑) (𝑽_𝑺/𝒏𝒅)^(𝟎.𝟒𝟖) (𝑸/(𝒏𝒅^𝟑 ))^(𝟎.𝟕𝟗)
𝒌_𝑳 𝒂=𝟏.𝟐𝟗 (〖𝒉_(𝑳,𝒅)〗^(𝟎.𝟎𝟖) 〖𝑽_𝑺〗^(𝟎.𝟒𝟖) 𝑸^(𝟎.𝟕𝟗) 𝒏^(𝟏.𝟒𝟏))/𝑷^(𝟎.𝟗𝟎)
𝒅’=−𝟎.𝟎𝟖+(−𝟓×−𝟎.𝟗)+(𝟐×−𝟎.𝟎𝟑)+(−𝟏×𝟎.𝟒𝟖)+(−𝟑×𝟎.𝟕𝟗)
𝒅’=𝟏.𝟓𝟑
𝒏’=(−𝟑×−𝟎.𝟗)+(−𝟎.𝟎𝟑)+(−𝟏×𝟎.𝟒𝟖)+(−𝟏×𝟎.𝟕𝟗)
𝒏’=𝟏.𝟒𝟏
𝑲_𝟏’=𝑲_𝟏×𝝆^[(−𝟏×−𝟎.𝟗)+(−𝟎.𝟎𝟑)] ×𝝁^((−𝟏×−𝟎.𝟎𝟑))×𝒅^(𝟏.𝟓𝟑)
𝝆=𝟗𝟗𝟑.𝟑𝟑 𝒌𝒈.𝒎^(−𝟑), 𝝁=𝟎.𝟎𝟎𝟎𝟔𝟗𝟏𝟓 𝒌𝒈𝒎^(−𝟏) 𝒔^(−𝟏), 𝒅=𝟎.𝟎𝟓 𝒎
𝑲_𝟏’=𝟎.𝟑𝟔×〖𝟗𝟗𝟑.𝟑𝟑〗^[(−𝟏×−𝟎.𝟗)+(−𝟎.𝟎𝟑)] ×〖𝟎.𝟎𝟎𝟎𝟔𝟗𝟏𝟓〗^((−𝟏×−𝟎.𝟎𝟑))×〖𝟎.𝟎𝟓〗^(𝟏.𝟓𝟑)
𝑲_𝟏’=𝟏.𝟐𝟖𝟖𝟗

## 8
Sample calculations for the determination of mass transfer coefficient (KLa) for 2 L (42 % impeller imersion)at 100 rpm with 0.2 LPM air flow rate. | [11 empty] | Standard uncertainty |  | Square |  | Differentials
Estimation of KLa from DO-time curve |  |  |  |  | kLa (Exp) | 0.089213 | min-1 | Error |  |  | σT | 0.3 | 0.09 | K | dkH/dT | -1.6686e-07
|  | 1.3e-05 |  |  |  | 5.3528 | h-1 | 0.193 | h-1 |  | σkH | 5.0058e-08 | 2.5058e-15 | molPa-1m-3 | dC*/dkH | 21278.25
|  | 1500 |  |  | KLa (Model) | 4.7458 | h-1 |  |  |  | σC | 0.0022769 | 5.1842e-06 | mol m-3
[11 empty] | σt | 0.001 | 1e-06 | s
[6 empty] | Uncertainty associated with kLa measurements (σkLa) |  |  |  |  | σC* | 0.0035505 | 1.2606e-05 | mol m-3
[11 empty] | σCo | 0 | 0 | mol m-3
Steady state variables
|  | -11.4452 | [9 empty] | T | 310.15 |  | K
|  | 1.0701e-05 |  |  |  | Resonse time of DO probe | [5 empty] | C* | 0.22769 |  | mol m-3
[12 empty] | Co | 0 |  | mol m-3
Oxygen solubility at room temperature
|  | 0.22769 |  | 0.00022769
7.286
Parameters |  |  |  |  | Aeration tube calculations
Gass flow rate (LPM) | 0.2 |  |  |  | di | 0.0032 | m
Impeller speed (rpm) | 100 |  |  |  | Tube area | 8.0425e-06 | m2
Cs or C*O2 | 0.22769 |  |  |  | Slope | 0.079096
DO2 recording (min/reading) | K1 | 0.50267 |  |  | response time | 0.78
| K2 | 0.51454 |  |  | Error | 0.193
Average |  | 0.5086
| Superficial Gas velocity | 0.41447 | m/s |  |  |  | Cs-Co
0.22769
| Impeller speed | 1.6667 | rps | O2(%sat) mean | C (mol m-3) | Cs - C | LN(Cs-Co/ Cs-C) | dkLa/dC | dkLa/dC* | dkLa/dt | dkLa/dCo | σkLA | Cm | Error C^2
| Time (min) | O2 (%sat) 1 | O2 (%sat) 2 | [15 empty] | Table 3‑3 Experimental matrices for oxygen mass transfer
0 | 0 | 0 | 0 | 0 | 0 | 0.22769 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 0 | 0 |  | Volume |  | 2L |  |  |  | 3L
1 | 1.0172 | 0.2 | 0.4 | 0.3 | 0.00068307 | 0.22701 | 0.0030045 | 0.072178 | -0.00021653 | -8.0659e-07 | -0.071961 | 0.00016434 | 0.0078513 | 5.1384e-05 |  | Air flow rate [LPM] |  | 0.2 | 0.3 | 0.4 | 0.5 | 0.2 | 0.3 | 0.4 | 0.5
2 | 2.0344 | 0.9 | 2.3 | 1.6 | 0.003643 | 0.22405 | 0.016129 | 0.036566 | -0.00058505 | -1.0825e-06 | -0.035981 | 8.3282e-05 | 0.0222 | 0.00034435 |  | Impeller Speed (rpm) | 100 | Sample | x | x | x | x | x | x | x
3 | 3.0516 | 3 | 5.9 | 4.45 | 0.010132 | 0.21756 | 0.045521 | 0.025104 | -0.0011171 | -1.3578e-06 | -0.023987 | 5.7297e-05 | 0.037367 | 0.00074172 |  |  | 200 | bar: 2L 0.2 to 3L 0.5
4 | 4.0688 | 7.1 | 10.4 | 8.75 | 0.019923 | 0.20777 | 0.091567 | 0.019715 | -0.0017251 | -1.5364e-06 | -0.01799 | 4.5306e-05 | 0.051885 | 0.0010216 |  |  | 300 | bar: 2L 0.2 to 3L 0.5
5 | 5.086 | 11.6 | 15.6 | 13.6 | 0.030966 | 0.19672 | 0.14618 | 0.016658 | -0.0022654 | -1.5698e-06 | -0.014392 | 3.8771e-05 | 0.065422 | 0.0011873 |  |  | 400 | bar: 2L 0.2 to 3L 0.5
6 | 6.1032 | 16.8 | 21.3 | 19.05 | 0.043375 | 0.18431 | 0.21134 | 0.014816 | -0.0028224 | -1.576e-06 | -0.011994 | 3.5191e-05 | 0.077952 | 0.0011956 |  |  | 500 | bar: 2L 0.2 to 3L 0.5
7 | 7.1204 | 22.9 | 27.2 | 25.05 | 0.057036 | 0.17065 | 0.28835 | 0.013716 | -0.0034359 | -1.5798e-06 | -0.01028 | 3.3528e-05 | 0.089524 | 0.0010554 |  | Experimental results
8 | 8.1376 | 29.1 | 33.4 | 31.25 | 0.071153 | 0.15654 | 0.37469 | 0.013084 | -0.0040887 | -1.5717e-06 | -0.0089951 | 3.3139e-05 | 0.1002 | 0.00084395
9 | 9.1549 | 35.3 | 39.5 | 37.4 | 0.085156 | 0.14253 | 0.4684 | 0.012773 | -0.004777 | -1.5524e-06 | -0.0079957 | 3.3666e-05 | 0.11006 | 0.00062015 |  | Volume | bar: 2L 0.2; 3L 0.2
10 | 10.1721 | 41.4 | 45.2 | 43.3 | 0.098589 | 0.1291 | 0.5674 | 0.012692 | -0.0054954 | -1.5232e-06 | -0.0071961 | 3.4868e-05 | 0.11915 | 0.00042282 |  | Air flow rate [LPM] |  | 0.2 | 0.3 | 0.4 | 0.5 | 0.2 | 0.3 | 0.4 | 0.5
11 | 11.1893 | 47.3 | 51.1 | 49.2 | 0.11202 | 0.11567 | 0.67727 | 0.012878 | -0.0063359 | -1.5027e-06 | -0.0065419 | 3.6957e-05 | 0.12754 | 0.00024085 |  | Impeller Speed (rpm) | 100 | 5.35 | 6.64 | 7.42 | 8.22 | 5.34 | 5.91 | 6.04 | 6.78
12 | 12.2065 | 53.3 | 56.1 | 54.7 | 0.12455 | 0.10314 | 0.79186 | 0.013238 | -0.0072411 | -1.4763e-06 | -0.0059968 | 3.9617e-05 | 0.13528 | 0.00011531 |  |  | 200 | 6.92 | 8.75 | 10.03 | 9.97 | 6.89 | 8.22 | 8.54 | 9.68
13 | 13.2237 | 58.7 | 61.3 | 60 | 0.13661 | 0.091076 | 0.91629 | 0.013839 | -0.0083032 | -1.4555e-06 | -0.0055355 | 4.315e-05 | 0.14243 | 3.3806e-05 |  |  | 300 | 7.16 | 10.07 | 10.62 | 12.77 | 7.84 | 10.19 | 10.5 | 10.01
14 | 14.2409 | 63.9 | 66.3 | 65.1 | 0.14823 | 0.079463 | 1.0527 | 0.014728 | -0.009588 | -1.4419e-06 | -0.0051401 | 4.7785e-05 | 0.14902 | 6.2934e-07 |  |  | 400 | 7.64 | 10.85 | 12.97 | 12.54 | 8.94 | 10.85 | 12.06 | 13.17
15 | 15.2581 | 68.7 | 70.8 | 69.75 | 0.15881 | 0.068876 | 1.1957 | 0.015859 | -0.011062 | -1.4266e-06 | -0.0047974 | 5.3352e-05 | 0.1551 | 1.3783e-05 |  |  | 500 | 8.45 | 11.43 | 13.7 | 15.97 | 9.29 | 11.6 | 14.25 | 15.666
16 | 16.2753 | 73.2 | 75.3 | 74.25 | 0.16906 | 0.05863 | 1.3567 | 0.017466 | -0.012969 | -1.4228e-06 | -0.0044976 | 6.0842e-05 | 0.16071 | 6.9674e-05
17 | 17.2925 | 77.4 | 79.3 | 78.35 | 0.17839 | 0.049295 | 1.5302 | 0.019552 | -0.015319 | -1.4214e-06 | -0.004233 | 7.0286e-05 | 0.16589 | 0.00015637 |  | kLa by model to include oxygen probe time constant
18 | 18.3097 | 81.5 | 83.6 | 82.55 | 0.18796 | 0.039732 | 1.7458 | 0.02291 | -0.018912 | -1.4466e-06 | -0.0039978 | 8.503e-05 | 0.17067 | 0.00029895
19 | 19.3269 | 85 | 86.4 | 85.7 | 0.19513 | 0.032559 | 1.9449 | 0.026486 | -0.022698 | -1.4463e-06 | -0.0037874 | 0.00010065 | 0.17507 | 0.00040217 |  | Volume | bar: 2L 0.2; 3L 0.2
20 | 20.3441 | 89 | 90 | 89.5 | 0.20378 | 0.023907 | 2.2538 | 0.034267 | -0.030669 | -1.5126e-06 | -0.0035981 | 0.00013396 | 0.17914 | 0.00060709 |  | Air flow rate [LPM] |  | 0.2 | 0.3 | 0.4 | 0.5 | 0.2 | 0.3 | 0.4 | 0.5
21 | 21.3613 | 91.7 | 93.1 | 92.4 | 0.21038 | 0.017304 | 2.577 | 0.045088 | -0.041662 | -1.5688e-06 | -0.0034267 | 0.00018005 | 0.1829 | 0.00075566 |  | Impeller Speed (rpm) | 100 | 4.7457 | 6.0382 | 6.4921 | 6.4921 | 5.4021 | 4.8392 | 4.2438 | 5.9957
22 | 22.3785 | 94.7 | 95.9 | 95.3 | 0.21699 | 0.010701 | 3.0576 | 0.069595 | -0.066324 | -1.696e-06 | -0.003271 | 0.00028383 | 0.18636 | 0.00093816 |  |  | 200 | 6.3954 | 7.9987 | 9.37 | 9.37 | 7.0342 | 7.4384 | 7.2659 | 10.3727
23 | 23.3957 | 97.3 | 98.5 | 97.9 | 0.22291 | 0.0047815 | 3.8632 | 0.14899 | -0.14586 | -1.9605e-06 | -0.0031287 | 0.00061909 | 0.18955 | 0.0011125 |  |  | 300 | 6.9351 | 9.6448 | 10.6875 | 10.6875 | 7.8194 | 9.87 | 11.3683 | 7.6254
24 | 24.4129 | 99.7 | 100 | 99.85 | 0.22735 | 0.00034153 | 6.5023 | 1.9989 | -1.9959 | -3.0306e-06 | -0.0029984 | 0.0084222 | 0.1925 | 0.0012142 |  |  | 400 | 6.8736 | 10.3543 | 12.9341 | 12.9341 | 8.1887 | 12.9628 | 11.6251 | 10.3891
25 | 25.4301 | 100 |  | 100 | 0.22769 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -0.0028784 | #DIV/0! | 0.19522 | 0.0010541 |  |  | 500 | 8.0102 | 11.1003 | 13.2993 | 13.2993 | 8.4649 | 12.8165 | 12.2337 | 12.2448
[9 empty] | 5.352 | h-1
ANOVA
2L, 0.2L/min, 100rpm, Run1
| Gas Flow Rate |  | 0.2 | LPM | kLa | 0.088096 | min-1
| Impeller speed |  | 100 | rpm | kLa | 5.2857 | h-1
Cs-Co
0.22769
| Time (min) |  |  | O2(%sat) | C (mol/m3) | Csat - C | LN(Cs-Co/Cs-C)
| 0 |  |  | 0 | 0 | 0.22769 | 0
| 1.0053 |  |  | 0.2 | 0.00045538 | 0.22723 | 0.002002
| 2.0107 |  |  | 0.9 | 0.0020492 | 0.22564 | 0.0090407
| 3.016 |  |  | 3 | 0.0068307 | 0.22086 | 0.030459
| 4.0213 |  |  | 7.1 | 0.016166 | 0.21152 | 0.073647
| 5.0267 |  |  | 11.6 | 0.026412 | 0.20128 | 0.1233
| 6.032 |  |  | 16.8 | 0.038252 | 0.18944 | 0.18392
| 7.0373 |  |  | 22.9 | 0.052141 | 0.17555 | 0.26007
| 8.0427 |  |  | 29.1 | 0.066257 | 0.16143 | 0.3439
| 9.048 |  |  | 35.3 | 0.080374 | 0.14731 | 0.43541
| 10.0533 |  |  | 41.4 | 0.094263 | 0.13343 | 0.53444
| 11.0587 |  |  | 47.3 | 0.1077 | 0.11999 | 0.64055
| 12.064 |  |  | 53.3 | 0.12136 | 0.10633 | 0.76143
| 13.0693 |  |  | 58.7 | 0.13365 | 0.094035 | 0.88431
| 14.0747 |  |  | 63.9 | 0.14549 | 0.082196 | 1.0189
| 15.08 |  |  | 68.7 | 0.15642 | 0.071267 | 1.1616
| 16.0853 |  |  | 73.2 | 0.16667 | 0.061021 | 1.3168
| 17.0907 |  |  | 77.4 | 0.17623 | 0.051458 | 1.4872
| 18.096 |  |  | 81.5 | 0.18557 | 0.042122 | 1.6874
| 19.1013 |  |  | 85 | 0.19354 | 0.034153 | 1.8971
| 20.1067 |  |  | 89 | 0.20264 | 0.025046 | 2.2073
| 21.112 |  |  | 91.7 | 0.20879 | 0.018898 | 2.4889
| 22.1173 |  |  | 94.7 | 0.21562 | 0.012068 | 2.9375
| 23.1227 |  |  | 97.3 | 0.22154 | 0.0061476 | 3.6119
| 24.128 |  |  | 99.7 | 0.22701 | 0.00068307 | 5.8091
| 25.1334 |  |  | 100 | 0.22769 | 0 | #DIV/0!
2L, 0.2L/min, 100rpm, Run2
| Gas Flow Rate |  | 0.2 | LPM | kLa | 0.090401 | min-1
| Impeller speed |  | 100 | rpm | kLa | 5.4241 | h-1
Cs-Co
0.22769
| Time (min) |  |  | O2(%sat) | C (mol/m3) | Csat - C | LN(Cs-Co/Cs-C)
| 0 |  |  | 0 | 0 | 0.22769 | 0
| 1.0291 |  |  | 0.4 | 0.00091076 | 0.22678 | 0.004008
| 2.0582 |  |  | 2.3 | 0.0052368 | 0.22245 | 0.023269
| 3.0872 |  |  | 5.9 | 0.013434 | 0.21426 | 0.060812
| 4.1163 |  |  | 10.4 | 0.02368 | 0.20401 | 0.10981
| 5.1454 |  |  | 15.6 | 0.035519 | 0.19217 | 0.1696
| 6.1745 |  |  | 21.3 | 0.048498 | 0.17919 | 0.23953
| 7.2035 |  |  | 27.2 | 0.061931 | 0.16576 | 0.31745
| 8.2326 |  |  | 33.4 | 0.076048 | 0.15164 | 0.40647
| 9.2617 |  |  | 39.5 | 0.089937 | 0.13775 | 0.50253
| 10.2908 |  |  | 45.2 | 0.10292 | 0.12477 | 0.60148
| 11.3199 |  |  | 51.1 | 0.11635 | 0.11134 | 0.71539
| 12.3489 |  |  | 56.1 | 0.12773 | 0.099955 | 0.82326
| 13.378 |  |  | 61.3 | 0.13957 | 0.088116 | 0.94933
| 14.4071 |  |  | 66.3 | 0.15096 | 0.076731 | 1.0877
| 15.4362 |  |  | 70.8 | 0.1612 | 0.066485 | 1.231
| 16.4652 |  |  | 75.3 | 0.17145 | 0.056239 | 1.3984
| 17.4943 |  |  | 79.3 | 0.18056 | 0.047132 | 1.575
| 18.5234 |  |  | 83.6 | 0.19035 | 0.037341 | 1.8079
| 19.5525 |  |  | 86.4 | 0.19672 | 0.030966 | 1.9951
| 20.5816 |  |  | 90 | 0.20492 | 0.022769 | 2.3026
| 21.6106 |  |  | 93.1 | 0.21198 | 0.015711 | 2.6736
| 22.6397 |  |  | 95.9 | 0.21835 | 0.0093352 | 3.1942
| 23.6688 |  |  | 98.5 | 0.22427 | 0.0034153 | 4.1997
| 24.6979 |  |  | 100 | 0.22769 | 0 | #DIV/0!
| 25.7269 |  |  | 0 | 0 | 0.22769 | 0
𝑪_𝒎=𝑪^∗+(𝑪^∗−𝑪_𝒐)/(𝟏−𝝉_𝒓 𝒌_𝑳 𝒂) [𝝉_𝒓 𝒌_𝑳 𝒂.𝐞𝐱𝐩{(−(𝒕−𝒕_𝒐 ))/𝝉_𝒓 }−𝐞𝐱𝐩{−𝒌_𝑳 𝒂(𝒕−𝒕_𝒐 )}]
𝜎_𝑦^2=(𝜕𝑓/(𝜕𝑥_1 ))^2 𝜎_1^2+(𝜕𝑓/(𝜕𝑥_2 ))^2 𝜎_2^2+…(𝜕𝑓/(𝜕𝑥_𝑛 ))^2 𝜎_𝑛^2 " "
𝑘_𝐿 𝑎=0.0892 〖𝑚𝑖𝑛〗^(−1)=0.0892 〖𝑚𝑖𝑛〗^(−1)×60𝑚𝑖𝑛/1ℎ=5.352759319
𝒌_𝑯 (𝟐𝟗𝟖.𝟏𝟓𝑲)
(𝒅 𝒍𝒏⁡〖(𝒌_𝑯 )_𝑪𝒑 〗)/𝒅(𝟏/𝑻)
𝒍𝒏⁡〖𝒌_𝑯 (𝟑𝟏𝟎.𝟏𝟓𝑲)〗=[𝒍𝒏⁡(𝟏.𝟑×〖𝟏𝟎〗^(−𝟓 ) ) ]+𝟏𝟓𝟎𝟎(𝟏/(𝟑𝟏𝟎.𝟏𝟓)−𝟏/(𝟐𝟗𝟖.𝟏𝟓))
𝒎𝒐𝒍.𝒎^(−𝟑) 〖𝑷𝒂〗^(−𝟏)
𝒍𝒏⁡〖𝒌_𝑯 (𝟑𝟏𝟎.𝟏𝟓𝑲)〗
𝒎𝒐𝒍.𝒎^(−𝟑) 〖𝑷𝒂〗^(−𝟏)
𝒌_𝑯 (𝟑𝟏𝟎.𝟏𝟓𝑲)
𝑪_𝑶𝟐^∗=𝒌_𝑯 𝑷_(𝑨,𝒔𝒂𝒕)=〖𝒌_𝑯 𝒚〗_𝑨 𝑷_𝒕𝒐𝒕𝒂𝒍
𝑪_𝑶𝟐^∗
𝒎𝒐𝒍.𝒎^(−𝟑)
𝒎𝒐𝒍.𝑳^(−𝟏)
𝒎𝒈.𝑳^(−𝟏)
𝒎𝒐𝒍.𝒎^(−𝟑)

## 9
Sample calculations for overall heat transfer coefficient (UA) for HTB at 2 L and 3 L volume
| Rep1 | Rep2 | Avg.
m | 2 | 2 | kg | [7 empty] | m | 3 | 3 | kg
| 4180 |  | J/kg-k | [8 empty] | 4180 |  | J/kg-k
T1 | 328 | 322 | k | [7 empty] | T1 | 323 | 325.5 | k
T2 | 301 | 300.5 | k | [7 empty] | T2 | 302 | 302 | k
T∞ | 297.1651 | 297.175 | k | sum of square errors | 0.86854 |  |  |  | sum of square errors |  | T∞ | 296.2333 | 296.275 | k | sum of square errors | 36.3483 |  |  |  | sum of square errors
UA | 0.20923 | 0.21432 | 0.21177 | [6 empty] | 0.70979 | UA | 0.23457 | 0.27202 | 0.25329 | [6 empty] | 36.1244
T∞ oC | 24.1651 | 24.175 | [8 empty] | T∞ oC | 23.6164 | 23.2286
Time (h) | T∞ (oC) | T (oC) | Predicted Temp. | SSE |  | Time (h) | T∞ (oC) | T (oC) | Predicted Temp. | SSE | Time (h) | T∞ (oC) | T (oC) | Predicted Temp. | SSE |  | Time (h) | T∞ (oC) | T (oC) | Predicted Temp. | SSE
0 | 24.0858 | 55 | 55 | 0 |  | 0 | 25.2 | 49 | 49 | 0 | 0 | 23.4 | 50 | 50 | 0 |  | 0 | 23.7 | 52.5 | 52.5 | 0
1 | 23.6979 | 52.5 | 52.3433 | 0.024547 |  | 1 | 24.8 | 47 | 46.8114 | 0.035568 | 1 | 23.2 | 47.5 | 48.2818 | 0.61126 |  | 1 | 23.4 | 48 | 50.3011 | 5.2951
2 | 23.0314 | 49.5 | 49.9155 | 0.17268 |  | 2 | 24 | 45 | 44.8158 | 0.033944 | 2 | 23.1 | 46 | 46.6756 | 0.45638 |  | 2 | 23.3 | 46 | 48.2674 | 5.1411
3 | 23.8468 | 47.5 | 47.6969 | 0.038785 |  | 3 | 23.8 | 43 | 42.9961 | 1.5586e-05 | 3 | 23.1 | 44.5 | 45.1739 | 0.45412 |  | 3 | 23.1 | 44 | 46.3865 | 5.6952
4 | 23.5784 | 45.5 | 45.6695 | 0.028724 |  | 4 | 23.2 | 41.5 | 41.3368 | 0.026644 | 4 | 23 | 43 | 43.77 | 0.59291 |  | 4 | 23 | 43 | 44.6468 | 2.712
5 | 24.7401 | 43.5 | 43.8167 | 0.1003 |  | 5 | 24.5 | 39.5 | 39.8238 | 0.10483 | 5 | 23 | 41.5 | 42.4576 | 0.91691 |  | 5 | 23 | 41.5 | 43.0379 | 2.365
6 | 24.5694 | 42 | 42.1236 | 0.015268 |  | 6 | 24.2 | 38.5 | 38.4442 | 0.0031176 | 6 | 23 | 40.5 | 41.2306 | 0.53373 |  | 6 | 23.1 | 40 | 41.5498 | 2.4018
7 | 24.623 | 40.5 | 40.5763 | 0.005821 |  | 7 | 24.6 | 37 | 37.1862 | 0.034664 | 7 | 23 | 39 | 40.0835 | 1.174 |  | 7 | 23.2 | 39 | 40.1735 | 1.377
8 | 24.2178 | 39 | 39.1623 | 0.026354 |  | 8 | 24 | 36 | 36.0391 | 0.0015292 | 8 | 23 | 38 | 39.0111 | 1.0223 |  | 8 | 23.4 | 38 | 38.9006 | 0.811
9 | 24.5995 | 38 | 37.8702 | 0.016847 |  | 9 | 23.6 | 35 | 34.9932 | 4.6861e-05 | 9 | 23 | 37 | 38.0086 | 1.0172 |  | 9 | 23.5 | 37 | 37.7233 | 0.52311
10 | 24.7799 | 36.5 | 36.6894 | 0.035872 |  | 10 | 24.2 | 34 | 34.0394 | 0.0015536 | 10 | 23.1 | 36 | 37.0713 | 1.1477 |  | 10 | 23.5 | 36.5 | 36.6344 | 0.018067
11 | 24.8243 | 35.5 | 35.6103 | 0.012172 |  | 11 | 24.1 | 33 | 33.1698 | 0.028819 | 11 | 23.1 | 35.5 | 36.1951 | 0.48316 |  | 11 | 23.5 | 35.5 | 35.6274 | 0.016219
12 | 24.0632 | 34.5 | 34.6242 | 0.015433 |  | 12 | 24.6 | 32 | 32.3768 | 0.14196 | 12 | 23.2 | 34.5 | 35.3759 | 0.76728 |  | 12 | 23.4 | 34.5 | 34.6959 | 0.038396
13 | 24.1682 | 34 | 33.7231 | 0.07668 |  | 13 | 23.8 | 31.5 | 31.6537 | 0.023623 | 13 | 23.2 | 34 | 34.6101 | 0.37226 |  | 13 | 23.4 | 34 | 33.8345 | 0.027387
14 | 23.187 | 33 | 32.8996 | 0.010082 |  | 14 | 23.2 | 31 | 30.9944 | 3.1698e-05 | 14 | 23.3 | 33 | 33.8942 | 0.79958 |  | 14 | 23.4 | 33.5 | 33.0378 | 0.21364
15 | 24.789 | 32 | 32.147 | 0.021621 |  | 15 | 23.6 | 30.5 | 30.3932 | 0.011413 | 15 | 23.3 | 32.5 | 33.2249 | 0.52545 |  | 15 | 23.3 | 32.5 | 32.3009 | 0.039638
16 | 23.9084 | 31.5 | 31.4593 | 0.001654 |  | 16 | 24 | 30 | 29.845 | 0.024035 | 16 | 23.4 | 32 | 32.5992 | 0.35898 |  | 16 | 23.2 | 32 | 31.6194 | 0.14487
17 | 24.4384 | 31 | 30.8309 | 0.028604 |  | 17 | 24.4 | 29.5 | 29.3451 | 0.023994 | 17 | 23.4 | 31.5 | 32.0142 | 0.26437 |  | 17 | 23.1 | 31.5 | 30.9891 | 0.26106
18 | 24.8639 | 30.5 | 30.2566 | 0.059263 |  | 18 | 24.8 | 29 | 28.8893 | 0.012255 | 18 | 23.4 | 31 | 31.4673 | 0.21836 |  | 18 | 23 | 31 | 30.4061 | 0.35273
19 | 23.141 | 30 | 29.7317 | 0.071969 |  | 19 | 24.8 | 28.5 | 28.4737 | 0.00069264 | 19 | 23.4 | 30.5 | 30.956 | 0.20796 |  | 19 | 23.2 | 30.5 | 29.8669 | 0.40081
20 | 24.8416 | 29.5 | 29.2521 | 0.061446 |  | 20 | 24.6 | 28 | 28.0947 | 0.0089693 | 20 | 23.4 | 30 | 30.478 | 0.22853 |  | 20 | 23.4 | 30 | 29.3682 | 0.39913
21 | 23.8777 | 29 | 28.8138 | 0.03466 |  | 21 | 24.2 | 28 | 27.7491 | 0.06293 | 21 | 23.5 | 30 | 30.0312 | 0.00097362 |  | 21 | 23.3 | 29.5 | 28.907 | 0.35163
22 | 24.1761 | 28.5 | 28.4133 | 0.0075169 |  | 22 | 24 | 27.5 | 27.434 | 0.0043505 | 22 | 23.5 | 29.5 | 29.6135 | 0.012872 |  | 22 | 23.2 | 29.5 | 28.4804 | 1.0395
23 | 23.9128 | 28 | 28.0473 | 0.0022355 |  | 23 | 24 | 27.5 | 27.1467 | 0.12481 | 23 | 23.6 | 29 | 29.2229 | 0.049691 |  | 23 | 23 | 29 | 28.0859 | 0.83554
|  |  | Sum of errors | 0.86854 | [6 empty] | 24 | 24.2 | 29 | 28.8578 | 0.020219 |  | 24 | 23 | 28.5 | 27.721 | 0.60679
[11 empty] | 25 | 24.4 | 28.5 | 28.5165 | 0.0002714 |  | 25 | 22.9 | 28.5 | 27.3836 | 1.2465
[11 empty] | 26 | 24.2 | 28.5 | 28.1974 | 0.091584 |  | 26 | 22.8 | 28 | 27.0714 | 0.86224
[11 empty] | 27 | 24.6 | 28.5 | 27.899 | 0.36114 |  | 27 | 23.1 | 28.5 | 26.7828 | 2.9489
[11 empty] | 28 | 24.4 | 28 | 27.6202 | 0.14428 | [5 empty] | 36.1244
[11 empty] | 29 | 24.3 | 28 | 27.3594 | 0.41034
[11 empty] | 30 | 24.1 | 27.5 | 27.1157 | 0.14771
[11 empty] | 31 | 24.1 | 27.5 | 26.8878 | 0.3748
[11 empty] | 32 | 24 | 27.5 | 26.6747 | 0.68104
[11 empty] | 33 | 24 | 27 | 26.4756 | 0.27501
[11 empty] | 34 | 24 | 27 | 26.2894 | 0.50497
[11 empty] | 35 | 23.9 | 27 | 26.1153 | 0.78266
[11 empty] | 36 | 23.9 | 27 | 25.9526 | 1.0971
[11 empty] | 37 | 23.9 | 26.5 | 25.8004 | 0.48937
[11 empty] | 38 | 23.9 | 26.5 | 25.6582 | 0.70859
[11 empty] | 39 | 23.8 | 26.5 | 25.5253 | 0.95013
[11 empty] | 40 | 23.8 | 26 | 25.4009 | 0.35886
[11 empty] | 41 | 23.8 | 26 | 25.2847 | 0.5116
[11 empty] | 42 | 23.8 | 26 | 25.1761 | 0.67883
[11 empty] | 43 | 23.7 | 26 | 25.0745 | 0.85651
[11 empty] | 44 | 23.7 | 26 | 24.9796 | 1.0413
[11 empty] | 45 | 23.7 | 25.5 | 24.8908 | 0.37113
[11 empty] | 46 | 23.7 | 25.5 | 24.8078 | 0.47913
[11 empty] | 47 | 23.7 | 25.5 | 24.7302 | 0.59256
[11 empty] | 48 | 23.7 | 25.5 | 24.6577 | 0.70949
[11 empty] | 49 | 23.7 | 25.5 | 24.5899 | 0.82832
[11 empty] | 50 | 23.8 | 25.5 | 24.5265 | 0.94773
[11 empty] | 51 | 23.8 | 25 | 24.4672 | 0.28385
[11 empty] | 52 | 23.8 | 25 | 24.4118 | 0.34596
[11 empty] | 53 | 23.8 | 25 | 24.36 | 0.40957
[11 empty] | 54 | 23.8 | 25 | 24.3116 | 0.4739
[11 empty] | 55 | 23.8 | 25 | 24.2663 | 0.53828
[11 empty] | 56 | 23.8 | 25 | 24.224 | 0.60217
[11 empty] | 57 | 23.7 | 25 | 24.1844 | 0.66514
[11 empty] | 58 | 23.7 | 25 | 24.1474 | 0.72685
[11 empty] | 59 | 23.7 | 25 | 24.1129 | 0.78701
[11 empty] | 60 | 23.7 | 25 | 24.0805 | 0.84541
[11 empty] | 61 | 23.7 | 24.5 | 24.0503 | 0.20222
[11 empty] | 62 | 23.6 | 24.5 | 24.0221 | 0.22843
[11 empty] | 63 | 23.6 | 24.5 | 23.9956 | 0.25437
[11 empty] | 64 | 23.6 | 24.5 | 23.9709 | 0.27989
[11 empty] | 65 | 23.6 | 24.5 | 23.9479 | 0.30486
[11 empty] | 66 | 23.5 | 24.5 | 23.9263 | 0.32916
[11 empty] | 67 | 23.4 | 24.5 | 23.9061 | 0.35271
[11 empty] | 68 | 23.4 | 24.5 | 23.8872 | 0.37548
[11 empty] | 69 | 23.4 | 24.5 | 23.8696 | 0.3974
[11 empty] | 70 | 23.4 | 24.5 | 23.8531 | 0.41846
[11 empty] | 71 | 23.4 | 24.5 | 23.8377 | 0.43864
[11 empty] | 72 | 23.4 | 24.5 | 23.8233 | 0.45793
36.3483
𝐶_𝑝
𝐶_𝑝

## 10
Sample calculation for power utilisation in HTB at 2 L and 3 L with 200 rpm
m | 3 | kg | [5 empty] | m | 3 | kg |  |  |  |  | m | 2 | kg |  |  |  |  | m | 2 | kg
| 4184 | J kg-1 K-1 | [6 empty] | 4184 | J kg-1 K-1 | [5 empty] | 4184 | J kg-1 K-1 | [5 empty] | 4184 | J kg-1 K-1
UA | 0.25329 | W K-1 | [5 empty] | UA | 0.25329 | W K-1 |  |  |  |  | UA | 0.21177 | W K-1 |  |  |  |  | UA | 0.21177 | W K-1
P | 0.15 | W | [5 empty] | P | 0.18091 | W |  |  |  |  | P | 0.086762 | W |  |  |  |  | P | 0.088831 | W
T∞ | 23.1775 | K | [5 empty] | T∞ | 23.9935 | K |  |  |  |  | T∞ | 24.7917 | K |  |  |  |  | T∞ | 26 | K
a | -49555.0191 | [6 empty] | a | -49555.0191 | [5 empty] | a | -39513.795 | [5 empty] | a | -39513.795 |  |  |  |  | 5.6502
τ | 26.2303 | K |  |  |  | 1.4549 |  | τ | 19.7922 | K |  |  |  | 0.015702 | τ | 24.7986 | K |  |  |  |  | τ | 18.0805 | K
Time | T∞ | Temperature | τ Exp | τ model | T model | SSE |  | Time | T∞ | Temperature | τ Exp | τ model | T model | SSE | Time | T∞ | Temperature | Touexp | Toumodel | Tmodel | SSE | Time | T∞ | Temperature | Touexp | Toumodel | Tmodel | SSE
0 | 25.2 | 50 | 26.2303 | 26.2303 | 50 | 0 |  | 0 | 26.4 | 44.5 | 19.7922 | 19.7922 | 44.5 | 0 | 0 | 28.5 | 50 | 24.7986 | 24.7986 | 50 | 0 | 0 | 29.4 | 44.5 | 18.0805 | 18.0805 | 44.5 | 0
1 | 24.6 | 48 | 24.2303 | 24.3923 | 48.162 | 0.026253 |  | 1 | 24.8 | 43 | 18.2922 | 18.4054 | 43.1131 | 0.012803 | 0.5 | 28 | 49 | 23.7986 | 23.6943 | 48.8957 | 0.010884 | 0.5 | 28.5 | 44.5 | 18.0805 | 17.2754 | 43.6948 | 0.64828
2 | 24.2 | 46.5 | 22.7303 | 22.6831 | 46.4528 | 0.0022235 |  | 2 | 24.5 | 42 | 17.2922 | 17.1157 | 41.8235 | 0.031161 | 1 | 28 | 47.5 | 22.2986 | 22.6392 | 47.8405 | 0.11596 | 1 | 27.4 | 43.5 | 17.0805 | 16.5061 | 42.9255 | 0.33
3 | 24 | 45 | 21.2303 | 21.0937 | 44.8634 | 0.018653 |  | 3 | 25 | 40.5 | 15.7922 | 15.9164 | 40.6242 | 0.015418 | 1.5 | 27.5 | 46.5 | 21.2986 | 21.631 | 46.8324 | 0.11047 | 1.5 | 26.8 | 43 | 16.5805 | 15.771 | 42.1905 | 0.6553
4 | 23.6 | 43.5 | 19.7303 | 19.6157 | 43.3854 | 0.013138 |  | 4 | 24.6 | 39.5 | 14.7922 | 14.8011 | 39.5089 | 7.9214e-05 | 2 | 27.5 | 45.5 | 20.2986 | 20.6677 | 45.8691 | 0.13623 | 2 | 26.4 | 42 | 15.5805 | 15.0687 | 41.4882 | 0.26195
5 | 24 | 42 | 18.2303 | 18.2412 | 42.0109 | 0.00011872 |  | 5 | 24.8 | 38.5 | 13.7922 | 13.764 | 38.4718 | 0.00079647 | 2.5 | 27 | 44.5 | 19.2986 | 19.7474 | 44.9487 | 0.20136 | 2.5 | 26.4 | 41.5 | 15.0805 | 14.3977 | 40.8171 | 0.46629
6 | 24.4 | 41 | 17.2303 | 16.963 | 40.7327 | 0.071436 |  | 6 | 23.8 | 37.5 | 12.7922 | 12.7995 | 37.5073 | 5.3697e-05 | 3 | 27 | 43.5 | 18.2986 | 18.868 | 44.0693 | 0.32415 | 3 | 25.8 | 40.5 | 14.0805 | 13.7565 | 40.176 | 0.10498
7 | 24.6 | 39.5 | 15.7303 | 15.7744 | 39.5441 | 0.0019463 |  | 7 | 23.6 | 36.5 | 11.7922 | 11.9027 | 36.6105 | 0.012201 | 3.5 | 26.5 | 42.5 | 17.2986 | 18.0278 | 43.2291 | 0.53162 | 3.5 | 26 | 40 | 13.5805 | 13.1439 | 39.5634 | 0.19063
8 | 25 | 38.5 | 14.7303 | 14.6691 | 38.4388 | 0.003746 |  | 8 | 23.5 | 36 | 11.2922 | 11.0686 | 35.7764 | 0.049983 | 4 | 26.5 | 42 | 16.7986 | 17.225 | 42.4263 | 0.18174 | 4 | 25.5 | 39.5 | 13.0805 | 12.5586 | 38.9781 | 0.27241
9 | 25.4 | 37.5 | 13.7303 | 13.6412 | 37.4109 | 0.0079345 |  | 9 | 23 | 35 | 10.2922 | 10.2931 | 35.0008 | 7.1305e-07 | 4.5 | 26 | 41 | 15.7986 | 16.4579 | 41.6593 | 0.43462 | 4.5 | 25.8 | 39 | 12.5805 | 11.9994 | 38.4188 | 0.33778
10 | 25 | 36.5 | 12.7303 | 12.6854 | 36.4551 | 0.0020182 |  | 10 | 23.6 | 34.5 | 9.7922 | 9.5718 | 34.2796 | 0.048574 | 5 | 25.5 | 40.5 | 15.2986 | 15.725 | 40.9264 | 0.18178 | 5 | 26.6 | 38 | 11.5805 | 11.465 | 37.8845 | 0.013349
11 | 24.5 | 36 | 12.2303 | 11.7965 | 35.5662 | 0.18818 |  | 11 | 23.2 | 33.5 | 8.7922 | 8.9011 | 33.6089 | 0.01186 | 5.5 | 25 | 39.5 | 14.2986 | 15.0247 | 40.2261 | 0.52721 | 5.5 | 25.8 | 37.5 | 11.0805 | 10.9544 | 37.3739 | 0.0159
12 | 24 | 35 | 11.2303 | 10.9699 | 34.7396 | 0.067799 |  | 12 | 23.6 | 33 | 8.2922 | 8.2774 | 32.9852 | 0.00021917 | 6 | 25 | 39 | 13.7986 | 14.3557 | 39.557 | 0.31027 | 6 | 25.4 | 37 | 10.5805 | 10.4666 | 36.8861 | 0.012976
13 | 24 | 34 | 10.2303 | 10.2013 | 33.9709 | 0.00084394 |  | 13 | 23.8 | 32.5 | 7.7922 | 7.6974 | 32.4052 | 0.0089884 | 6.5 | 24.5 | 38.5 | 13.2986 | 13.7164 | 38.9177 | 0.1745 | 6.5 | 25.6 | 36.5 | 10.0805 | 10.0005 | 36.42 | 0.0064017
14 | 24.4 | 33.5 | 9.7303 | 9.4864 | 33.2561 | 0.059467 |  | 14 | 24.4 | 32 | 7.2922 | 7.158 | 31.8658 | 0.018001 | 7 | 24.5 | 38 | 12.7986 | 13.1056 | 38.3069 | 0.094201 | 7 | 25.6 | 36 | 9.5805 | 9.5552 | 35.9746 | 0.00064269
15 | 24.4 | 33 | 9.2303 | 8.8217 | 32.5914 | 0.16694 |  | 15 | 24.4 | 31.5 | 6.7922 | 6.6565 | 31.3643 | 0.018424 | 7.5 | 25 | 37.5 | 12.2986 | 12.522 | 37.7233 | 0.049867 | 7.5 | 25.6 | 36 | 9.5805 | 9.1297 | 35.5491 | 0.20328
16 | 25 | 32 | 8.2303 | 8.2036 | 31.9733 | 0.00071401 |  | 16 | 24.2 | 31 | 6.2922 | 6.1901 | 30.8978 | 0.010437 | 8 | 24.5 | 37 | 11.7986 | 11.9643 | 37.1657 | 0.027451 | 8 | 25 | 35.5 | 9.0805 | 8.7231 | 35.1426 | 0.12775
17 | 24.6 | 31.5 | 7.7303 | 7.6288 | 31.3984 | 0.010312 |  | 17 | 23.6 | 30.5 | 5.7922 | 5.7563 | 30.4641 | 0.0012888 | 8.5 | 24 | 36 | 10.7986 | 11.4315 | 36.6329 | 0.40055 | 8.5 | 25.7 | 35 | 8.5805 | 8.3347 | 34.7541 | 0.060455
18 | 24.6 | 31 | 7.2303 | 7.0942 | 30.8639 | 0.018523 |  | 18 | 23.6 | 30 | 5.2922 | 5.353 | 30.0608 | 0.0036907 | 9 | 24 | 36 | 10.7986 | 10.9225 | 36.1238 | 0.015333 | 9 | 25.3 | 34.5 | 8.0805 | 7.9635 | 34.383 | 0.013697
19 | 24 | 30.5 | 6.7303 | 6.5971 | 30.3668 | 0.017741 |  | 19 | 23.6 | 29.5 | 4.7922 | 4.9779 | 29.6857 | 0.034472 | 9.5 | 24.5 | 35.5 | 10.2986 | 10.4361 | 35.6374 | 0.018887 | 9.5 | 25.4 | 34 | 7.5805 | 7.6089 | 34.0283 | 0.00080307
20 | 24 | 30 | 6.2303 | 6.1348 | 29.9045 | 0.0091123 |  | 20 | 23 | 29.5 | 4.7922 | 4.6291 | 29.3369 | 0.026614 | 10 | 24.5 | 35 | 9.7986 | 9.9713 | 35.1727 | 0.029823 | 10 | 25.7 | 34 | 7.5805 | 7.27 | 33.6895 | 0.096409
21 | 23.5 | 29.5 | 5.7303 | 5.705 | 29.4747 | 0.00064161 |  | 21 | 23.2 | 29 | 4.2922 | 4.3047 | 29.0125 | 0.00015628 | 10.5 | 25 | 34.5 | 9.2986 | 9.5273 | 34.7287 | 0.052282 | 10.5 | 26.5 | 33.5 | 7.0805 | 6.9463 | 33.3658 | 0.018022
22 | 23.8 | 29 | 5.2303 | 5.3052 | 29.0749 | 0.0056129 |  | 22 | 23.4 | 28.5 | 3.7922 | 4.0031 | 28.7109 | 0.044465 | 11 | 24 | 34 | 8.7986 | 9.103 | 34.3044 | 0.092651 | 11 | 26.5 | 33 | 6.5805 | 6.637 | 33.0564 | 0.0031837
23 | 24 | 29 | 5.2303 | 4.9335 | 28.7032 | 0.088102 |  | 23 | 23.4 | 28.5 | 3.7922 | 3.7226 | 28.4304 | 0.0048485 | 11.5 | 24 | 33.5 | 8.2986 | 8.6977 | 33.899 | 0.15921 | 11.5 | 26.5 | 33 | 6.5805 | 6.3414 | 32.7609 | 0.057184
24 | 23.4 | 28.5 | 4.7303 | 4.5878 | 28.3575 | 0.02031 |  | 24 | 24 | 28 | 3.2922 | 3.4617 | 28.1695 | 0.028739 | 12 | 25 | 33.5 | 8.2986 | 8.3103 | 33.5117 | 0.00013667 | 12 | 26.8 | 32.5 | 6.0805 | 6.059 | 32.4785 | 0.00046332
25 | 23.2 | 28 | 4.2303 | 4.2663 | 28.036 | 0.0012974 |  | 25 | 24 | 28 | 3.2922 | 3.2192 | 27.927 | 0.005335 | 12.5 | 24.5 | 33 | 7.7986 | 7.9403 | 33.1416 | 0.020055 | 12.5 | 26.4 | 32.5 | 6.0805 | 5.7892 | 32.2087 | 0.084881
26 | 23.4 | 28 | 4.2303 | 3.9674 | 27.7371 | 0.069129 |  | 26 | 24 | 27.5 | 2.7922 | 2.9936 | 27.7014 | 0.040558 | 13 | 25 | 32.5 | 7.2986 | 7.5867 | 32.788 | 0.082958 | 13 | 26.2 | 32 | 5.5805 | 5.5314 | 31.9509 | 0.0024153
27 | 20.9 | 27.5 | 3.7303 | 3.6894 | 27.4591 | 0.0016745 |  | 27 | 23.4 | 27.5 | 2.7922 | 2.7838 | 27.4916 | 7.0121e-05 | 13.5 | 24 | 32.5 | 7.2986 | 7.2488 | 32.4502 | 0.0024824 | 13.5 | 25.7 | 32 | 5.5805 | 5.2851 | 31.7045 | 0.087302
28 | 20.8 | 27 | 3.2303 | 3.4309 | 27.2006 | 0.040225 |  | 28 | 24.2 | 27.5 | 2.7922 | 2.5888 | 27.2966 | 0.041388 | 14 | 24 | 32 | 6.7986 | 6.926 | 32.1274 | 0.016224 | 14 | 25.5 | 31.5 | 5.0805 | 5.0497 | 31.4692 | 0.00094998
29 | 20.9 | 27 | 3.2303 | 3.1905 | 26.9602 | 0.0015872 |  | 29 | 24.6 | 27 | 2.2922 | 2.4074 | 27.1152 | 0.013263 | 14.5 | 25 | 32 | 6.7986 | 6.6176 | 31.8189 | 0.03278 | 14.5 | 25.2 | 31.5 | 5.0805 | 4.8248 | 31.2443 | 0.06538
30 | 20.9 | 26.5 | 2.7303 | 2.9669 | 26.7366 | 0.055981 |  | 30 | 24.6 | 27 | 2.2922 | 2.2387 | 26.9465 | 0.0028647 | 15 | 22 | 31.5 | 6.2986 | 6.3229 | 31.5243 | 0.00058825 | 15 | 24.6 | 31 | 4.5805 | 4.61 | 31.0294 | 0.00086713
31 | 20.9 | 26.5 | 2.7303 | 2.759 | 26.5287 | 0.00082425 | [8 empty] | 15.5 | 23 | 31.5 | 6.2986 | 6.0413 | 31.2427 | 0.066211 | 15.5 | 24.6 | 31 | 4.5805 | 4.4047 | 30.8242 | 0.030921
32 | 20.9 | 26.5 | 2.7303 | 2.5657 | 26.3354 | 0.027098 | [8 empty] | 16 | 24 | 31 | 5.7986 | 5.7723 | 30.9737 | 0.00069414 | 16 | 25 | 30.5 | 4.0805 | 4.2085 | 30.628 | 0.016386
33 | 21 | 26 | 2.2303 | 2.3859 | 26.1556 | 0.024213 | [8 empty] | 16.5 | 23 | 31 | 5.7986 | 5.5152 | 30.7166 | 0.080314 | 16.5 | 25 | 30.5 | 4.0805 | 4.0211 | 30.4406 | 0.0035291
34 | 21 | 26 | 2.2303 | 2.2187 | 25.9884 | 0.00013402 | [8 empty] | 17 | 25 | 30.5 | 5.2986 | 5.2696 | 30.471 | 0.00084101 | 17 | 25.5 | 30 | 3.5805 | 3.8421 | 30.2615 | 0.068396
35 | 21 | 25.5 | 1.7303 | 2.0633 | 25.833 | 0.11086 | [8 empty] | 17.5 | 22 | 30.5 | 5.2986 | 5.035 | 30.2363 | 0.06952 | 17.5 | 25.5 | 30 | 3.5805 | 3.671 | 30.0904 | 0.008178
36 | 21 | 25.5 | 1.7303 | 1.9187 | 25.6884 | 0.035488 | [8 empty] | 18 | 25 | 30 | 4.7986 | 4.8108 | 30.0121 | 0.00014683 | 18 | 25 | 30 | 3.5805 | 3.5075 | 29.927 | 0.0053352
37 | 21 | 25.5 | 1.7303 | 1.7842 | 25.5539 | 0.0029094 | [8 empty] | 18.5 | 25 | 30 | 4.7986 | 4.5965 | 29.7979 | 0.04085 | 18.5 | 25.8 | 29.5 | 3.0805 | 3.3513 | 29.7708 | 0.073312
38 | 21 | 25 | 1.2303 | 1.6592 | 25.4289 | 0.18397 | [8 empty] | 19 | 23 | 30 | 4.7986 | 4.3918 | 29.5932 | 0.16549 | 19 | 26 | 29.5 | 3.0805 | 3.2021 | 29.6215 | 0.014768
39 | 21 | 25 | 1.2303 | 1.543 | 25.3127 | 0.097752 | [8 empty] | 19.5 | 23 | 29.5 | 4.2986 | 4.1963 | 29.3976 | 0.010482 | 19.5 | 26 | 29.5 | 3.0805 | 3.0595 | 29.4789 | 0.00044394
[6 empty] | 1.4549 | [8 empty] | 20 | 25 | 29.5 | 4.2986 | 4.0094 | 29.2108 | 0.083665 | 20 | 26.4 | 29 | 2.5805 | 2.9232 | 29.3427 | 0.11743
[15 empty] | 20.5 | 22 | 29.5 | 4.2986 | 3.8309 | 29.0322 | 0.21883 | 20.5 | 26 | 29 | 2.5805 | 2.793 | 29.2125 | 0.045161
[15 empty] | 21 | 23 | 29 | 3.7986 | 3.6603 | 28.8616 | 0.019151 | 21 | 26.2 | 29 | 2.5805 | 2.6687 | 29.0881 | 0.0077671
[15 empty] | 21.5 | 24 | 29 | 3.7986 | 3.4973 | 28.6986 | 0.090833 | 21.5 | 25.8 | 28.5 | 2.0805 | 2.5498 | 28.9693 | 0.22023
[15 empty] | 22 | 22 | 29 | 3.7986 | 3.3415 | 28.5429 | 0.20896 | 22 | 26 | 28.5 | 2.0805 | 2.4363 | 28.8557 | 0.12655
[15 empty] | 22.5 | 25 | 28.5 | 3.2986 | 3.1927 | 28.3941 | 0.011221 | 22.5 | 26.6 | 28.5 | 2.0805 | 2.3278 | 28.7473 | 0.061133
[15 empty] | 23 | 25 | 28.5 | 3.2986 | 3.0505 | 28.2519 | 0.061556 | 23 | 26.6 | 28 | 1.5805 | 2.2241 | 28.6436 | 0.41421
[15 empty] | 23.5 | 25 | 28.5 | 3.2986 | 2.9147 | 28.116 | 0.14742 | 23.5 | 26.4 | 28 | 1.5805 | 2.1251 | 28.5445 | 0.29653
[21 empty] | 5.6125 | [6 empty] | 5.6502
Power table for all the volumes and at all the impeller speeds
Impeller immersion | Volume (L) | Volume (m3) | Parameters | Power (W)
|  |  |  | Rep 1 | Rep2 | Rep 1 | Rep2 | Rep 1 | Rep2 | Rep 1 | Rep2
|  |  |  | 200 |  | 300 |  | 400 |  | 500
42% | 2 L | 0.002 | UA | 0.21177 | 0.21177 | 0.212 | 0.212 | 0.212 | 0.212 | 0.212 | 0.212
|  |  | T∞ | 24.7917 | 26 | 24.5 | 24 | 23.5 | 23.5 | 24 | 24
|  |  | P (W) | 0.086762 | 0.088831 | 0.217 | 0.201 | 0.276 | 0.313 | 0.474 | 0.482
|  |  | AVG Power | 0.087796 |  | 0.209 |  | 0.2945 |  | 0.478
|  |  | Std. Dev | 0.0014633 |  | 0.011314 |  | 0.026163 |  | 0.0056569
|  |  | P/V | 43.8981 |  | 104.5 |  | 147.25 |  | 239
79% | 3 L | 0.003 | UA | 0.25329 | 0.25329 | 0.253 | 0.253 | 0.253 | 0.253 | 0.253 | 0.253
|  |  | T∞ | 23.1775 | 23.9935 | 22.4 | 22.1 | 21.2 | 23.5 | 23.1 | 22
|  |  | P | 0.15 | 0.18091 | 0.4 | 0.38 | 0.54 | 0.63 | 0.95 | 0.99
|  |  | AVG | 0.16546 |  | 0.39 |  | 0.585 |  | 0.97
|  |  | Std. Dev | 0.021858 |  | 0.014142 |  | 0.06364 |  | 0.028284
|  |  | P/V | 55.1519 |  | 130 |  | 195 |  | 323.3333
Fully immersed | 4 L | 0.004 | UA | 0.274 | 0.274 | 0.274 | 0.274 | 0.274 | 0.274 | 0.274 | 0.274 | [6 empty] | Figure 4-25
|  |  | T∞ | 24.3 | 24.5 | 21.6 | 22.3 | 22.4 | 22.2 | 22.9 | 23.1
|  |  | P | 0.18 | 0.25 | 0.63 | 0.76 | 0.96 | 0.85 | 1.19 | 1.33
|  |  | AVG | 0.215 |  | 0.695 |  | 0.905 |  | 1.26
|  |  | Std. Dev | 0.049497 |  | 0.091924 |  | 0.077782 |  | 0.098995
|  |  | P/V | 53.75 |  | 173.75 |  | 226.25 |  | 315
Mixing time data
| Impeller speed | Mixing time | Std |  | Impeller speed | Mixing time | Std |  | Power
1 L (MT) | 100 | 659.3333 | 181.9924 | 3 L (MT) | 100 | 70.6667 | 1.9001
| 200 | 162 | 15.3609 |  | 200 | 24.6667 | 1.781 |  | Volume (L) | Volume(m^3) | RPM | Power (W) | Std. dev | Power (P/V)
| 300 | 136.6667 | 25.7872 |  | 300 | 19.6667 | 1.3478 |  | 2 | 0.002 | 100 | Not done
| 400 | 84.6667 | 4.7428 |  | 400 | 17.6667 | 0.68968 |  |  |  | 200 | 0.087 | 0.0028 | 43.5
| 500 | 81.6667 | 4.1366 |  | 500 | 16.3333 | 2.6717 |  |  |  | 300 | 0.209 | 0.011 | 104.5
[11 empty] | 400 | 0.2945 | 0.026 | 147.25
2 L (MT) | 100 | 72 | 6.1524 | 4L (MT) | 100 | 135.3333 | 3.7784 |  |  |  | 500 | 0.478 | 0.006 | 239
| 200 | 41.3333 | 2.9352 |  | 200 | 68.6667 | 2.986
| 300 | 27.6667 | 0.7218 |  | 300 | 46 | 1.243 |  | 3 | 0.003 | 100 | Not done
| 400 | 20.6667 | 1.7759 |  | 400 | 35.3333 | 4.8204 |  |  |  | 200 | 0.165 | 0.021 | 55
| 500 | 20.3333 | 2.6744 |  | 500 | 33.3333 | 0.74536 |  |  |  | 300 | 0.39 | 0.014 | 130
[11 empty] | 400 | 0.585 | 0.064 | 195 |  |  |  | Figure 4-26
Volume (L) | Kla (h^-1) | [9 empty] | 500 | 0.97 | 0.028 | 323.3333
| Air flow rates (LPM) | [7 empty] | 4 | 0.004 | 100 | Not done
| 0.2 LPM | 0.3 LPM | 0.4 LPM | 0.5 LPM | [6 empty] | 200 | 0.215 | 0.0495 | 53.75
2 | 5.35 | 6.64 | 7.42 | 8.22 | [6 empty] | 300 | 0.695 | 0.09192 | 173.75
| 6.92 | 8.75 | 10.03 | 9.97 | [6 empty] | 400 | 0.905 | 0.07778 | 226.25
| 7.16 | 10.07 | 10.62 | 12.77 | [6 empty] | 500 | 1.26 | 0.09899 | 315
| 7.64 | 10.85 | 12.97 | 12.54
| 8.45 | 11.43 | 13.7 | 15.97
3 | 5.34 | 5.91 | 6.04 | 6.78
| 6.89 | 8.22 | 8.54 | 9.68
| 7.84 | 10.19 | 10.5 | 10.01
| 8.94 | 10.85 | 12.06 | 13.17
| 9.29 | 11.6 | 14.25 | 15.666
[9 empty] | a | [7 empty] | b
Figure 4-27 Mass transfer efficiency Vs power input at (a) 42 % and (b) 79 % impeller immersion
𝐶_𝑝
𝐶_𝑝
𝐶_𝑝
𝐶_𝑝